## Slide 1: Dynamic Signal Measurement Basics Hardware Basics
Filters and Delta Sigma Converters
Notes: Faithful signal acquisition is essential to any successful sound and vibration measurement application. Dynamic Signal Acquisition products feature several hardware features designed to guarantee accurate frequency-domain measurements. Amplifiers, filters and converters play an essential role on this digitalization process. We will go over some of the basics of these components and their roles in dynamic signal acquisition.

## Slide 2: NI DSA Hardware Basics
Architecture
Filters
 Converters
Notes: We’ll talk about the architecture of the National Instruments 44xx and 45xx DSA Products. We’ll talk about the specifics of the digital anti-aliasing and anti-imaging filtering capabilities as well as the advantages of Delta Sigma converter technology.

## Slide 3
I/O Connector
Delta Sigma
DAC
Anti-Imaging
Atten
Amp
Ch Out
DSA
Measurement
Engine
Delta Sigma
DAC
Anti-Imaging
Atten
Amp
Ch Out
MITE
PCI Interface
Gain
Amp
AC/DC
Coupling
Ch In
LP
Filter
Gain
Amp
AC/DC
Coupling
Ch In
LP
Filter
Digital I/O
PCI Bus
Delta Sigma
ADC
Anti-Aliasing
Delta Sigma
ADC
Anti-Aliasing
DSA Hardware Architecture
RTSI Bus
Notes: The simpified architecture shown here is a two inputs / two outputs such as the one used on the 4551 and 4451. The 4452 and 4552 have four inputs and they have the same architecture than the two (upper part ) inputs shown here.
The board has AC / DC coupling capabilities that you can change from you software, you have also a minimum range of +-10mv and a maximum of +-42v on the inputs with several ranges in between (gain selections between -20dB and 60dB in 10dB increments).
All channels are simultaneous and coupling and gain can be set on a per channel basis, there is a delta sigma converter per channel and digital and analog filter per channel as well.
The 4551 and 4552 instruments have DSA Measurement Engine Processor on board and the TIO digital I/O chip . The 4451 and 4452 boards will not have the processor and will use the DAQ-STC counter timer chip. Besides that the architecture is basically the same.
The NI 4472 device (not shown in this diagram) features 8 24-bit delta-sigma analog-o-digital converters. It relies on the National Instruments DAQ-STC ASIC for acquisition timing. Like the 4451 and 4452, the 4472 does not use an on-board DSP processor.

## Slide 4: Anti-Alias Filtering
Analog Filter
  4MHz antialiasing analog low pass filter
  Eliminate out of band high frequencies
 ADC converters (1 bit x 128 oversampling)
  Immune to DNL distortion (linear transfer func.)
  Low noise,low distortion,almost perfect linearity
Digital Filter
  Auto-adjustable cutoff frequency low pass filter
  Flat ,linear phase digital filter
Notes: According to Shannon sampling theorem, the highest frequency (Nyquist frequency: fN) that can be analyzed is fN = fs/2, where fs is the sampling frequency. Any analog frequency greater than fN will, after sampling, appear as a frequency between 0 and fN. Such a frequency is known as an “alias” frequency. In the digital (sampled) domain, there is no way to distinguish these alias frequencies from the frequencies that actually lie between 0 and fN. Therefore, these alias frequencies need to be removed from the analog signal before sampling by the A/D converter.
To remove these components present at higher frequencies than the Nyquist frequency, an analog lowpass filter may be used. This anti-aliasing filter should exhibit a flat in-band frequency response with a good high frequency alias rejection and a fast roll-off in the transition band.
A key feature of delta-sigma converters is that they can incorporate robust digital sampling in the conversion process. Thus, an acquisition device based on delta-sigma technology does not need to rely on analog filtering alone to eliminate aliasing. Advantages of digital filtering include aggressive filter characteristics (sharp roll-off), temperature invariance, and a cut-off frequency that automatically tracks the sampling rate to provide robust alias protection at all sampling rates.

## Slide 5: Anti-Image Filtering
Digital Filter
  Eight times oversampling interpolators
  Linear phase digital filter (remove images)
 DAC converters (1 bit x 64 oversampling)
  Immune to DNL distortion
  Excellent amplitude flatness and THD
Analog Filters
  Eliminate quantization noise (5th order .52 s.r.)
  Eliminate images (2th order 80kHz)
Notes: Analog filtering and digital filtering with oversampling are also useful on the analog output side to provide clean stimulus signals for your unit under test (UUT). “Images” are unwanted output frequencies beyond Nyquist, and delta-sigma digital-to-analog converter (DAC) technology can minimize these unwanted signal components.
Specifics and detailed explanation of how filtering and converting is performed on the DSA boards for inputs and outputs as well as specs of filters and overall performance of the board can be found on the NI 4472, NI 4551/4552 and PCI 4451/4452 Users Manuals downloadable from our web page at www.ni.com.

## Slide 6: Analog / Digital Conversion  Converters
What is a  converter?
  1-bit oversampling digitizer
  High speed to high resolution conversion
DNL = 0 vs high DNL for traditional 16 bits ADC
Notes: Delta sigma converters are used for analog to digital and digital to analog conversion on NI DSA boards. This state of the art technology provides excellent conversion performance when dealing with dynamic signals. Its key advantages are outstanding linearity and excellent digital alias protection.
Delta Sigma converters in these boards provide 16 or 24 bits resolution. However at a fundamental hardware level these converters function as high speed 1-bit ADCs running at 64 to 128 times the user-defined sampling rate (This multiple of the user-given sampling rate is known as the “oversample frequency.”) Some of the hardware basics behind these converters are explained here.
The results of the 1-bit oversampling digitizer are noise-shaped, filtered, and decimated to provide a high resolution output.
This type of digitizer exhibits excellent differential and integral linearity and no trimming for linearity is required. The key to this excellent linearity is that the transfer function contains only two points for the two possible states of a single bit (more on this later).

## Slide 7: Where is it useful?
Covers a broad range of applications
High resolution of 24 or 16 bits for audio and vibration (dynamic signal acquisition)
Control applications where loop times are slow enough
Notes: In any application that requires analog to digital conversion, understanding the application’s dynamic range requirements is essential. In dynamic signal acquisition (DSA) applications, the requirements is typically a minimum of 90 db SNR. The input dynamic range on National Instruments DSA hardware ranges from 90 bits for the 445x and 455x devices to an outstanding 110 dB for the 4472.
There are many DSA applications where you must resolve a very weak signal component from a large excitation signal. A failure analysis on a bearing would require resolving specific minute frequency components indicating a possible future failure in the presence of many other more intense vibrations. Thus, a wide dynamic range is essential for monitoring mechanical vibrations.
Users do not typically think of delta-sigma converters as being useful in control type applications. Careful understanding of the required loop times may actually show that these types of converters are indeed useful. Special attention should be paid to the group delay incurred in the analog anti-aliasing filter which you must when a non delta-sigma converter is being used. Sometimes this delay can actually be greater than the delay through a delta-sigma converter. Delta-sigma converters can solve most control applications requiring greater than a 10 ms loop time. If the loop time is between 1 and 10 ms then you must carefully consider the delays in order to determine the usefulness of delta-sigma converters.

## Slide 8: How does it work?
Oversampling
Noise shaping (heart of )
Digital filtering
Decimation
Putting it all together

## Slide 9: Oversampling Without Noise-Shaping
Reduces requirements on analog anti-aliasing filter
Spreads quantization noise across entire bandwidth to Kfs/2 (K = oversampling ratio)
Increases data rates for a modest increase in S/N by 10log10(K) (K = 4 -> 6.02dB, K = 16 -> 12.04dB, K = 64 -> 18.06dB)
Notes: Oversampling simply redistributes the noise energy across a larger frequency range. The noise density remains the same.

## Slide 10: Effect of oversampling on filter requirements
fs
fs/2
fa
Standard Nyquist Sampling
fs
fs/2
fa
Oversampling
Kfs/2
Kfs
Kfs-fa
analog
filter
analog
filter
digital
filter
Notes: Total noise energy = q2/12 and that energy is spread out evenly across from DC to fs/2. This assumes a non noise-shaped signal which is sufficiently busy through dithering or activity in the input signal.

## Slide 11: Effect of oversampling on quantization noise
fs
fs/2
Standard Nyquist Sampling
fs
fs/2
Oversampling
Kfs/2
Kfs
rms quantization noise
q/sqrt(12)
out of band quantization noise
Notes: Total noise energy = q2/12 and that energy is spread out evenly across from DC to fs/2. This assumes a non noise-shaped signal which is sufficiently busy through dithering or activity in the input signal.

## Slide 12: Noise Shaping (heart of )
Used to keep oversampling ratio reasonable
Further redistributes quantization noise so that majority of noise lies between fs/2 and Kfs/2, i.e. above the band of interest
Function of  modulator
The higher the order (number of stages), the more noise is pushed out, but also the more unstable the modulator becomes
Notes: Again, like oversampling, noise shaping only redistributes the noise. The total noise energy remains the same. Think of a tubular balloon and imagine squeezing one end of it. This is analogous to noise-shaping, where the air has been pushed towards the other end of the balloon.

## Slide 13: Modulator
1-bit sampler running at Kfs
  Inherently linear two point transfer function (good)
  Very high RMS quantization noise of q/sqrt(12) where q = weight of LSB (bad)
 – measures difference between input signal and 1-bit feedback
 – sums up (integrates) the error voltage
Gain and integrator acts as a high pass filter on quantization noise (see modulator gain equ.)
Notes: In general this is what the Delta and Sigma heart of the converter are used for, and some of its advantages and disadvantages.

## Slide 14:  Modulator (1st order)
Modulator Gain Equation
Y = (X - Y) H(f)g + q
if loop gain H(f)g >> 1 then
Y = X + q / [H(f)g ]
Transfer function of
1 bit A/D converter
Notes: This shows the linear transfer function presented by the Delta Sigma converters and the simplified architecture of such 1 bit converters that will provide the 16 or 24 bits desired resolution and measurement quality.

## Slide 15: Effect of modulator on quantization noise
fs
fs/2
Kfs/2
Kfs
out of band quantization noise
fa

## Slide 16: Digital Filter – What It Does
Acts as an anti-aliasing filter with respect to final sampling rate, fs
Filters out the high frequency noise generated by the noise-shaping process
Filter dominates signal behavior of A/D
  Signal bandwidth, flatness of passband, phase response of passband, group delay, transition band characteristics, and out-of-band attenuation
  Settling time characteristics of transient signals
Notes: The DSA boards can be programmed to have a desired Alias Free Frequency Span, this is, the cutoff frequency of the digital filters on the board can achieve filtering on user selectable frequency spans (DC -to->Alias Free Span).

## Slide 17: Digital Filter – Advantages
Linear phase / constant group delay (FIR)
Stable over time and temperature
Tracks changes in the sample rate
Exactly reproducible
Aggressive filter characteristics at a low cost
(compared to an equivalent analog filter)
Notes: Some of the advantages of digital filtering on board.

## Slide 18: Decimation
A resampling method by which redundant signal information introduced by oversampling is removed
  Only every Kth sample is output from filter (decimate) at fs
  Can be made part of filter (FIR filter only)

## Slide 19: Effect of modulator on quantization noise
fs
fs/2
aliasing
fa
Final output is a waveform output at fs with very high S/N (low quantization noise)
Notes: The filter roll-off of the actual filter will look almost like an ideal filter. We have exaggerated the slope to illustrate the effects of quantization noise and aliasing.