## Slide 1: Architecting Hybrid Test Systems for Longevity and Performance

## Slide 2: Outline
Hybrid ATE Systems Overview
Evolution of ATE Systems
ATE Developer Needs
Five Layer ATE Architecture
System Benefits
Demonstration
Summary

## Slide 3: Hybrid ATE Systems Overview
A hybrid system combines components from multiple ATE platforms
  Streamlines system transition and maintenance
  Protects investment in existing software and hardware
  Allows easy integration of advances in ATE system development
Layered architecture is important
Multiple hybrid topologies
Notes: Hybrid ATE systems provide developers with a long-term ATE system that can evolve with short commercial technology lifecycles. Hybrid ATE systems take advantage of a 5 layer architecture to increase reusability, connectivity, extensibility, and module replacement.
By following the layered architecture, you create a test framework that allows you to replace and/or add components without having to re-architect the system. By clearly defining the system, you can take advantage of multiple buses using a variety of hybrid topologies. This allows you to meet multiple needs – protect the investment in existing hardware and software while being able to integrate the latest technologies.

## Slide 4: Evolution of Automated Test Platforms
“Big Iron” ATE
Traditional GPIB-Based ATE
User-Defined ATE
Instrument on
a Card
Proprietary Tester
Stand Alone
VXI
PC/PXI
Notes: <number>
<footer>
<header>
<date/time>
The introduction of new platforms for test and measurement has been steady for in recent history. Many of these platforms remain popular today; some cater to particular applications while others are utilized for many applications. The availability of various platforms, both new and old, can force test engineers to utilize existing test equipment with newer equipment that may be offered different platforms. Test equipment manufacturers offer various solutions to combine multiple platforms into hybrid automated test systems.

## Slide 5: Issues For System Developers
Best System for
DUT Measurement
Needs
Integrate Latest
Technologies
Preserve
Investment in
Hardware and
Software
Streamline
System Upgrades
and Maintenance
Notes: System developers want to design the best test system for their DUT measurement needs. Thus, it is important to choose the right instruments to meet those measurement needs. Beyond the measurements, a large part of creating the best test system is designing one with longevity - with the ability to preserve investment in hardware and software as well as streamline the processes for maintenance and upgrades. In addition, a good test system not only meets measurement needs but also provides the developer with the ability to enhance their test system later on with the ability to integrate the latest technologies.

## Slide 6: Goal of ATE System
ATE HW &SW
PASS
FAIL
DUT
DUT
DUT
Replacement
Reusability
Connectivity
Productivity
Extensibility
Notes: To achieve the goals of an ATE systems, the developer needs a clearly designed system. By structuring the system accordingly, the developer also benefits from an architecture that provides
Productivity – Ability to more quickly develop test programs through ADEs chosen for the test application as well as test management software that provides common components to speed up develop of a test framework.
Reusability – Provides a framework so that you can easily reuse modular test code for varying DUTs and varying tests.
Connectivity – Ability to connect between multiple platforms.
Extensibility – Ability to add additional hardware (of varying platforms if needed) and software to the system without having the reconstruct the whole test system.
Replacement – Ability to replace existing instruments without having to change whole test system. Aids in system maintenance and mid-life upgrades.

## Slide 7: Five Layer ATE Architecture
- TestStand, DIAdem
System Management
- LabVIEW, LabWindows/CVI, C++
Application
- VISA, LV PnP, IVI, MAX, DAQmx
Measurement and Control Services
- PXI, VXI, Desktop PC, GPIB
Computing
- Instruments, Data Acquisition, Vision, Motion
Device I/O
Notes: Moving from the bottom up, we’ll start with the actual hardware devices that you use to perform testing and automation. This layer includes the various instruments and data acquisition used in your system. Developers want to choose the devices that meet measurement needs and also scale with changing system needs. Oftentimes, the devices used will require multiple instrumentation buses. The computing layer above refers to the instrumentation buses used and the method and topologies that connect these buses. Users should choose a topology that meets timing and synchronization needs, allows them to protect investment in hardware, and add new components to benefit from advancing technologies. The measurement and control services is the layer that integrate the hardware with the software. It provides configuration utilities, soft front panels, and drivers with APIs to ease programming. With tight integration with the application layer, these APIs should appear as a seamless part of the application development environment. The application layer refers to the software used to create the test programs. For instance, developers might create a program to take a measurement with a DMM and another program to generate a signal. The system management layer allows users to call these programs, manage various user levels, log and manage data, and generate reports.

## Slide 8: Device I/O
DUT drives Measurement Requirements
Measurement Requirements drive Instrument Selection
Instrument Selection drives ATE bus choices
Optimize Price/Performance
Choose Measurement Functionality
DUT
Instrument
ATE bus
Price/
Performance
Measurement
Requirements
Notes: When designing a system, the first consideration is looking for instruments that meet the measurement functionality required for the system. The device under test (DUT) will drive the measurement requirements which will then drive the instruments available. When selecting instruments, developers should not constrain themselves to one bus, but rather should let the instrument selection drive the ATE buses used in the test system. From there, the developer can optimize price and performance options and iterate on the process until determining the best instruments for the system.

## Slide 9: Device I/O
Resolution
Frequency
Accuracy
Channel count
Complete measurement capability
Bus bandwidth and latency
DUT
Instrument
ATE bus
Price/
Performance
Measurement
Requirements
Choose Measurement Functionality
Notes: When evaluating the best instrument specific measurement needs, there are many factors to consider. Not only does the instrument need to meet basic needs in terms of resolution, frequency, accuracy, and channel count, but the instrument should also provide complete measurement capability. For instance, with RF instruments, the instrument might provide only common spectral measurements like peak/multipeak search, power-in-band, adjacent channel power, and occupied bandwidth. However, if application needs changed and the system now needed other spectral measurements like carrier-to-noise, phase noise, and noise-floor measurements, the developer would be forced to buy a new instrument to meet needs. Choosing an instrument with complete measurement capability ensures that the ability to adjust for changing measurement needs. Another important factor to consider is ease of use – this includes documentation, support, instrument drivers, example programs. Lastly, it’s important to consider how the instrument will scale if system needs change. If more channels are required, how easy would it be to easily add channels with tight timing and synchronization or would a new instrument need to be purchased?

## Slide 10: Device I/O
Other Considerations
Mass Interconnect
Mix and Volume
  Optimizing for low mix, high volume
  Designing for high mix environments
Product Lifecycle
  How does the product lifecycle compare to the lifecycle of test components (both platforms and devices)
Planning for replacement
Notes: Beyond measurement functionality, there are still other considerations when selecting an instruments. When designing test systems, developers should factor whether mass interconnect options are available to smoothly and easily connect to all of the instrument IO. When selecting instruments, another factor is to consider which instruments support mix and volume needs. Some hardware components are more modular and flexible and thus can be used to test a mixture of tests as opposed to instrument boxes that are designed to test just one device. Performance and throughput capabilities are also a consideration in a high volume environment.
Product lifecycle in comparison to the lifecycle of the test components is also a consideration factor. If the product/device lifecycle is expected to be longer than test system components, it is important that the developer identify components that have long form-fit-function replacements available. This will allow the user to upgrade components without having to redesign the test system.

## Slide 11: Computing
Let measurement needs drive the bus
Bus factors will affect system capabilities
  Latency
  Bandwidth
  Timing and synchronization
Various ATE bus options
  PXI
  VXI
  GPIB
  cPCI
  PCI
  Etc…
Notes: Measurement needs should drive the bus used. With the many ATE bus options available, developers should not limit themselves to the instruments available on just one bus. When choosing instruments, however, developers should acknowledge that bus factors like latency, bandwidth, and timing and synchronization will affect system capabilities. These factors determine whether you will be able perform applications like digital multimeter (DMM) measurements, switching and instrument configuration quickly enough as well as support a high channel count or a data streaming application to the extent needed.
Various ATE bus options are available including PXI, VXI, GPIB, PCI, cPCI, PXI, and more.

## Slide 12: ATE Bus Options
 | GPIB | VXI | USB | TCP/IP Ethernet | Standard PCs | CompactPCI | PXI
Transfer Width (bits) | 8 | 8, 16, 32, 64 | Serial | Serial | 8, 16 (ISA) 8, 16, 32, 64 (PCI) 8, 16, 32, 64 (PCI-X 1.0) 8, 16, 32, 64 (PCI-X 2.0) | 8, 16, 32, 64 | 8, 16, 32, 64
Throughput (Mbytes/s) | Up to 8 | Up to 160 | Up to 1.5 (USB 1.0) Up to 60 (USB 2.0) | Up to 1.25 (10BaseT) Up to 12.5 (100BaseT) Up to 125 (1000BaseT) | 1-2 (ISA) 132-512 (PCI) 264-1024 (PCI-X 1.0) 1064-4264 (PCI-X 2.0) | Up to 528 | Up to 528
Timing and Synchronization | None | Defined 10 MHz Ref Clock 8 TTL Trigger Lines 2 ECL Trigger Lines Local Bus (sub-microsecond) | None | IEEE-1588 Synchronization Protocol (microseconds) | Proprietary | None | Defined 10 MHz Ref Clock 8 TTL Trigger Lines STAR Trigger Local Bus (nanosecond)
Control Loop Rates | Seconds | Microseconds | Seconds | Seconds | No | No | Microseconds
Standard Software Frameworks | VISA available | VXIplug&play Defined | VISA available | VISA available | None | None | Defined
Modular | No | Yes | No | No | No | Yes | Yes
EMI Shielding | Optional | Defined | Optional | Optional | Board Specific | Module Specific | Module Specific
Notes: This chart provides a comparison of some of the popular buses for test and measurement applications. Some of the different consideration factors when examining buses include the throughput, timing and synchronization capabilities, and the provision of standard software frameworks.

## Slide 13: Computing
Advantages of a multi-platform system
  Extensibility
  Longevity
  Integration with latest technologies
Many system topologies and connectivity options to interconnect buses provided via controller hardware and drivers
  VXI-USB, GPIB-ENET, PXI-MXI-2, etc.
Design this layer to account for instrument connectivity needs
  Connect to legacy or specialized instruments
  Enables reuse of ATE components
Notes: Through the computing layer, developers are able to interconnect multiple bus platforms into one system. If any instruments need to be added or replaced, a multi-platform system is easily extensible and allows developers to mix and match ATE buses . Multi-platform systems also meet the need for longevity by providing a system that protects investments in hardware and software and integrate the latest technologies.
Many connectivity options are available so that developers can design a system with a topology to meet specific system needs. Many controllers use standard interfaces so that multiple connectivity options are there. You should design this layer carefully to account for any instrument connectivity issues. This allows you to incorporate legacy or specialized instruments as well as integrate newer technologies.

## Slide 14: ATE Connectivity Options
Control Interface | Bus Type
Embedded Control | PXI, VXI
PCI-to-PXI | MXI-4
PCI/PXI-to-VXI | MXI-2, USB
VXI-to-PXI1 | MXI-4
Ethernet-to-PXI | Ethernet
PC to Stand-Alone Instrument | GPIB,USB, Ethernet , Serial
Many connectivity options
Try to stay simple and modular
Keep PC and PXI in the center
1Requires VXIpc with built-in PCI expansion slot
Notes: There are many methods of control available for VXI and PXI. VXI and PXI both have embedded and remote control options available. Some of the more popular tools available to connect platforms include Ethernet, GPIB, MXI-4, and MXI-2 solutions.
Keeping PC and PXI in the center gives developers the most connectivity and interface options. With the most standard interface options available, the developer can easily connect to other instrumentation buses.

## Slide 15: Measurement and Control Services
Benefits
Separate hardware and software when debugging
Aid in replacement
Help in planning for longevity
Measurement and Control Services
Configuration
Manager
Measurement and Automation eXplorer
Application Development Environments
Diagnostic Tools
NI-Spy, GPIB Analyzer
Instrument Drivers
LabVIEW and CVI Plug and Play, IVI
APIs
NI-VISA, NI-488, NI-VXI, NI-DMM, IVI
Driver Engine
VISA, GPIB, VXI
Notes: The Measurement and Control Services Software plays a critical role in delivering the key benefits of a modular computer and networked-based measurement system. The components of this software – hardware drivers, flexible high-level application programming interfaces (APIs), and a configuration manager – must all integrate within the application development environments (ADEs) to attain maximum system performance and development productivity. The specific tasks of the Measurement and Control Services Software include integration of measurement devices, as well as local and distributed configuration and programming of the measurement devices.
Some of the benefits provided by this layer include the ability to distinguish between hardware and software debugging to simplify troubleshooting, as well drivers to help aid in replacement of hardware. Developers can use open standards like VISA, IVI, and plug&play drivers to help plan for longevity needs.

## Slide 16: Measurement and Control Services
Measurement and Automation Explorer
Configuration for hardware and software
View devices connected to your system
Create and edit channels, task, and interfaces
System diagnostics
Notes: Measurement & Automation Explorer (MAX) provides access to your National Instruments CAN, DAQ, FieldPoint, GPIB, IMAQ, IVI, Modular Instruments, Motion, NI Switch Executive, VI Logger, VISA, and VXI devices. With MAX, you can:
Configure your National Instruments hardware and software
Create and edit channels, tasks, interfaces, scales, and virtual instruments
Execute system diagnostics
View devices and instruments connected to your system
Update your National Instruments software

## Slide 17: Measurement and Control Services
  Platform independent
  Interface independent
  Must know SCPI command set to program directly with VISA
GPIB
VXI
Serial
PXI
VISA
VISA (Virtual Instrument Software Architecture)
NI-Spy
  Tracks application calls to NI-VISA, NI-488, NI-VXI, IVI class drivers, NI-Scope, NI-DMM, and NI-Switch
  Displays return code and parameters for each call
  Highlights failed calls
Notes: For many years, industry has moved toward purchasing instrumentation from a variety of vendors. This allows engineers to select the best possible equipment for their applications without being locked into a specific vendor. This trend required the definition of hardware standards to ensure the compatibility between different modules. This was one of the factors leading to the development of the VXI specification. But even with these improved hardware standards, a system was time consuming and expensive to put together. Successful integration of a multivendor system requires all hardware and software products work together, eliminating system-level compatibility issues for end-users. VISA is the underlying layer of function calls used in the LabVIEW instrument driver VIs to communicate with the driver software.
NI Spy is a valuable trouble shooting utility program that ships with the Win 32 versions of NI-488.2, NI-VXI and NI-VISA. With this tool, you can capture all function calls that your application makes into the following APIs: NI-VISA, NI-488, NI-VXI, IVI class drivers, NI-Scope, NI-DMM, NI-Switch, and more. For each call, you can see its return code and verify the parameters. Failed calls are highlighted in red to make them easy to locate during your development. Detailed information includes contents of data buffers, time stamps, process IDs, and thread IDs. You can quickly debug multithreaded test applications. You can also save this information to a file or print out a hard copy.

## Slide 18: Measurement and Control Services
Plug and Play Drivers (LabVIEW and CVI)
  Use VISA for bus interchangeability
  Include native LabVIEW or CVI source code
  Widely available
Plan for instrument replacement
  Develop separate program or functions for controlling each instrument
  Modify only these programs or functions when replacing an instrument
Notes: LabVIEW and CVI Plug and Play drivers use VISA providing bus interchangeability. This means that you could change the bus or controller used to control the instrument, but the instrument driver would still support the instrument. Plug and play drivers are widely available. When designing test programs, users should develop separate VIs for controlling each instrument so that only these programs need to be modified for instrument replacement purposes.

## Slide 19: Measurement and Control Services
IVI (Interchangeable Virtual Instrument)
  Standard instrument driver programming interfaces for common classes of instruments
  Instrument interchangeability
  Consistent API
Plan for instrument replacement
  Avoid using instrument-specific features
  Replace with an instrument of similar capabilities
Notes: The goal of IVI is to reduce the overall cost of test and increase test system longevity by defining standard instrument driver programming interfaces to common classes of instruments. By creating these standard programming interfaces for common classes of instruments, it gives users instrument interchangeability, a consistent API, simulation, and state caching.
IVI drivers allow replacement of instruments of the same type with minimal or no application software changes, but only if the application avoids using instrument-specific features and the instruments have similar capabilities.
Compared to LabVIEW or CVI Plug & Play, IVI drivers offer more features in addition to interchangeability; however, they are more complex to develop and debug, do not always come with source code, and are less prevalent. In addition, only 8 instruments classes are defined, and there is only sparse coverage of functionality within the classes. To cover all instruments, developers might have to pay for driver development.

## Slide 20: Application Development Environments
Factors to consider in choosing an ADE
  Open architecture
  Flexibility
  Ease of use
  Productivity
  Driver availability
Mix and match multiple Application Development Environments as needed
  Platform independence
  Scalability
  Technical support
  Stability
  Compiled performance
Notes: There are many requirements for ADEs that developers should look for. Some of the more important ones include an open architecture that is flexible and offers platform independence. This provides users with a application development environment that scales with their needs.. Developers should also look for ADEs that are not only easy to use but also provide technical support if needed. A key factor is the availability of instrument drivers in the ADE. Having many instrument drivers to choose from can greatly simplify programming.
Developers should mix and match ADEs as needed to meet application requirements. This way developers can incorporate existing code but also take advantage of other ADEs that offer ease of use, productivity, flexibility etc benefits.

## Slide 21: Application Development Environments
Keep applications and tests independent of ATE bus
  Write modular test programs
  Avoid direct ATE bus level calls in tests
  Use test-oriented abstractions
    Create GenerateWfm.vi instead of AWG4321_Start.vi
Benefits of modular test code
  Reusability: modular code is usable in other test programs
  Replacement: interchanging instruments is simplified
Notes: When designing a test program, it is important to create modular test code to maximize reusability. Developers should separate UUT level and system level tasks so that the programs can be easily reused for other test systems and to simplify replacement of instruments. Applications and test should be kept independent of the ATE bus by avoiding direct ATE bus level calls in tests. Using test-oriented abstractions should help to simplify maintenance of software when replacing instruments.
By creating modular test code, developers benefit from high reusability and plan for replacement. With modular test code, users can continue to use the same code in other places in the same application or in a different application without having to spend excess development time. Modular test code also simplifies the process of replacing instruments because the any code that needs to be modified is isolated from the rest of the test programs.

## Slide 22: System Management
  Supports different manufacturing models
  Archiving results
  Test reports
  Data management and report generation
  Test Framework
  Flexibility
  Operator interfaces
  User management
  Supports multiple ADEs
  Test sequencing
  Debugging
  Unit Under Test (UUT) tracking
  Test flow control
System management considerations
Notes: Many factors should be considered for system management including the user experience/environment for the developer, test engineer, and technician, as well as test program management, process management, data management, and report generation.
A system management software should not only provide a test sequencer but also a complete test framework. A test framework enables a cost-effective test development environment by allowing users to take advantage of reusable test code modules and simplify updates and maintenance while providing rapid development of test modules.
A key consideration in system management is flexibility. It should accommodate currents needs as well as adjust and scale with a growing system, changing users, and/or changing desired functionality.

## Slide 23: System Management
Benefits of system management
Improve test code reuse
Decrease development time
Extensibility
Simplify global design and manufacturing
Increase throughput
Allows merging of new test code with legacy test code
Accommodates software migration
Notes: Using system management software provides many benefits. It allows users to easily take advantage of modular test code so that they can reuse test code for varying DUTs for changing ATE systems. It helps to decrease development time by allowing users to quickly create a test sequence, incorporate report generation, and create interfaces for different users. With test management software, adding or extending the test sequence is easily accomplished by adding the new test programs to the sequence. It also provides an easy way to deploy test programs and manage data through database logging. System management software can improve test throughput by providing users with tools for parallel testing and auto-scheduling tests as opposed to limiting the user to just sequential tests. System management software also allows users to merge new test code with legacy code since it can various test programs within a sequence. This provides developers with a bridge to accommodate software migration.

## Slide 24: Demo: Wireless Communications Test System
3D Power Spectrum, Occupied Bandwidth, Auto Find (LabVIEW)
TestStand
MAX, VXI ResMan, VISA, IVI and PnP Drivers
PCI (NI-8350), PXI MXI-4, VXI-USB, GPIB
2.7 GHz PXI RF Signal Analyzer, 8.5 GHz VXI RF Signal Analyzer, PXI DMM, PXI Switch, GPIB Power Supply
Notes: For a wireless communications system, you can see how the layers work together to integrate the buses for this test application. Through the architecture the user can integrate VXI and GPIB instruments with some newer PXI instruments. This allows the user to take advantage of existing hardware and software and still incorporate newer technologies for the price/performance benefits.

## Slide 25: Demo: Wireless Communications Test System
Test Head
Test Receiver
(DUT)
2.4 GHz
900 MHz
NI-8350 (PCI)
GPIB Power Supply
PXI 4060
PXI MXI-4
VXI-USB
VXI RF Signal Analyzer
PXI 2592
PXI 4551
PXI 4552
MXI-4
USB
Power for DUT
Notes: By using MXI-4 and the VXI-USB, the developer can control PXI, VXI, and GPIB instruments from a PC. This topology allows the developer to take advantage of an existing VXI RF signal analyzer and GPIB power supply and incorporate some PXI instrumentation for the rest of the system. The controllers used provide connectivity to the various buses in this system. The measurement and control services layers provides the drivers to interface with the instruments as well as a configuration manager that provides access to all hardware and software in the system. From there, the develop can create test programs in multiple ADEs including both legacy and new test code that is managed and sequenced by the system management software.

## Slide 26: Hybrid ATE Summary
Streamlines system transition and maintenance
Maximizes return on investment in existing software and hardware
Use the five layer architecture
  Understand interfaces
  Carefully select or build components
Keep PC and PXI in the center
  Best PC connectivity
  Best instrumentation connectivity
PASS
FAIL
DUT
DUT
DUT
DUT
Notes: In summary, hybrid ATE systems allow developers to create test systems that streamline for system maintenance and transition and allow users to maximize investments in existing hardware and software. By using a five layer architecture, system developers create a system that can handle interconnecting multiple buses and extend and grow with needs. It is important to understand the various interfaces to interconnect buses and to carefully select components within the system. When designing the system, keeping the PC and PXI in the center provides the best connectivity options to PC interfaces and instrumentation.

## Slide 27: Questions?
Notes: <number>
<footer>
<header>
<date/time>
© National Instruments Corporation Acquire, Analyze, and Present with LabVIEW Seminar