## Slide 1: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
by
Jeff G. Schmitt, P.E., JGS Consulting
David A. Nelson, Nelson Acoustical Engineering

## Slide 2: About the Authors
Jeff G. Schmitt, P.E. David Nelson, P.E.
Combined 35 years experience in acoustic
measurements and noise control
Domain experts for NI Sound Power System™
Worked with most major ITTE manufacturers on
measurement and control of noise emissions
Independent consultants in the areas of:
+Laboratory Design +N and V Control +Sound Quality
+Quality Systems +Acoustical Seminars and Instruction

## Slide 3: National Instruments Developer Zone™
Visit the Developer Zone at
zone.ni.com
for detailed information on
the materials summarized
in this presentation !
http://zone.ni.com

## Slide 4: Acoustic Properties Presentation Overview
Basic Acoustics
  Sound power
  Sound intensity
  Sound pressure
Sound Pressure Level
Sound Intensity
Reverberation
Time
Absorption
Power
Transmission Loss
Acoustic Measurements
Basic Measurement
Intermediate Measurement
Material and Equipment Properties

## Slide 5: Acoustic Properties Presentation Overview
Quality systems for acoustic measurement applications
National Instruments Sound Power System application
LabVIEW™ application
Product test codes

## Slide 6: Basic Acoustics Illustrations Legend
Enclosure Walls
Finite Transmission Loss
No Sound Absorption
Enclosure lining
Finite Sound Absorption
No Transmission Loss
Acoustic Source
Microphone
P
Propagating Acoustic
Energy
Sound Pressure
at this point
General Acoustic Enclosure

## Slide 7: Measurement Block Diagrams Measurement Tools Legend
NI 455x series DAQ
Data Acquisition
with Real-Time Octave Driver
Microphone on Rotating Boom
(or microphone array)
Amplified Pink Noise Source
Computer with LabVIEW Measurement Control and Data Analysis
S and V Toolkit
Excel and Word
Custom Analysis and Reporting after Data Export
Real-Time Solution
Post-Processing
Solution
Alternate Implementation

## Slide 8: Typical Acoustical Situation
Equipment and Material Acoustical Properties
Enclosure
Reflected Sound Power = IRI2 x WI
(Absorption Coefficient  = 1 - IRI2 )
Transmitted Sound Power = t x WI
[t = Transmission Coefficient]
P1
P2
PN
Sound Pressure at Points 1 and 2
Sound Source with Acoustic Power = W
Incident Sound Power = Wi
HEAT
  Source Sound Power, W, in watts
  Sound Absorption Coefficient,  (%)
  Sound Transmission Loss, TL (dB)

## Slide 9: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
Basic
Acoustics

## Slide 10: Basic Acoustics Sound Power, Intensity, Pressure
Sound intensity— I [watt/sq. m.] is defined as the amount of sound power per square meter passing through a surface
Sound Pressure — is the expression of the acoustic power in the environment
Sound Power - W [watt], is a fundamental property of a sound source, i.e., the amount of acoustic power radiated into the environment.

## Slide 11: Sound Power The Space Heater Analogy
Space Heater
Objective
Go to the local hardware store and buy a 75 degree Fahrenheit space heater

## Slide 12: Sound Power The Space Heater Analogy
In the Arctic
Cannot maintain desired temperature

## Slide 13: Sound Power The Space Heater Analogy
In the Desert
Ambient temperature dominates.

## Slide 14: Sound Power The Space Heater Analogy
In a Small Enclosure
Desired temperature can be maintained

## Slide 15: Sound Power The Space Heater Analogy
In a Large Enclosure
Desired temperature cannot be maintained

## Slide 16: Sound Power The Space Heater Analogy
Analogies
Temperature
is a function of :
Heater power output
Distance from heater
Thermal properties
Ambient temperature
Sound Pressure
is a function of :
Source sound power
Distance from source
Acoustical properties
Ambient sound levels
Heater Power
(BTU, Watts)
Sound Power
(Watts)
Temperature
Sound Pressure
P

## Slide 17: Sound Power Basic Relationships
Sound Power — Watts
Typical Sound Power Levels
1x10-12 watts  1 W
0 dB  120 dB
Logarithmic relationship reduces range of magnitudes and tracks human perception
Sound Power Level — decibels (dB)
Lw = 10 * log (W/W0) W0 = 1 x 10-12 watts

## Slide 18: Basic Acoustics Sound Power, Intensity, Pressure
Sound power — W [watt], is a fundamental property of a sound source, i.e., the amount of acoustic power radiated into the environment.
Sound intensity — I [watt/sq. m.], is defined as the amount of sound power per square meter passing through a surface.
Sound pressure — is the expression of the acoustic power in the environment.

## Slide 19: Sound Intensity Illustration of Basic Concept
Energy distributed over spherical surface
Area = 4  r2
Free Space
r
Source
Sound Power= W
radiates uniformly
into free space
Intensity = W / 4  r2
(watts/m2)
In Free Field :
Intensity = p2 / c
Intensity is directly related to mean square sound pressure
P
Gauss’s Law
I

## Slide 20: Sound Intensity Illustration of Basic Concept
Source
Sound Power= W
radiates
into enclosed space
Energy reflected by enclosure walls back into space
P
Sound Pressure is a function of :
Incident intensity
Reflected intensity
Independent of direction of energy Flow
In non-Free Field Conditions:
No direct relationship between Sound Pressure and source Sound Power
Sound Intensity is VECTOR quantity
Guass’s Law still
holds:
I

## Slide 21: Sound Intensity Basic Relationships
Sound intensity — W/meters2
Sound intensity level — decibels (dB)
LI = 10 * log (I/I0) I0 = 1 x 10-12 watts/meters2
Magnitude and direction
  Sound pressure
  Particle velocity
  Phase relationship between P and V

## Slide 22: Basic Acoustics Sound Power, Intensity, Pressure
Sound power — W [watt], is a fundamental property of a sound source, i.e., the amount of acoustic power radiated into the environment.
Sound intensity — I [watt/sq. m.], is defined as the amount of Sound Power per square meter passing through a surface.
Sound pressure — is the expression of the acoustic power in the environment.

## Slide 23: Sound Pressure Basics
Notation/Units
Lp or SPL in decibels (dB)
Definition
The total instantaneous atmospheric pressure at a point minus the static (average) pressure at that point.
Significance
  Fundamental human hearing response parameter
  Fundamental acoustic measurement parameter (e.g. microphone is pressure transducer)
  Related to sound intensity under certain conditions

## Slide 24: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
  Measurement of
  Sound Pressure level

## Slide 25: Sound Pressure Level Measurement Basics
Acoustic Source
Generates pressure
variations
in space/time
Microphone
Generates voltage
proportional to
pressure
Voltage/Unit Pressure
Phase
L
e
v
e
l
Level
Frequency ~ 1/time
Frequency
V(t) ~ P(t)

## Slide 26: Sound Pressure Level Level Measurements
Range of typical Sound Pressures
  Threshold of hearing (0dB) = 2 x 10-5 Pascal
  Threshold of pain (120 dB) = 20 Pascal
Range is factor of 1,000,000 : 1
  Reduces range to 0 to 120 dB
  Tracks logarithmic response of human hearing
Sound pressure level
  SPL = 20 * log (p/pref); pref = 2 x 10-5 Pascal

## Slide 27: Sound Pressure Level Level Measurements
Level versus time measurements
  SPL recorded at regular time intervals.
Useful for making dynamic measurements such as
reverberation time or sound quality analysis,

## Slide 28: Sound Pressure Level Level Measurements
Time averaging
  Equivalent Sound Pressure Level = LEQ
Useful for making overall energy level measurements (Sound Power, Sound Transmission Loss, and more)
LEQ

## Slide 29: Sound Pressure Level Frequency Measurements
FFT analysis
  All bins have same width, determined by the number of “lines” in the FFT, as follows:
  Bin Width = Frequency Range / # Lines
  For Example:
    0-20kHz with 800 lines Bandwidth=25 Hz
    0-20kHz with 1600 lines  Bandwidth=12.5 Hz/line
    5kHz-10kHz with 1600 lines  Bandwidth=3.2 Hz/line
Common FFT analyzers handle 400,800,1,600 lines
Post-Process FFT analyzers handle 16,000 + lines

## Slide 30: Sound Pressure Level Frequency Measurements
FFT analyzer Display
from National Instruments Sound Power System

## Slide 31: Sound Pressure Level Frequency Measurements
Proportional bandwidth filtering
A band of frequencies around a center frequency (f0) that extends from:
  f0* SQRT (21/N) to f0 * 1/SQRT(21/N)
  1/1 Octave = 21/1
  Std Octaves:125, 250, 500, 1k, 2k, 4 k,8k (Hz)
  1/3 Octave = 21/3
  Std 1/3 OB :100,125,160,200,315,400,500,630,
  800,1k,1.25k,1.6k,2k,2.5k,3.15k,4k,5k,6.3k,8k,10k (Hz)

## Slide 32: Sound Pressure Level Frequency Measurements
1/3 Octave Band Real-Time Analyzer Display
from National Instruments Sound Power System

## Slide 33: Sound Pressure Level A-Weighting
31.5
63
125
250
500
1000
2000
4000
8000
log Frequency [Hz]
-60
-40
-20
0
R
e
s
p
o
n
s
e
(
d
B
)
A
B
C
A-weighting most often used
A-weighting approximates response of human ear
B and C weighting networks shown
Expresses overall level as single value

## Slide 34: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
  Measurement of
  Sound intensity
  reverberation time

## Slide 35: Sound Intensity Basics
Definition
Sound power radiated by a source, per unit area
Significance
Sound power can be determined by integrating Sound intensity over a surface
Notation/Units
LI or SIL in decibels relative to 10-12 watts/meter2.
May be expressed as a single A-weighted level, in 1/3-octave bands or octave bands.

## Slide 36: Sound Intensity Block Diagram
MIC
MIC
Spacer
V2(t) ~ P2(t)
V1(t) ~ P1(t)
NI 455x and DSA with Real-Time Octave driver
PC with LabVIEW Control and Analysis Software
1/3 OB L/T Display
Notes
Microphones must be phase matched
Spacer distance may vary to cover entire frequency range

## Slide 37: Sound Intensity Analysis
Time Domain
Average Pressure
Pressure Differential
Frequency Domain

## Slide 38: Sound Intensity Advantages/Disadvantages
Fundamental measurement
No special test environment
Steady state noise rejection
Noise source identification
Expensive to instrument
Multiple spacers for full range
Mathematically complicated
Many sources of error

## Slide 39: Laboratory Acoustical Environments
Anechoic Test Environment
Reverberant Test Environment
In both cases, the Sound Power of the source can be from Sound Pressure Level Measurements
Sound power from:
Sound pressure
Surface size
Sound power from:
Sound pressure
Reverberation time

## Slide 40: Reverberation Time Basics
Definition:
At any point within an enclosure, the time required for the mean square pressure to decay by 60 dB, after the source has been stopped
Significance
The total sound absorption in the enclosure is proportional to the Reverberation Time
Notation/Units:
T60 or RT (in seconds)

## Slide 41: Reverberation Time Electrical Circuit Analogy
Reverberation Chamber
Chamber Volume V
Capacitor C
Resistor R
Total Absorption A
Decay Rate  Absorption * (1/Volume)
Time Constant  Resistance * Capacitance

## Slide 42: Reverberation Time Block Diagram
Reverberation Chamber
Test Setup Legend
NI 455x and DSA with Real-Time Octave driver
V(t) ~ P(t)
PC with LabVIEW Control and Analysis Software
1/3 Octave Band L/T
Amplified Pink Noise Source
(Interrupted)
Microphone
Rotating Boom
1/3 OB PWL
Results Display
Curve Fitting

## Slide 43: Reverberation Time Analysis
Sound Pressure Level Versus Time
Least Square Error fit to early portion of level vs time decay.
slope = decay rate
d (dB/sec)
Reverberation Time
RT = 60/d (sec)

## Slide 44: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
Sound absorption
Sound transmission loss
Sound power

## Slide 45: Sound Absorption Basics
Definition
The sound absorption coefficient is the percentage of incident sound energy that, rather than being reflected, is converted to other forms of energy such as heat
Notation/units
Sound Absorption Coefficient () in percent
(dimensionless)
Basic measurements used
Reverberation Time

## Slide 46: Sound Absorption Block Diagram
Reverberation Chamber
Test Setup Legend
Export to Word/Excel
for custom analysis/report
1/3 OB PWL
Results Display
Absorption Results
NI 455x and DSA with Real-Time Octave driver
V(t) ~ P(t)
PC with LabVIEW Control and Analysis Software
1/3 Octave Band L/T
Amplified Pink Noise Source
Test Specimen
Microphone
Rotating Boom

## Slide 47: Sound Absorption Analysis
Mathematical Relationship
Aspecimen= Achamber+specimen- Achamber
A = .92*V*d / c
where:
A = total absorption in chamber
V = volume to test chamber (m3)
d = decay rate of sound pressure level (in dB/sec)
c = speed of sound in air (343m/sec)

## Slide 48: Sound Power Level Basics
Definition
Sound energy radiated by a device, per unit time.
Notation/Units:
Sound power level (Lw or PWL) in decibels (dB) relative to a reference level (usually 1x10-12 watts)
May be expressed in terms of an overall level (linear), weighted level (Awtd) or band levels (1/1 octave, 1/3 octave, etc.)

## Slide 49: Sound Power Level Block Diagram - Free — Field Method
Hemi-Anechoic Chamber
UUT
Unit Under Test
Test Setup Legend
PC with LabVIEW Control and Analysis Software
1/3 Octave Band LEQ
Export to Word/Excel
for custom analysis/report
1/3 OB RTA Display
1/3 OB Sound Power Results
Microphone(s)
NI 455x and DSA with Real-Time Octave driver
Multiplexer
V(t) ~ P(t)

## Slide 50: Sound Power Level Analysis - Free — Field Method
where:
  Lw = Sound Power of device (in dB re 1x10-12 watts)
  = Space/Time Average SPL (in dB re 2 x 10-5 Pa)
  S = Surface Area of measurement surface (in m2)
  K= Correction for effect of test environment (in dB)
  (comparison method only; direct method K=0)
Mathematical Relationship
LP

## Slide 51: Sound Power Level Block Diagram - Reverberant Field Method
Reverberation Chamber
Test Setup Legend
Export to Word/Excel
for custom analysis/report
1/3 OB PWL
Results Display
1/3 OB Sound Power Results
NI 455x and DSA with Real-Time Octave driver
V(t) ~ P(t)
Pink Noise Source Reference Source
Microphone(s)
Rotating Boom
UUT
Unit Under Test
UUT
PC with LabVIEW Controland Analysis Software
1/3 Octave Band LEQ
1/3 Octave Band L/T

## Slide 52: Sound Power Level Analysis — Reverberant Room Method
Mathematical Relationships
Direct Method
Lw = Lp-10*log(T)+10*log (V)+10*log[1+(S/8V)]-14.0
Comparison Method
Lw = Lw (RSS) – [Lp(RSS) – Lp(UUT)]
where:
  Lw= Sound Power of device (in dB re 1x10-12 watts)
  Lw(RSS)=Calibrated sound power of reference sound source
  Lp = Average SPL or RSS or UUT(in dB re 2 x 10-5 Pa)
  T= Reverberation Time in chamber (in sec)
  V = Volume of Reverberation chamber (in m3)
   = wavelength at test frequency (in meters)

## Slide 53: Sound Transmission Loss Basics
Definition
Ten times the logarithm of the reciprocal of the ratio of the energy transmitted through the specimen to the energy incident on the specimen
Notation/units
Transmission loss (TL) in decibels (dB) in octave or 1/3 octave bands
Also may be expressed in terms of the single number ratings STC, OITC or Rw
Sound transmission coefficient, dimensionless

## Slide 54: Sound Transmission Loss Measurement Block Diagram
Receive Chamber
Source Chamber
Test Setup Legend
1/3 OB Transmission Loss and STC
Test Specimen
Export to Word/Excel
for custom analysis/report
Amplified Pink Noise Source
NI 455x and DSA with Real-Time Octave driver
V(t) ~ P(t)
PC with LabVIEW Control and Analysis Software
1/3 OB LEQ
1/3 OB L/T
Rotating Boom
Microphone(s)

## Slide 55: Sound Transmission Loss Analysis
Mathematical Relationship
    TL = Lp(src) – Lp(rcv) + 10*log (S/A)
    Lp(src) = SPL in source chamber
    Lp(rcv) = SPL in receive chamber
    S = surface area of the test specimen
    A = absorption in receive chamber
    (A is determined from reverberation time in receive chamber)

## Slide 56: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
Test codes

## Slide 57: Test Codes Overview
International standard or consensus industry method
Class of machinery or equipment (e.g., computers, vehicles, mechanical equipment, doors, windows, and more)
Operational conditions and machine/material specific test requirements
Underlying standards for the measurement of the specified parameter (i.e. sound power, sound pressure, transmission loss, and more)
Measurements of other parameters

## Slide 58: Product Test Codes Examples
ISO 7779 - noise emissions of computer and business equipment
  Sound Power Level
  Sound Pressure Level at operator/bystander positions
  Prominent tone and impulsive noise analysis
  Operational condition of device

## Slide 59: Material Test Codes Examples
ASTM E1408 - sound transmission loss of door assemblies
  TL of door system
  Installation and operational conditions
  Evaluation of door seals
  Door opening force

## Slide 60: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
Quality
Systems

## Slide 61: Quality Systems in Acoustic Measurements
Calibration
  Level
  Frequency response
Documentation
  Product
  Test conditions
Test procedure
Quality assurance
ISO 17025 (Guide 25) Standard
Quality Systems in Test Laboratories

## Slide 62: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
Example
Application

## Slide 63: National Instruments Sound Power System ISO 7779 Measurement Application
ISO 7779 - noise emissions from ITT equipment
Fully integrated LabVIEW application
Utilizes NI 445x DAQ
Automates data acquisition, analysis and export

## Slide 64: National Instruments Sound Power System ISO 7779 Measurement Application
Sound Power Measurement
Real-Time Analyzer
Sound Power Analysis
1/3 Octave and Awtd PWL Results

## Slide 65: National Instruments Sound Power System ISO 7779 Measurement Application
Sound Pressure Level Measurement
Sound Pressure Level at Operator and Bystander Positions

## Slide 66: National Instruments Sound Power System ISO 7779 Measurement Application
FFT Analysis
12,800 line Post-Process FFT with “On- the-Fly” Zoom
Automated Prominent Tone Analysis Module

## Slide 67: National Instruments Sound Power System ISO 7779 Measurement Application
Quality Systems
Automated Level Calibration
Control Chart for Calibration QA
Microphone Frequency Response Corrections

## Slide 68: National Instruments Sound Power System ISO 7779 Measurement Application
Quality Systems
Test Procedure Control
Test Documentation
Module

## Slide 69: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment Presentation Summary
Basic Sound Pressure measurement :
  Sound Pressure Level
  Sound Intensity
  Reverberation Time
Material and equipment properties
  Sound Transmission Loss
  Sound Absorption
  Sound Power
Test Codes for materials and equipment
National Instruments 455x DAQ and LabVIEW solution for ISO 7779

## Slide 70: Basic Measurement Techniques for Evaluating the Acoustical Properties of Materials and Equipment
JGS Consulting
Acoustics, Technology, Quality and Business
Virtual Acoustic Instrument Development Team
JeffS@Prodigy.net; 512 282-6016