## Slide 1: Dynamic Signal Measurement Basics Sensors
Microphones
Notes: Typical sensors used for dynamic signal measurements are accelerometers, microphones, and tachometers. Other common sensors include displacement probes, velocity probes, and hammers. We will briefly introduce you to microphones, their principle of operation, and their typical characteristics. It is important to select your sensors carefully and to make sure that they will be well suited for the analysis you plan to perform.

## Slide 2: Microphones
Measure
  Sound Pressure Level (SPL) (that is, variation around the atmospheric pressure)
Result is expressed in dB (ref 20 Pa)
  SPL = 20 log10(p/p0) with p0=20 Pa
  Typical levels: 30 dB (whisper) to 120 dB (pain)
  Audible frequency range: 20 Hz to 20 kHz
Calibration is performed with a reference noise source (94 dB typical)
Notes: Microphones are basically pressure sensors, but they are dedicated to the measurement of very small variations of pressure around the atmospheric pressure.
The sound pressure level (SPL) measured by a microphone is expressed in decibels (dB ref 20 Pa), as given by the following formula:
SPL = 20 log10(p/p0) with p0=20 Pa
This reference value (giving 0 dB) corresponds roughly to the minimal sensitivity of the human ear. A sound pressure level around 120 dB corresponds to the threshold of pain.
The calibration of a microphone is performed with a reference noise source generating a constant and known sound pressure level (typically 94 dB, which corresponds almost to 1 Pa) at a particular frequency.

## Slide 3: How a Microphone Works
Most widely used type: condenser (capacitor)
Charge generated by an external polarizing voltage, or
Generated by the properties of the material itself (prepolarized)
Notes: The condenser microphone is the most widely used type. It is based on the principle that the capacitance of two electrically charged plates will change with their separation distance. In microphones, this capacitor is formed by the back plate and the light diaphragm, which moves in response to acoustic pressure variations.
The charge of this capacitor is generated either by an external polarizing voltage, or by the properties of the material itself in the case of prepolarized microphones.
Some microphones, called piezoelectric microphones, can use the same signal conditioning as accelerometers. Some microphones even make use of ICP signal conditioning to provide the polarization voltage. Many other microphones require that the polarization voltage be provided externally.

## Slide 4: Microphones
Characteristics are specific to the sound field investigated: free field (1,2) or diffuse field (3)
Microphone type:
1. Free field
2. Pressure
3. Random incidence
1. 2. 3.
Increase in size (diameter) improves sensitivity, but affects frequency range and omnidirectivity
Notes: Microphone characteristics are optimized to compensate for the disturbance caused by the presence of the microphone itself in the sound field. Free field microphones should be used when measuring in a free field (outdoor, indoor if reflection is weak) and should be pointed in the direction of the noise source. Pressure microphones can also be used in a free field, but should instead be held perpendicular to the direction of the propagation of sound. Random incidence microphones are designed to make measurements in a diffuse field (highly reflective room), where the sound arrives from any direction.
Increasing the size of a microphone improves its sensitivity, but it also affects its frequency response and its omnidirectivity. The smaller a microphone is, the less it will interfere with the sound field, especially at high frequency where the size of the microphone becomes of the same order of magnitude as the wavelength.

## Slide 5: Typical Frequency Response
Notes: A typical frequency response of a microphone is displayed on this graph.

## Slide 6: Typical Characteristics
 | Typical
Frequency range | 0.1 Hz to 140 kHz
Sensitivity | 1 to 100 mV/Pa
Polarization voltage | 28 V; 200 V 0V (prepolarized)
Dynamic range | 10 to 187 dB
Diameter | 1/8” to 1”
Notes: Some typical characteristics of microphones are given in this table.
No single microphone can cover the whole range of characteristics mentioned here. The choice of a particular sensor is thus always a compromise.

## Slide 7: Signal Conditioning
Integrate Electronic Piezoelectric
Charge amplifier
Filter cutoff frequencies
Integration
Microphone polarization
Notes: IEPE signal conditioning provides a constant-current source to power internally amplified sensors. These transducers are simple to use and set up and relatively inexpensive.
Charge mode accelerometers do not have built-in signal conditioning and require external charge amplification and conditioning to convert charge to voltage.
Vout = q/Cf
The advantage of charge mode sensors is that they can typically be employed in harsher and hotter environments that would damage the sensitive electronics in IEPE devices. However,charge mode amplification poses some limitations, especially when driving long cables whose capacitance will affect the measurement. Also, the per channel cost is typically much higher than IEPE sensors.