## Slide 1
Using the Internet to Collect and Monitor
Data at The San Francisco Bay Model

## Slide 2
Who Is Cal-Bay Systems?
National Instruments Alliance Program Member since 1993
General Purpose Systems Integrator familiar with most all of the NI product line
Software expertise in National Instruments LabVIEW, LabWindows/CVI a component of Measurement Studio, Microsoft Visual Basic and Visual C++, and Internet Programming
Custom programming, hardware integration, driver development, custom hardware development, field installations, startup, and general training
Industries — Semiconductor, biomed, pharmaceutical , automotive, petrochemical, chemical, manufacturing, automated test
Applications — contract programming, driver development, process control, data acquisition, embedded control, instrument interface, turn-key systems

## Slide 3
Using the Internet to Collect and Monitor
Data at The San Francisco Bay Model

## Slide 4
What Is the Bay Model?
The San Francisco Bay and Delta are the West Coast’s largest estuary encompassing 1,600 sq. miles of waterways, draining over 40 percent of California’s fresh water.
The Bay Model is a three dimensional representation of the San Francisco Bay.
Horizontal scale — 1 foot = 1,000 feet
Vertical scale — 1 foot = 100 feet
Time scale — 1 minute = 1 hour and 40 minutes (14.9 minutes = 1 lunar day)
Model dimensions — 320 feet x 400 feet
The model simulates tides, currents, river inflows, mixing of salt and fresh water, and sedimentary movement.
The model is used as a scientific, hydraulic laboratory to measure the effects of changes on the physical tide forces of the Bay and Delta region.
Creation of new levees/dams
Demolition of old levees/dams
Dredging new shipping channels
Sedimentary movement

## Slide 5
Automation Requirements
Tide control — precise tidal conditions are replicated within .1 ft (.012 in. in the model) with 28 day lunar tides used as setpoint data. This includes spring (full moon) and neap (no moon) tides. Tidal variations of up to 7 ft (.84 in.).
High channel count data acquisition — there are over 150 sensors around the model for collecting water depth, salinity, and velocity data.
Calibration — all 150 sensors are calibrated before every test.
Data presentation — data is presented to a wide variety of engineers for analysis over the Internet.
Distributed control — flow controllers are used to simulate river inflows from the Sacramento-San Joaquin Delta and are located at several distant places around the model.
Consistent, easy-to-use programming interface that handles all aspects of the application

## Slide 6
Automation Requirements
(Map of Sensors and Transducers)

## Slide 7
Automation Requirements
(Map of System)
Internet
Wireless Ethernet
Ethernet

## Slide 8
Data Acquisition
Water Level Sensor in the Model

## Slide 9
There are 150 analog sensors around the model. All sensors are conditioned to output a 0-5 VDC signal.
Sensor types are
  Salinity (temperature and conductivity)
  Water depth
  Velocity (X and Y direction)
Used eight SCXI 1100 (32-channel multiplexers ) with AT-MIO-16XE50
Data is collected every two seconds.
Tests are run in the model over a period of two weeks (equivalent of 3.6 years in real life)
Raw data is logged to an SQL server database and archived onto writeable CD-ROMs.
Calibration constants are stored in the database (over wireless Ethernet), and data is scaled when recalled for viewing.
Data Acquisition
Salinity (temperature and conductivity)
Water Depth
Velocity (X and Y direction)

## Slide 10
Calibration
Wireless
Ethernet

## Slide 11
Calibration
A LabVIEW app handles calibration procedures, using “Wizard” like interface.
Calibration is done on laptops using PCMCIA DAQCards™ and LabVIEW for data acquisition/interface.
Calibration constants are written to calibration table in SQL database on server.
Because of the large area, we needed a way to transfer calibration data directly to the server. The calibration can take a long time.
Using wireless Ethernet to transfer data back to the DAQ server

## Slide 12
How can we use
the local Intranet?
User wanted an easy way to monitor the data from different locations
The tides are run almost every day (especially Friday and Saturday)
The visitor center could use the web-based monitoring to show visitors water depth data in real time

## Slide 13
Using the Intranet
Internet

## Slide 14
Data Presentation on
the Internet — server
Windows/NT server as the internet server engine
Internet server doubles as the data acquisition server, running LabVIEW and housing all database tables
Using Internet information server (comes with Windows/NT server), we wrote “server side scripting,” which queries the database and pumps data to the browser
Fully integrated with all the other Microsoft tools
VB-script and Active-X controls (Measurement Studio) are used for presentation and user interface in the browser
The Access database acts as a ‘global variable’. We write to it (the new data) and read from it (to display the latest data on the web).

## Slide 15
Data Presentation on
the Internet — server
Server code written using ASP (Active Server Pages)
The server builds the page dynamically based on what the client requests. (The processing is all done on the server side.)
The server then serves pure HTML back to the client.
The client could be a low-end PC running Internet Explorer.

## Slide 16
Data Presentation
on the Internet — Client
Any PC on the internal network that can access the server
The client first gets a map of the model with sensor locations highlighted.
User clicks on a sensor and gets more information on it.
User can also view the data in real time using a Measurement Studio graph (ActiveX object) embedded in the page. The graph runs on the client, but the data is pumped from the server.
The Web page displays dynamic information (updates every 10 seconds).

## Slide 17
Challenges
We needed to build the map of the sensors dynamically. A separate application on the server creates the map, places the sensors on the map (as red dots), and then saves it as a JPEG file. This file is then transferred to the client.
Using ASP, we do both monitoring and control. When the user clicks on the map, we translate the x and y coordinates of the map into a sensor location. The sensor data is then queried from the database and displayed on the graph.
We’ve applied the same technology to other projects where the customer needed to monitor a building’s fire alarms, temperature, humidity etc.

## Slide 18
Summary
We used the strengths of many software and hardware tools in this project:
  LabVIEW for data acquisition
  Access database for storing and retrieving data
  SQL toolkit for communicating between Access and LabVIEW
  SCXI hardware for signal conditioning and multiplexing
  PCMCIA hardware for “on-the-spot” calibration
  Microsoft Internet information server (IIS)
  Microsoft ASP language for creating dynamic HTML pages
  Measurement Studio for displaying data in a browser
We’ve applied the same technology to other projects where the customer needed to monitor a building’s fire alarms, temperature, humidity etc.
In both cases, the enabling technology is National Instrument hardware and software, Microsoft ASP language.

## Slide 19
Visiting the Bay Model
2100 Bridgeway, Sausalito, CA
Hours — Tues. — Sat. 9:00 a.m. — 5:00 p.m.
Closed Sunday and Monday
http://www.spn.usace.army.mil/bmvc/