## Slide 1: Dynamic Signal Measurement Basics Sensors
Accelerometers
Notes: Typical sensors used for dynamic signal measurements are accelerometers, microphones, and tachometers. Other common sensors include displacement probes, velocity probes, and hammers.
Even before doing any actual measurements, it is important to select your sensors carefully and to make sure that they will be well suited for the analysis you plan to perform.
In the next slides, we will briefly introduce you to accelerometer transducers, their principle of operation, and their typical characteristics.

## Slide 2: Vibration Components
Amplitudes are related by a function of frequency and time for sinusoidal excitations
Acceleration is the most commonly measure component
Acceleration is used to compute the frequency response due to a known force input
Acceleration, Velocity and Displacement

## Slide 3: Why Measure Vibration?
Predict structural response of an object in order to modify or reduce vibrations
  Sometimes small vibrations can excite resonant frequencies of a structure and amplify them into major vibrations and noise sources
In an ultrasonic cleaning bath, this is GOOD.
In an airplane, this is BAD.
Notes: Vibration tests are run for a number of reasons including:
1. Engineering development testing
2. Qualification testing
3. Reliability qualification testing
4. Production screening testing
5. Machinery condition monitoring
Vibration testing is the art and science of measuring and understanding a structure’s response while exposed to a specific dynamic environment; and if necessary, simulating this environment in a satisfactory manner to ensure that the structure will either survive or function properly when exposed to this dynamic environment under field conditions.

## Slide 4: Accelerometers
Measure
  Acceleration
  Velocity and displacement (via integration versus time)
Result is expressed in m/s2 or g
  1g : acceleration at the surface of Earth
  1g = 9.81 m/s2
1D or 3D (triaxial) accelerometers
Calibration is performed with a vibration shaker
Notes: The original purpose of any accelerometer is of course to measure the acceleration level at a particular point of the system under test. But this signal, if properly integrated, also allows the simultaneous measurement of the velocity and the displacement.
Vibration parameters are almost universally measured in metric units (m/s2). The gravitational constant “g” is often used for acceleration, but is not a standard metric unit of measure. Fortunately, a conversion factor exists with which you can perform quick mental calculations with only 2% of error.
1.0g = 9.81 m/s2
Simple (1D) accelerometers measure the component of the acceleration in one particular direction. If three of these accelerometers are combined in a single unit (3D accelerometer), the acceleration can me measured simultaneously in three orthogonal directions.
The calibration of an accelerometer is typically performed with a shaker generating a constant and known level of acceleration at a particular frequency.

## Slide 5: How an Accelerometer Works
Relies on the piezoelectric effect
Other design: shear mode (less temperature sensitive)
Conditioning: built-in microelectronics (ICP) or external (charge mode)
Typical compression
mode accelerometer
Notes: Most accelerometers rely on the use of the piezoelectric effect. In the case of compression mode accelerometers, the acceleration of the test structure is transmitted to a seismic mass that generates a proportional force on the piezoelectric crystal. This external stress on the crystal then generates an electrical charge proportional to the applied force, and thus also proportional to the acceleration. Another common design is the shear mode, which is less sensitive to temperature effects (the charge is picked up in a direction perpendicular to the polarization direction).
The piezoelectric accelerometer has some very important characteristics:
Very wide frequency and dynamic range with good linearity throughout
Robust and reliable because there are no moving parts
Stable characteristics over time
Self-generating, as it does not need an external power supply
Some accelerometers use built-in microelectronics (ICP) and require only a constant-current source to power internally amplified sensors. Other accelerometers (charge mode) do not have built-in signal conditioning and require external charge amplification and conditioning to convert charge to voltage.

## Slide 6: ICP Accelerometers
Advantages
  Simple and easy to use
  Built-in microelectronics
  Simple constant current signal conditioner (18-30 VDC ; 2 mA)
Limitations
  Temperature range: max 250 °F (some 325 °F)
  Fixed sensitivity
Notes: ICP signal conditioning provides a constant-current source to power internally amplified sensors.
ICP is a registered trademark of PCB Piezotronics and stands for Integrated Circuit Piezoelectric. All major accelerometer suppliers have their own equivalent solution, such as Isotron from Endevco, DeltaTron from Brüel & Kjaer, and Piezotron from Kistler.
The major advantage of the ICP accelerometer is that it is simple and easy to use due to the built-in microelectronics and the simple conditioner required.
Limitations are a limited temperature range, especially for high temperatures, and a fixed sensitivity.

## Slide 7: Charge Mode Accelerometers
Advantages
  Temperature up to 1000 °F
  Variable sensitivity
Limitations
  External conditioning required
  Need low-noise cabling
  Sensitive to environmental influences
Notes: Charge mode accelerometers are used essentially for measurements in high-temperature environments. They also offer variable sensitivity.
However, as they do not have built-in signal conditioning, they require an external charge amplification and conditioning to convert charge to voltage.
Vout = q/Cf
They also require low noise cabling and are sensitive to environmental influences.

## Slide 8: Typical Frequency Response
Frequency/natural resonant frequency
Acceleration output/acceleration of the structure
Notes: The upper frequency limit is determined by the resonant frequency of the mass-spring system of the accelerometer itself. When using any accelerometer, filtering is a must to remove this resonant frequency peak. Failure to do so could create an overload condition that will corrupt data and measurements.

## Slide 9: Typical Characteristics
 | Typical | Specific
Frequency range | 0.2 Hz to 25 kHz | 
Sensitivity | 10 to 1,000 mV/g | 1 to 10 V/g (seismic)
Mass | 5 to 40 grams | 200 grams (seismic)
Amplitude range | 500 g (peak) | 100,000 g (shock)
Notes: Some typical characteristics of accelerometers are given in this table. Two important parameters are the frequency range, which should be carefully selected based on the problem under investigation, and the mass of the sensor, which could affect the dynamical behavior of the structure under test, especially for light parts.
No single accelerometer can cover the whole range of characteristics mentioned here. The choice of a particular sensor is thus always a compromise.
Specific characteristics are available for special measurements, like seismic and shock analysis.

## Slide 10: The Piezoelectric Accelerometer
Commonly used for vibration measurements
Generates an electrical charge (q) across its pole faces which is proportional to the applied acceleration
Notes: The word piezoelectricity literally means “pressure electricity” in Greek.
Piezoelectric accelerometers use a polarized material, which when mechanically stressed, will generate an electrical charge which is proportional to the applied force. There are two basic materials used – quartz and ceramics. Each type of crystal possesses unique characteristics that can enhance the performance of an accelerometer.
The piezoelectric accelerometer has some very important characteristics including:
1) Very wide frequency and dynamic range with good linearity throughout
2) Robust and reliable because there are no moving parts
3) Stable characteristics over time
4) Self-generating as it does NOT need an external power supply
The piezoelectric accelerometer is also a good choice because you can integrate its acceleration-proportional output to get velocity- and displacement- proportional signals.
In practical accelerometer designs, the piezoelectric element is arranged so that when the accelerometer is vibrated, an internal mass applies a force to the element which is proportional to the vibratory acceleration according to the Force equation:
Force = Mass * Acceleration

## Slide 11: Accelerometer Types
Tension
Compression
Shear
Notes: In tension/compression deformation the charge is picked up in the polarization direction. This has the distinct DISADVANTAGE that non-vibrational inputs like temperature fluctuations cause charge to be developed in the polarization direction and thus induce error.
In shear deformation type transducers, the charge is picked up in a direction perpendicular to the polarization direction and therefore temperature effects are not picked up.

## Slide 12: Accelerometer Material Types
Piezoelectric man-made CERAMICS
  Begins to depolarize if temperature range is exceeded (usually above 250° C)
  Temperature dependent sensitivity
  High charge sensitivity
Piezoelectric QUARTZ
  Naturally piezoelectric and therefore very stable
  Temperature stable
  Low capacitance and low charge sensitivity
Notes: Quartz is widely known for its ability to perform accurate measurement tasks. Since quartz is naturally piezoelectric it has no tendency to relax to an alternative state and is considered the most stable of piezoelectric materials. Also, quartz does not suffer from temperature change effects, which provides stability in thermally active environments. The useful temperature range of quartz is limited to approximately 350º C.
A variety of ceramic materials are used for accelerometers, depending on the requirements of the particular application. All ceramics are man-made and are forced to be piezoelectric by a polarization process. If the material is exposed to out-of-range temperatures, strong electric fields, or accumulates high levels of static charge, then the material may depolarize. The main advantage of ceramics is that you can tailor it to fit a particular application or environment.

## Slide 13: Accelerometer Signal Conditioning
High impedance (voltage mode) amplifiers
  High impedance devices, like piezoelectric, require attention to wiring capacitance and input characteristics of amplifiers
Charge type amplifiers
  These produce a voltage proportional to the input charge
  Most widely used type of amplifier because of superior characteristics
Notes: The choice of accelerometer amplifier depends on whether we want to detect charge or voltage as the electrical output of the accelerometer.
The distinct advantage of the charge amplifier is that you can use both very long and very short cables without changing the overall sensitivity of the system. However, when you use a voltage amplifier, a change of cable type or cable length will necessitate the recalibration of the system sensitivity. This recalibration is necessary as the voltage produced by the accelerometer is divided between the accelerometer’s capacitance and the cable’s capacitance. Though the sensitivity of a charge amp will not be significantly attenuated by changes in cable length, the use of very long cables will slightly attenuate the high frequency response.
Because of the high impedance output characteristics of the accelerometer, external noise can easily couple into the cable and cause error. One method of limiting this error is to include a preamplifier inside the accelerometer, which lowers the output impedance of the accelerometer and reduces its sensitivity to external noise coupling. The major disadvantage of this system is that this type of accelerometer is not rugged and is restricted in its usable temperature and dynamic range. Brüel & Kjær uses the Line-Drive system, PCB uses the ICP system and Endevco uses the ISOTRON system to integrate the preamp in the accelerometer.

## Slide 14: Accelerometer Signal Conditioning
Amplify input signal to take advantage of full dynamic range of converters
Warn of overload conditions at the input of preamp
Filter the input signal rejecting unwanted low and high frequency components
Provide constant current power source if the transducer contains an integrated preamp (ICP)
Notes: Integrated circuit piezoelectric (ICP) – Term describing a voltage mode sensor from PCB Piezotronics that combines an integrated preamp circuit and the piezoelectric sensing element in a single housing to provide a voltage output.
These types of transducers require a constant current power source for proper operation. This signal conditioner consists of a well-regulated 18 to 30 VDC source in series with a current-regulating circuit like a diode and a capacitor to remove the DC-bias from the signal going to the measuring system. The current supplied should be a minimum of 2 mA and a max of 20 mA per transducer.
The Brüel & Kjær in-line system requires a 15 VDC power supply and recommends using one of their line-drive supplies like the model 2813.
Endevco’s ISOTRON system requires 18 to 24 VDC with currents of 2 mA to 20 mA.
To be compatible with most integrated preamps, it is sufficient to supply 4 mA at 24 VDC.

## Slide 15: Some Rules of Thumb
Transducer mass < 0.1 mass being measured
You determine the useful upper frequency limit by 1/10 of the mounted resonant frequency (fr) of the accelerometer
Use a lowpass filter to remove frequency data amplified by the resonant peak
Notes: High sensitivity (high output levels) usually means using a large piezoelectric assembly and therefore a large and heavy unit. Because you want the assembly mass to be a maximum of 0.1 * mass of object being measured, pick the accelerometer required based on weight first. You can compensate for the lower output level by using a good analog preamplifier stage. Once you have chosen a weight class, select the transducer with the highest sensitivity you need or can afford.
The useful acceleration range is determined at the lower limit by the total system noise including cables and circuitry. You determine the upper limit by the structural strength of the transducer.
You determine the useful frequency range at the lower limit by either the electronics or temperature effect fluctuation. Because in shear-type accelerometers the temperature effect is minimal, the lower limit is typically well below 1 Hz. At the upper frequency the limit is determined by the resonant frequency of the mass-spring system of the accelerometer itself. There are differing opinions on what you should set the upper range to, but the most conservative is the 1/10 fr of mounted accelerometer. By choosing the useful limit to be 1/3 fr of mounted accelerometer , the components measured at the upper frequency will have no more than 12 percent error (1dB). You must ultimately decide what upper limit makes sense for your application and what amount of error you can tolerate.

## Slide 16: Some Rules of Thumb
Maximum vibration level should not exceed 1/3 of shock rating
Operating a transducer beyond its useful range will probably cause non-linear results