## Slide 1: Using Calibration to Improve Measurement Accuracy
Paul Packebush
Calibration Engineer
Wed Aug 16
12:00-1:15 p.m., 3:30-4:45 p.m.
Hickory (9C)
Notes: Welcome to our presentation on instrumentation accuracy and calibration.
For many applications, the accuracy of your measurement system can have a dramatic effect on your company’s productivity and profitability.
Example – manufacturing
Effects on production yield:
  passing bad parts – deliver bad parts to your customers
  failing good parts – throw away usable products, and waste time making more to suppliment those that failed
So, understanding your instrument’s accuracy affects your bottom line.

## Slide 2: Calibration & Accuracy Agenda
Defining instrumentation
Determining measurement accuracy
Maintaining accuracy via calibration
Proving your calibration
Improving accuracy via system calibration
Calibration services
Notes: Specifically, we will address:
What components make a device an instrument
How to determine the uncertainty of your measurements via accuracy tables
What external factors can alter a device’s accuracy
The role of various calibration technologies and services to ensure an instrument delivers accurate measurements

## Slide 3: Defining Instruments
Measurement circuitry
Onboard memory
Voltage reference source
Calibration circuitry
Power supply regulation
Processor/computer
Vendor defined
User defined
Notes: Think about a device at your company that you consider as an instrument.
What components does it have that make it an instrument?
  Processor/computer – controls the instrument
  Measurement circuitry – A/D, signal conditioning
  Power supply regulation – maintains constant, steady supply of power to your circuitry
  Calibration circuitry – so you can adjust your measurement circuitry when necessary
  Voltage reference source – compensates for changes in your environment
  Onboard memory – stores calibration results and uses them
An instrument will have all of these components!
Which one of these devices is an instrument? – BOTH!
  One is vendor-defined (functionality defined by the manufacturer)
  Other is user-defined
    You decide the acquisition, analysis, and presentation
    You leverage the advancements in computers to build your own personal measurement solution to meet your app. requirements

## Slide 4: Calibration & Accuracy Terms
Absolute accuracy
  Specifies the maximum possible uncertainty in a measurement
Relative accuracy
  Difference between two or more measurements
Calibration
  Process of determining and adjusting an instrument’s accuracy
Notes: From this point on, we’ll talk in terms of instruments
Now, we’ll define some terms that we’ll use throughout the presentation
Absolute accuracy – Describes the maximum possible uncertainty in your measurement
  We’re not describing error (a wrong measurement), but a window where we know the signal lies
  Really describes how close your measurement is to the actual value
Relative accuracy – Describes the difference between two measurements
  Also know as repeatability or precision
  A function of the device’s A/D and system noise
  Really describes the smallest change in a signal that you can detect
  A device can have poor absolute accuracy, and good relative accuracy
  Imagine a 16-bit device (good relative accuracy) that always measures 1 V too high (poor absolute accuracy)
Calibration – Process of determining and adjusting an instrument’s accuracy
  We’ll cover this in more detail later

## Slide 5: Taking Measurements
Measurement + uncertainty
Uncertainty is defined by the device’s Absolute Accuracy
2V
+/- 0.0002V
Notes: With these terms defined, consider an instrument:
You take a measurement – 2 Volts
Does that mean your signal was exactly 2 Volts? – NO
Your system has some uncertainty – example: +/- 0.0002 V
The uncertainty is defined by the device’s Absolute Accuracy.
Question – How do you determine your instrument’s uncertainty?

## Slide 6: Calculating Accuracy
Absolute accuracy = (Input voltage x % of Reading) + Offset + System noise + Temperature drift
Device
Nominal Range
Absolute Accuracy
% of Reading
Offset (uV)
System Noise (uV)
Temperature Drift
% of Reading/C
Offset (V/C)
6052E
 10 V
+0.0304
1067
981.0
 0.0006
 24
 5 V
+0.0304
536.2
490.6
 0.0006
 14
 2.5 V
+0.0304
270.6
245.3
 0.0006
 8
 1 V
+0.0304
111.2
98.1
 0.0006
 6
 0.5 V
+0.0304
58.1
56.22
 0.0006
 5
Notes: Answer – Use the instrument’s Absolute Accuracy Table
All instrument manufacturers provide absolute accuracy tables
The values in these columns combine multiple measurement factors (gain errors, offset errors, non-linearities, etc.)
Plug these numbers into the equation to determine your instrument’s accuracy at a given input range
  Temperature effects are included for specified ranges
  *Outside this range, you include the temperature drift affect

## Slide 7: Example Calculation
Absolute accuracy = (Input voltage x % of Reading) + Offset + System noise + Temperature drift
Device: PCI-6052E Multifunction I/O Device
Signal: +/- 2 V
Environment: 25 °C
For the +/- 2.5 V input range:
% of Reading = 0.0304%
Offset = 270.6 uV
Noise = 245.3 uV
Absolute accuracy = +/- 1.12 mV
Notes: (Hidden)
In the example above, we are trying to read a 2 V signal
Use the terms in the table for the +/- 2.5 volt input range
We do not need to include temperature effects for this example, already factored in.
That’s 0.045% accurate

## Slide 8: Calculating Accuracy (Multiple Components)
Calculate accuracy for each stage
Calculation based on signal level at the input of that stage
Required for systems with multiple-gain stages
System accuracy =
(Stage 1)2 + (Stage 2)2
Stage 1
Stage 2
Notes: You may have a system that has multiple components
Stage must have an amplifier to be considered (terminal blocks or connector blocks don’t count)
You calculate the uncertainty for each stage:
Consider the input signal value at that stage
Example
  1 mV signal into stage 1, calculate accuracy for 1 mV input signal
  Stage 1 applies a gain of 1000, so input signal is 1V at stage 2
  Calculate accuracy for stage 2 based on 1V input signal
Convert each stages uncertainty into a percentage relative to the input signal
  (Accuracy)/(Input Signal)
Take the Root-Sum-of-the-Squares (RSS) of each stage to determine system uncertainty

## Slide 9: Example Calculation
Stage 1 – SCXI-1102 Amp/Mux
Input Range – 100 mV
% of Reading – 0.02%
Offset – 25 uV
Noise – 20 uV
Stage 1 = 95 uV
= 0.38% RTI
System accuracy =
(Stage 1)2 + (Stage 2)2
Stage 1
Stage 2
Stage 2 – PCI 6052E MIO
Input Range – 2.5 V
% of Reading – 0.0304%
Offset – 270 uV
Noise – 245 uV
Stage 2 = 1.12 mV
= 0.04% RTI
25 mV
System accuracy = 0.382%
Notes: (Hidden)
Example – measure a 25 mV signal
Calculate Stage 1 – SCXI-1102 Amplifier/Multiplexer in this example
95 uV for the 25 mV signal
0.38% Relative to Input (RTI) – also called Relative to Source (RTS)
Calculate Stage 2 – PCI 6052E Multifunction I/O Device for this example
Remember, the input signal is now 2.5 V when it reaches the DAQ device
Calculation (from previous) – 1.12 mV
0.04% RTI
System Uncertainty – RSS(stage1 & stage2)
0.382% Relative to input of the system

## Slide 10: Calculating Accuracy (Sensor)
Convert sensor value to voltage
Calculate accuracy for that voltage
Convert absolute accuracy back to unit of measure
Result
5.268 mV  ± 0.5%
5.242 mV  99 °C
5.294 mV  101 °C
Accuracy = ± 1 °C
100 °C  5.268 mV
Notes: Accuracy tables are commonly given in terms of voltage.
If you need it in terms of another unit of measure, you can calculate it.
Example – measuring a thermocouple
Convert thermocouple temperature into a voltage (from standard TC table)
Calculate accuracy for that voltage
Calculate the possible voltages from the uncertainty
Convert these values back to temperature
Now that we understand your instrument’s accuracy – What can affect its accuracy?

## Slide 11: Effects of Environment
2.0001 V
1.9999 V
Notes: Components behave differently in different environments
These effects can cause your device to report measurements outside of your calculated uncertainty
In the above example, which answer is right?
Both – as long as the board has been calibrated to its environment.
Any measurement differences should be within the “uncertainty” limits of your device
The process of calibrating your device to your environment is called “internal calibration”, (covered soon)

## Slide 12: Effects of Time
Time
Error
24 Hrs
90 Days
1 Year
0.01%
- 0.01%
Notes: Component’s behavior tends to drift over time
Common Question – What is this time line with respect to?
  In other words, when does the clock start? It starts when you use the product. You could have the board sitting on a shelf in a controlled environment,
  for up to a year, plug it in, and expect the 24 hour accuracy data (provided you perform an “internal calibration” after you install the device, which will cover soon.
(More slide building to come…)

## Slide 13: Effects of Time
?
Time
Error
24 Hrs
90 Days
0.01%
- 0.01%
1 Year
Notes: The problem is that it is hard to know exactly how the components will drift
This leads to greater uncertainty as time progresses
Some instruments account for this by publishing accuracy for the worst case
I.E – take the uncertainty at one year, and apply it everywhere
More sophisticated instruments supply graduated accuracy tables
This provides a better uncertainty profile for the life of your instrument
At some point, the drift causes the instrument’s uncertainty to become undefined
This just means the manufacturer can no longer predict the uncertainty, and guarantee the results
So, what do these effects mean to you?

## Slide 14: Temperature and Time Considerations
Increased measurement uncertainty
Lack of measurement consistency
Reduced production yields
Failure to meet ISO-9000 requirements
Resolve these issues with calibration
Notes: Environmental changes and aging means the manufacturer can no longer predict the uncertainty of your measurements.
This can result in:
Increased measurement uncertainty – larger possibility of error in your measurements
Reduced measurement consistency – The measurements of today are not the same quality as those of yesterday or those tomorrow
Lower production yields – you pass parts that are bad, and fail parts that are alright. Or, you may receive bad data, and make decisions based on that bad data, costing you time and money
Failure to meet ISO-9000 requirements – These standards require you use equipment with known uncertainties.
You can resolve these issues with calibration

## Slide 15: Calibration
Component calibration
  Consider each component as individual units
    Internal calibration
    External calibration
System calibration
  Consider entire system as a unit
  Compensates for other system errors
Notes: There are different levels of calibration
Component Calibration – consider each device as a separate unit
Common examples of component calibration are internal calibration and external calibration (we’ll discuss soon)
System Calibration – consider the entire system as a unit
With this, you can compensate for other system errors (cables, sensor, differences in each unit of a multi-device system)
We’ll discuss this soon also

## Slide 16: Component Calibration
Calibration of individual system components
Determines each component’s uncertainty
With calibrated components, you can:
  Calculate system’s absolute accuracy
  Replace devices
  Change computers
0.01%
0.01%
0.014%
Notes: Component Calibration
Calibrate each component separately
NOTE – Your system may only have one component
Determine the uncertainty of each component
Use the “Calculating System Accuracy” equation to determine overall accuracy
When each component is calibrated, you are now free to:
Swap in and out any calibrated components without changing your system accuracy
Move your system to other environments without changing your system accuracy

## Slide 17: Internal Calibration
Adjusts the instrument to your environment
  Change in temperature
Maintains uncertainty defined by calibration
Requires an onboard reference source and temperature sensor
Notes: You must always adjust your instrument to changes in the environment
Examples include:
Change in temperature
Changing computers (really just another change in temperature)
You compensate for these changes with an “Internal Calibration”
Adjusts your measurement circuitry
Maintains the uncertainty specification of your instrument
Requires an onboard reference source where you know how temperature affects the source’s value. Also requires a temperature sensor so you know the temperature in your environment.
Please note – this is required for traditional instruments (like scopes) and our instruments. In other words, it is not something special we have to do because we are computer-based. Every instruments has to do it.
Our instruments do this – we call it “self-calibration”

## Slide 18: External Calibration
Defines uncertainty of a measurement device
Can include adjustments
Requires a traceable source
Valid for a limited time
Valid for temperature range
Verification
Adjustment*
Reverification
* Not Possible with all devices
Notes: To combat time, you perform an “External Calibration”
This process determines the “uncertainty” of your instrument (because your board is out of calibration)
The actual calibration process includes 3 possible steps
Verification – determine how accurate the instrument measures
Adjustment – Adjusts the measurement circuitry to bring it back within specification (not always possible, but very common to have adjustment)
Reverification – Prove that you meet the defined specifications
Calibration must be performed using an accurate, traceable source
Source must be more accurate than your instrument
Source must be traceable back to a National or known standard
The calibration is only good for
A specified amount of time (from manufacturer)
A specified temperature range
Please note – once your device has been externally calibrated, it is recommended to perform an internal calibration for your specific environment.

## Slide 19: Proof of Calibration
To meet ISO-9000 guidelines, your measurement devices must be:
Calibrated according to manufacturer’s recommendations
Calibrated with an NIST traceable source
  Meet ANSI/NCSL Z540-1
Documented with a traceable calibration certificate
Notes: Certifications (like ISO-9000), require proof of calibration for instruments used to develop or manufacture products.
This proof is a Traceable Calibration Certificate, which describes:
The environmental conditions during calibration
The external source (also called a standard) used to verify/adjust the instrument
  NIST traceable source means the source can show proof of calibration.
  ANSI/NCSL Z540-1 & ISO Guide 25 – specifies the allowed sources and requires actual measurement data as proof
When the device must be calibrated next
All instruments must provide calibration certificates
Please note – NIST is an American Standard. Other countries may require other standards.

## Slide 20: System Calibration
Referred to as “end-to-end” calibration
Reduces system measurement uncertainty
  Cables, software, sensor
Apply adjustments in software
Cannot change system components
0.002%
Notes: To recap – We’ve covered
Internal calibration – compensates for changes in temperature
External calibration – compensates for changes in time.
These are both types of component calibration
Now we’ll cover System Calibration…
System calibration considers the entire system as a single unit
Also called an “end-to-end” calibration
Quantifies all of the measurement errors in your system, and compensates for them in software
Therefore, it reduces the overall measurement uncertainty
Does not replace Component Calibration, but compliments it
You can no longer replace components without performing another system calibration
Common Questions – How often do I need to perform a system calibration (how long does it last)?
  The same factors apply as component calibration – temperature (environment) and time (component drift). It basically comes down to how important it is to you. Typical response is 6 months to a year.

## Slide 21: Performing System Calibration
Apply a traceable source
Source
Source Value
Measurement
Take measurements across your input range
Create a compensation function
Apply the compensation function to all future measurements
Y = AX2 + BX + C
Notes: How do you perform a System Calibration?
Apply a traceable source that can generate signal values across your desired input range
  The source should behave like your signal will
  Example – thermocouple calibrator
Record measurements for each known setting of the source
  Set the thermocouple calibrator to 100 ° – measure 101 °
  Set the thermocouple calibrator to 200 ° – measure 200.5 °.
  You continue for as many points as you need (typically 10)
Create a polynomial function describing this behavior
Apply this function to all future measurements
Result – Your system has a smaller measurement uncertainty

## Slide 22: Calibration Services
Calibration certificates
  Basic certificates
  Detailed certificates
Worldwide calibration services
Metrology lab support
  Manual procedures
  Calibration executive automated software
Notes: National Instruments understands the importance of calibration, and offers various calibrations services for our products
We provide calibration certificates for our instrument products
  Basic – document which shows proof it was calibrated
  Detailed – shows before an after data
You can send your National Instruments instrument to us or several 3rd party metrology laboratories from anywhere around the world
We also provide calibration procedures and calibration executive tools if your company has its own internal metrology laboratory.

## Slide 23: Measurement Ready Instrumentation
Traceable calibration certificates
Absolute accuracy information
System calibration
Notes: We’ve covered a lot of topics, but we want to make sure you walk away with three concepts:
In order to ensure your instrumentation system provides accurate measurements:
Look for traceable calibration certificates – this ensures your system has the necessary technologies to provide accurate measurements throughout it’s life
Look for absolute accuracy tables – (not relative accuracy) as this provides a guaranteed uncertainty for calibrated devices
Consider system calibration – this procedure takes time, but can greatly reduce the uncertainty in your measurements.

## Slide 24: For More Information…
Application notes
Manual calibration procedures
National Instruments calibration services
Tutorials
System calibration procedures
Third party calibration houses
ni.com/calibration
Notes: If you need more information, please visit our Web site.
We have various documents to help you, including:
Calibration procedures
Information about our calibration services
Contact information for other metrology labs that support our devices
Thank you for attending!