## Slide 1: Choosing the Best Measurement Devices for Your System Needs

## Slide 2: Outline
Hybrid ATE Systems Overview
Measurement Selection Process
Measurement Functionality Considerations
Other Considerations
Demonstration
Summary

## Slide 3: Hybrid ATE Systems Overview
A Hybrid System
  Combines components from multiple ATE platforms
  Streamlines system transition and maintenance
  Protects investment in existing software and hardware
  Allows easy integration of advances in ATE system development
Layered architecture is important
Multiple hybrid topologies
Notes: Hybrid ATE systems provide developers with a long-term ATE system that can evolve with short commercial technology lifecycles. Hybrid ATE systems take advantage of a 5 layer architecture to increase reusability, connectivity, extensibility, and module replacement.
By following the layered architecture, you create a test framework that allows you to replace and/or add components without having to re-architect the system. By clearly defining the system, you can take advantage of multiple buses using a variety of hybrid topologies. This allows you to meet multiple needs – protect the investment in existing hardware and software while being able to integrate the latest technologies.

## Slide 4: Five Layer ATE Architecture
- TestStand, DIAdem
System Management
- LabVIEW, LabWindows/CVI, C++
Application
- VISA, LV PnP, IVI, MAX, DAQmx
Measurement and Control Services
- PXI, VXI, Desktop PC, GPIB
Computing
- Instruments, Data Acquisition, Vision, Motion
Device I/O
Notes: Moving from the bottom up, we’ll start with the actual hardware devices that you use to perform testing and automation. This layer includes the various instruments and data acquisition used in your system. Developers want to choose the devices that meet measurement needs and also scale with changing system needs. Oftentimes, the devices used will require multiple instrumentation buses. The computing layer above refers to the instrumentation buses used and the method and topologies that connect these buses. Users should choose a topology that meets timing and synchronization needs, allows them to protect investment in hardware, and add new components to benefit from advancing technologies. The measurement and control services is the layer that integrate the hardware with the software. It provides configuration utilities, soft front panels, and drivers with APIs to ease programming. With tight integration with the application layer, these APIs should appear as a seamless part of the application development environment. The application layer refers to the software used to create the test programs. For instance, developers might create a program to take a measurement with a DMM and another program to generate a signal. The system management layer allows users to call these programs, manage various user levels, log and manage data, and generate reports.

## Slide 5: Issues For System Developers
Longevity
  Preserve investment in test hardware
  Streamline system maintenance
    Ease hardware replacement
    Simplify upgrade process
Integration with latest technologies
  Ability to add newer devices for new measurements
  Interoperability
Notes: Developers want to design the best test system for their UUT measurement needs. A large part of creating the best test system is designing one with longevity - with the ability to preserve the investment in hardware and software as well as streamline the processes for maintenance and upgrades. In particular, with devices, users want to ease the process of replacing hardware and simplify the steps entailed with upgrading the system. In addition, a good test system provides the user with the ability to build on top of the test system to integrate the latest technologies. This allows developers to add on newer devices that provide better performance, new measurements, or affordability. In addition, another concern is productivity. Users want to choose devices that increase productivity by providing tools for assembling the system through configuration utilities and programming, through integrated drivers. Developers want devices that minimize the changes required in programs when replacing or upgrading instruments.

## Slide 6: Measurement Device Selection Process
UUT drives Measurement Requirements
Measurement Requirements drive Instrument Selection
Instrument Selection drives ATE bus choices
Optimize Price/Performance
UUT
Instrument
ATE bus
Price/
Performance
Measurement
Requirements
Notes: When selecting the devices to use in a system, a developer’s priority should be finding the devices that best meet measurement requirements. Developers should evaluate all possibilities, determine which meet specific requirements, and then optimize the price/performance. Because an integrated software framework simplifies multiple bus integration, developers do not need to constrain the system to just one instrumentation bus. Rather, developers should use the instruments that meet all needs.
We will examine some of the various features that developers should consider when selecting an instrument.

## Slide 7: Measurement Functionality
Resolution
Frequency
Accuracy
Sampling rate
Channel count
Dynamic range
Complete measurement capability
Bus bandwidth and latency
Ease of use
Notes: When evaluating instruments based on specific measurement needs, there are many factors to consider. Not only does the instrument need to meet basic needs in terms of resolution, frequency, accuracy, sampling rate, and channel count, but the instrument should also provide complete measurement capability. For instance, with RF instruments, the instrument might provide only common spectral measurements like peak/multipeak search, power-in-band, adjacent channel power, and occupied bandwidth. However, if application needs changed and the system now needed other spectral measurements like carrier-to-noise, phase noise, and noise-floor measurements, the developer would be forced to buy a new instrument to meet needs. Choosing an instrument with complete measurement capability ensures that the ability to adjust for changing measurement needs. Another important factor to consider is ease of use – this includes documentation, support, instrument drivers, example programs. Lastly, it’s important to consider how the instrument will scale if system needs change. If more channels are required, how easy would it be to easily add channels with tight timing and synchronization or would a new instrument need to be purchased?

## Slide 8: Traditional Test Approach is Breaking Down
Power Supply
Oscilloscope
Logic Analyzer
Spectrum Analyzer
DMM
Communications Analyzer
LCR Meter
Function Generator
Pattern Generator
Programmable Switch
Images courtesy of Fluke, Rohde & Schwarz, LeCroy, Ascor, Agilent, and Tektronix
Notes: <number>
<footer>
<header>
<date/time>
Another consideration in addition to measurement functionality is the overall test system flexibility and scalability. Due to rapidly changing devices, the traditional approach to testing is beginning to break down. Adding a new vendor-defined box to test each new feature is too slow and too expensive and many times, the box to test a cutting-edge feature hasn’t even been built yet. The only way to keep up with the changes in design is to employ a software oriented approach in the test system that mirrors the software oriented model of the devices themselves. A software defined approach allows developers to maintain system scalability and flexibility. The software defined approach takes advantage of hardware buses with tight timing and synchronization and high bandwidth and low latency to enable many complex measurements.

## Slide 9: System Scalability
Adding channels to the system
  Ability to add channels and maintain timing and synchronization in the system
Adding measurements to the system
  Ability to easily incorporate new measurement functionality into the system
  Ability to easily coordinate timing and triggering for different measurement components
Notes: An important consideration is system scalability. Is it designed to be modular so that adding channels would be easy? Can users still maintain timing and triggering specifications when adding channels? Developers should also plan for the measurement needs of the system changing – for instance, the system might require additional analysis. Does the device provide full measurement functionality and analysis or if additional measurements are needed, would the user be required to purchase new devices?

## Slide 10: Timing and Synchronization
Handling synchronous events
  For example, starting a signal generator and digitizer at same time
Performing asynchronous events
  For example, handshaking with switch and DMM, Arb, and so on
Clocks
  Shared
  Phase-lock looping
Triggers
  Propagation delay
  Skew
Notes: Often, the test system requires timing and triggering of the instruments in the system. Developers should make sure that they are able to synchronize various instruments together down to the level expected. Timing and synchronization methods is usually implemented through the instrumentation bus and are important for handling synchronous events like starting two instruments at the same time as well as asynchronous events like handshaking with a switch, DMM, and arbitrary waveform generator.
Developers should make sure that there is a way to correlate clocks between multiple instruments. This could be done either by sharing clocks or by phase-lock looping to a reference clock. When evaluating instruments, users should also take notice of the level of synchronization available with triggers. Factors like propagation delay and the skew of when the triggers are received are important in determining the level of synchronization that is achievable.

## Slide 11: Bus Capabilities
 | GPIB | VXI | USB | TCP/IP Ethernet | Standard PCs | CompactPCI/ PXI
Transfer Width (bits) | 8 | 8, 16, 32, 64 | Serial | Serial | 8, 16 (ISA) 8, 16, 32, 64 (PCI) 8, 16, 32, 64 (PCI-X 1.0 and 2.0) | 8, 16, 32, 64
Throughput (Mbytes/s) | Up to 8 | Up to 160 | Up to 1.5 (USB 1.0) Up to 60 (USB 2.0) | Up to 1.25 (10BaseT) Up to 12.5 (100BaseT) Up to 125 (1000BaseT) | 1-2 (ISA) 132-512 (PCI) 264-1024 (PCI-X 1.0) 1064-4264 (PCI-X 2.0) | Up to 528
Timing and Synchronization | None | Defined 10 MHz Reference Clock 8 TTL Trigger Lines 2 ECL Trigger Lines Local Bus (sub-microsecond) | None | IEEE-1588 Synchronization Protocol (microseconds) | Proprietary | None for CompactPCI Defined for PXI 10 MHz Reference Clock 8 TTL Trigger Lines STAR Trigger Local Bus (nanosecond)
Control Loop Rates | Seconds | Microseconds | Seconds | Seconds | No | No (CompactPCI) Microseconds (PXI)
Standard Software Frameworks | VISA available | VXIplug&play Defined | VISA available | VISA available | None | None (CompactPCI) Defined (PXI)
Modular | No | Yes | No | No | No | Yes
EMI Shielding | Optional | Defined | Optional | Optional | Board Specific | Module Specific
Notes: System flexibility and scalability as well as the timing and synchronization is influenced by the capabilities of the instrumentation bus. This chart compares some of the many instrumentation buses available. When selecting devices, determine if the bus will be able to adequately meet application needs. Because there are many hybrid topologies available, users should not constrain the system to just one instrumentation bus.

## Slide 12: Connectivity to PC
Native buses: PCI, USB, Serial
  Plugged into PC to use Windows native drivers
Register-based instruments: PXI and VXI
  Plugged into PC or connected with MXI for low latency access
Message-based, non-native instruments: GPIB, Ethernet
  Plugged into instrument bus controller
Notes: Depending on the device type, some PC connections are more appropriate than others. Thus, developers should also consider the way the instruments connect to the PC. Instruments that use buses native to the PC like PCI, USB, and serial should be directly plugged into the PC and use native Windows drivers. If the instrument is register-based it should be based on PXI or VXI and plugged into the PC (embedded controller) or use a MXI remote controller for low latency. If the instrument is message-based, like those common to GPIB and Ethernet, they should connect to the PC though an instrument bus controller.

## Slide 13: Mix and Volume
High mix environment
  Applicability of instruments for multiple UUTs
  Ability to design specific measurement system from instruments
High volume environment
  Instrument throughput
  Bus bandwidth
  Bus latency
Notes: For large test systems, for example those used in a manufacturing plant, users should consider the mix and volume of the items tested. If it is high mix, users should select devices that provide a breadth of measurements so that it can be used to test multiple UUTs. A modular test system is advantageous because it allows users to mix and match needed components to test various products. If it is a high volume environment, developers will want to select instruments with an instrumentation bus that can provide high throughput. Users should consider if the bus has the capability to support high bandwidth needs, especially for applications with waveform generation or acquisition as well as RF measurements. Users should also look for buses with low latency since this is a measure of the delay of transmission of data which directly affects applications like instrument configuration, switching, and digital multimeter measurements. With high volume, it is important to minimize the time it takes to execute each test.

## Slide 14: Mass Interconnect
Clean, reliable method of connecting to multiple UUTs
Designed for multiple connect/disconnect operations
Reusability
Vertical Hinge Mounting Frame
Test Equipment
Receiver
Cable Assemblies
ITA Frame
ITA Modules
ITA Patchcords
ITA w/ Enclosure
Notes: Another consideration is the availability of mass interconnect for the instruments in the system. Mass interconnect offers a clean, reliable method of connecting to multiple devices under test. This is especially valuable in ATE systems when users may have several different load boxes that need to be quickly interchanged. Mass interconnect is designed for multiple connect/disconnect operations – usually on the order of tens of thousands while standard connectors are designed for far fewer connects. Factors like whether mass interconnect is available and if the developer is able to reuse mass interconnect from another test system should be factors in the instrument selection process.

## Slide 15: Product Lifecycle
Unit Under Test (UUT) lifecycle in comparison to test component lifecycle
Shorter UUT lifecycle
  Instruments with a range of functionality for testing other UUTs
Longer UUT lifecycle
  Instruments with longevity and form-fit-function replacements
Notes: When selecting devices, developers should also consider how the UUT lifecycle compares with the test components (instruments) lifecycle. If the UUT has a short lifecycle in comparison, developers will want to choose devices that provide a range of functionality that could be used for other test systems. If the UUT has the longer lifecycle, users should choose devices that offer longevity so that when upgrades are needed, the devices are still available.

## Slide 16: Planning for Instrument Replacement
Provided via architectures and standards
  Plug&Play drivers based on VISA provide bus and controller interchangeability
  IVI drivers provide instrument interchangeability
  Test management software
Use new ATE buses
Enables upgrade of ATE components
Notes: In those cases where the UUT has a longer lifecycle than the test components, developers should plan for replacing and upgrading instruments as part of maintenance of the system. Users can ease the process of replacement by taking advantage of instrument drivers, both plug&play and IVI for their system. LabVIEW and CVI/LabWindows Plug&Play instrument drivers are based on VISA to provide bus and controller interchangeability. For instance, if using a VXI embedded controller for the VXI portion of the system, the developer could replace the embedded controller with a remote controller without having to change the software. IVI drivers allow engineers to interchange instruments in a system without modifying test software for specified classes of instruments such as oscilloscopes or switches. An IVI instrument driver that conforms to one of these classes may be substituted with another instrument of the same class, regardless of manufacturer or bus connection.
Developers can also plan for replacement by creating modular test code and using the test management software to sequence these individual test programs. This way, replacing an instrument has minimal impact on the software portion of the test system.
By taking advantage of these architectures and standards, developers can introduce new ATE buses when replacing instruments and upgrade the components of the test system.

## Slide 17: Multi-platform Systems
Let measurement needs drive bus
Various ATE bus options
  GPIB
  VXI
Many system topologies and connectivity options
Advantages of a multi-platform system
  Extensibility
  Longevity
  Integration with latest technologies
  PCI
  PXI
Notes: With all of these various measurement needs, developers should select the instruments best suited for the system. As a result, it is likely that multiple buses will be required. With the many ATE bus options available, users should not limit the system to the instruments available on just one bus. Many connectivity options are available so that developers can design a system with a topology to meet specific system needs.
If any instruments need to be added or replaced, a multi-platform system is easily extensible and allows users to mix and match ATE buses. Multi-platform systems also meet the need for longevity by providing a system that protects investment in hardware and software and integrates the latest technologies.

## Slide 18: Demo: Wireless Communications Test System
3D Power Spectrum, Occupied Bandwidth, Auto Find (LabVIEW)
TestStand
MAX, VXI ResMan, VISA, IVI and PnP Drivers
PCI (NI-8350), PXI MXI-4, VXI-USB, GPIB
2.7 GHz PXI RF Signal Analyzer, 8.5 GHz VXI RF Signal Analyzer, PXI DMM, PXI Switch, GPIB Power Supply
Notes: For a wireless communications system, you can see how the layers work together to integrate the buses for this test application. Through the architecture the user can integrate VXI and GPIB instruments with some newer PXI instruments. This allows the user to take advantage of existing hardware and software and still incorporate newer technologies for the price/performance benefits.
Here the user is able to take advantage of a 8.5 GHz RF signal analyzer for higher frequency needs and combine it in the same system with a PXI RF signal analyzer, switch, and DMM which offer a lower price/performance point. Here the user can continue to use some existing equipment like the GPIB power supply with newer instrumentation. This system also leverages instrumentation buses of PXI and VXI for their latency and bandwidth capabilities.

## Slide 19: Demo: Wireless Communications Test System
Test Head
Test Receiver
(UUT)
2.4 GHz
900 MHz
NI-8350 (PCI)
GPIB Power Supply
PXI 4060
PXI MXI-4
VXI-USB
VXI RF Signal Analyzer
PXI 2592
PXI 4551
PXI 4552
MXI-4
USB
Power for UUT
Notes: The block diagram shows how instruments of various buses can be interconnected to test this wireless communications test system.

## Slide 20: Hybrid ATE Summary
Use the five layer architecture
  Choose best instruments
  Factor bus capabilities in instrument selection process
  Let instrument selection drive ATE buses used
  Consider system scalability and flexibility
Streamlines system transition and maintenance
Maximizes return on investment in existing software and hardware
PASS
FAIL
UUT
UUT
UUT
UUT
Notes: In summary, hybrid ATE systems allow developers to create test systems that streamline system maintenance and transition and allow users to maximize investments in existing hardware and software. By using a five layer architecture, system developers can choose the best instruments for their system. Various factors like measurement functionality, as well as instrumentation bus capabilities, system scalability and flexibility, mass interconnect, mix, and volume should be considered. Developers do not need to constrain the system to one bus, but can instead create a system that can handle interconnecting multiple buses as well as extend and grow with needs. It is important to understand the various interfaces to interconnect buses and to carefully select components within the system..

## Slide 21: Questions?
Notes: <number>
<footer>
<header>
<date/time>
© National Instruments Corporation Acquire, Analyze, and Present with LabVIEW Seminar