## Slide 1: Seven Habits of Highly Effective LabVIEW™ DAQ Programmers
Reid Lee
Staff Software Engineer
Wed Aug 16
10:15-11:30 a.m., 12:00-1:15 p.m., 3:30-4:45 p.m.
Exhibit (3D)

## Slide 2: Agenda
Expectations
Seven key habits and exercises
Questions

## Slide 3: Expectations
  You should:
  be familiar with LabVIEW and data acquisition terminology
  have used LabVIEW to build a VI

## Slide 4: Habit #1
Use Measurement & Automation Explorer (MAX)
  Verify hardware
  Test panels
  Virtual channels
    Custom channel names
    Scaling
  I/O control in LabVIEW 6i
  Product updates wizard
Notes: The test panels in MAX allow you to quickly check your signals without any programming.
Virtual channels simplify channel strings and scaling from voltages to engineering units.
No need to use cryptic channel strings like ob0!sc1!md4!3. Instead, name the channel something descriptive like “Left front brake temperature”. Also, instead of manually converting a thermocouple’s reading in the mV to degrees, let MAX do it for you.
LabVIEW 6i’s new I/O control allows you to easily select and edit virtual channels.
MAX can contact ni.com to determine whether you have the latest software versions.

## Slide 5: Habit #2
Research and plan
  Start with examples
    Solution Wizard
    Helps verify signal connections
    Lays a solid foundation
  Think architecture
    Break application into logical tasks
    Consider a state machine approach
    Cookie cutting common constructs – Merge VI
Exercise 1
Notes: Starting with an example helps you to understand what VIs to call and in what order they should be called. Examples provide a solid foundation on which to base your application.
No matter how intuitive or easy LabVIEW is, it is still programming. Good programming principles are still important and will lead to code that is easier to debug, easier to add features to, and easier to understand.
A state machine can be defined in many different ways. For the purpose of this habit, we’ll define a state machine as a software model that divides an application into logical states such as initialize board, acquire waveform, set output voltage, save data to disk, and most likely, a “wait for user input” state. The transition from state to state depends on several factors including, but not limited to, success/failure of task, user input, time, or data acquired.
State machines in LabVIEW involve a case structure in a while loop. The case structure contains cases for each state. To move from state to state, you pass the next state of your application through a shift register wired back into the case structure. See NIWeek example State Machine.llb.
“Cookie cutting common constructs” refers to creating Merge VIs on your Functions palette that are often used. How many times have you created data acquisition program with “config/start/read/clear”? You will add a state machine template to your Functions palette in exercise 1 (see handout).

## Slide 6: Habit #3
Understand the DAQ VI palette
  Easy I/O, intermediate, & advanced VIs
  Configuration VI placement
Exercise 2
Notes: Easy I/O – Simple, one-shot acquisition VIs. These VIs can be called in a loop but you will take a performance hit. A typical acquisition consists of a “config, start, read, and clear”. That’s what the Easy I/O VIs call internally. Calling an Easy I/O VI in a loop causes unnecessary calls to AI Config, AI Start, and AI Clear for each iteration.
Intermediate DAQ VIs – The majority of DAQ users call into the intermediate-level VIs. These VIs offer flexibility while maintaining ease-of-use. The intermediate-level VIs consist of AI Config, AI Start, AI Read, AI Single Scan, and AI Clear. For typical buffered, continuous acquisitions, you call AI Config, AI Start, then AI Read in a while loop, followed by AI Clear. Although the intermediate VIs expose advanced features such as clock and trigger configuration, the default values for those inputs will satisfy most users.
Advanced DAQ VIs – Most intermediate DAQ VIs internally call the lowest level advanced DAQ VIs. The advanced DAQ VIs can be called when features are not exposed by the intermediate VIs. For example, you might use AI Buffered Read instead of AI Read in order to access the read specification cluster. You can also call advanced VIs to query for low level settings. For instance, if you need to know what gain, range, or polarity was set up by NI-DAQ, you can follow your call to AI Config with AI Hardware Config and interrogate the group channel settings cluster.

## Slide 7: Habit #4
Optimize analog input
  Intricacies of AI read
    Checking acquisition status
    Preventing ‘blocking’
Exercise 3
Notes: AI Read.vi can be used in many ways, some of which are outlined here. Depending on your application, altering the default behavior of AI Read can improve the performance of your VI.
Checking the acquisition status – You want to be able to check how many points your DAQ board has acquired before trying to read the data. If you are reading 1,000 point blocks from the board, you wouldn’t want to read the buffer when only 100 points were available. If you do, AI Read will block (data flow will cease) until the other 900 points are acquired. To check how many points are available, pass 0 for the number of scans to read and check the scan backlog output. Using the scan backlog, you can write your VI so that data is read from the buffer only when all of the data is present.

## Slide 8: Habit #4 (cont.)
Optimize analog input
  Intricacies of AI read
    Using read pointer/offset
Exercise 4
1
2
3
4
5
6
7
8
9
7 points acquired, 2 points read
Notes: AI Read’s read/search position input
  When AI Read reads N points, the default behavior is to read where it last left off. This is useful if you need to see every acquired point. Errors can occur if your acquisition rate is too fast (or your acquisition buffer is too small) resulting in a -10846 buffer overrun error.
  If every point is not important to you or if you want to maximize the DAQ board’s sample rate, you can have NI-DAQ read the latest acquired points rather than reading where it left off. To do so, specify relative to end of data for the read/search position. In the above hypothetical example, you would want to read relative to end of data to avoid a buffer overrun.
  This introduces a new problem. If you specify relative to end of data and you read two points out at a time, AI Read would block until points 8 and 9 are acquired. You can use a read offset of -2 to have AI Read immediately return points 6 and 7.

## Slide 9: Habit #5
Understand DAQ occurrences
  Applications
  Benefits
    background processing (eliminates polling)
    driver does the work
    more efficient
Exercise 5
Notes: Use the DAQ Occurrence Config VI to set up the occurrence conditions such as:
notify when the acquisition is complete
notify when N scans are acquired
notify every time N scans are acquired
Others include digital pattern conditions.
Wire into Wait On Occurrence function.
See Acquire N Scans – Async Occurrence.vi that ships with LabVIEW.

## Slide 10: Habit #6
Understand data efficiency
  Waveform data type (LabVIEW 6i)
  Binary vs. scaled data
  Channel to Index and AI Buffer Read VIs
  Conditional retrieval
Notes: Reasons to use waveform data type:
Everything on one wire (simpler diagrams!)
Built-in timing information (start time, and delta T), channel names, errors
Analysis VIs and charts/graphs can use timing information
Binary vs. scaled data – acquiring binary data is faster and more efficient for saving to disk.
Channel to index VI – use this VI along with AI Buffered Read to read selective channel data from the driver into LabVIEW memory. This is especially useful in Remote Device Access (RDA) applications. If you were acquiring 100 channels of data on your remote machine but you only needed to see five channels of data, you wouldn’t want to bog down your network with unnecessary traffic from the 95 other channels. By using the Channel to Index VI, you can select which channels of data actually get transferred over the network (even though the remote machine is still acquiring all the channels). In networked applications, this can affect throughput tremendously. See NIWeek example channel to index demo.vi.
Conditional Retrieval – use this cluster on AI Read to specify acquisition criteria. For boards that do not have HW analog triggers, you can use conditional retrieval to tell NI-DAQ “give me data once the signal crosses .7 volts with a rising edge slope”. See LabVIEW example Acquire N Scans Analog Software Trig.vi.

## Slide 11: Habit #7
Use Profiler
  LabVIEW’s Profiler helps you:
    Detect memory leaks
    Find what VIs your application spends the most time in
Exercise 6

## Slide 12: Related NIWeek Sessions
SP1A – Check Out The Latest LabVIEW
  All day today, Exhibit 3A
AT3B – LabVIEW Programming Techniques
  Friday, 10:15 AM and 1:45 PM, Ash (10A)
Making Measurements
  See page 7 of your NIWeek program
    Temperature : WS1A & SP1B
    Frequency : WS1B
    Strain and Displacement : WS1D
    Vibration : WS2D

## Slide 13: Closing
Questions?
Solutions for exercises
Visit Developer Zone™ – zone.ni.com
  Resource library
  Examples
  Technical presentations
  User forums