## Slide 1: Arbitrary Waveform Generators
Jerry Lopato
Applications Engineer
Fri Aug 18
10:15-11:30 a.m., 12:00-1:15 p.m., 1:45-3:00 p.m., 3:30-4:45 p.m.
Exhibit (3E)

## Slide 2: Agenda
Arb waveform generation
Arb multiple waveform generation
Creating AM/FM waveforms
DDS mode and frequency sweeping
Stimulus response example

## Slide 3: Inputs/Outputs
Arb out
Sync out
PLL ref in
50-pin digital connector
16 bit data (3.3 V/5 V), clock, external trigger, marker, +5 V/3.3 V, 5 GPIO (5 V only)
Notes: The arb/video out connector provides the waveform output. The maximum output levels on this connector depend on the type of load termination. If the output of your NI 5411/5431 terminates into a 50 Ωload, the output levels are ±5 V, as shown in Figure 1-2. If the output of your NI 5411/5431 terminates into a high-impedance load (HiZ), the output levels are ±10 V. If the output terminates into any other load, the levels are:
RL
V out = --------------- X 10 V
RL + Ro
V out is the maximum output voltage level
RL is the load impedance in ohms
Ro is the output impedance on the NI 5411/5431
By default, Ro = 50 , but you can use your software to set it to 75 Ω
The SYNC connector provides a TTL version of the sine waveform being generated at the output. You can think of the SYNC output as a very high-frequency resolution, software-programmable clock source for many applications. You can also vary the duty cycle of the SYNC output on the fly by software control.

## Slide 4
ni.com
PLL Ref Connector
Notes: The PLL ref connector is a phase-locked loop (PLL) input connector that can accept a reference clock from an external source and phase lock the NI 5411 internal clock to this external clock. The reference clock should not deviate more than ±100 ppm from its nominal frequency. The minimum amplitude levels of 1 V pk-pk are required on this clock. You can lock reference clock frequencies of 1 MHz and 5 to 20 MHz in 1 MHz steps.
Dig out is a 16-bit digital I/O connector that contains the 16-bit digital pattern outputs, digital pattern clock output, marker output, external trigger input, and +5 V power output.

## Slide 5: Hardware Features
Maximum update rate of 40 MHz
Waveform memory
  DRAM option – 2 MegSamples (standard)
  DRAM option – 8 MegSamples (optional)
12-bit high-speed DAC
  1 sample = 16 bits (lower 4 bits ignored)
80 MHz interpolating digital filter
Waveform looping and linking
Notes: On the NI 5411/5431, the high-speed DAC is always updated at 80 MHz, but the maximum update clock for waveform memory is 40 MHz. The update clock for the waveform memory can be further divided by a 16-bit counter, as shown in Figure 2-2. Therefore, the slowest update rate is 40 MHz divided by 65,536, which is 610.35 Hz.
The NI 5411/5431 uses a waveform memory that is 16 bits wide. The standard memory size for the NI 5411 is 2,000,000 samples, and for the NI 5431 is 8,000,000 samples. With a minimum standard memory size of 2,000,000 samples, you can store very long waveforms on the board and obtain reliable waveform generation even at full speed. You can upgrade the NI 5411 to an 8,000,000-sample waveform memory by installing the 16 MB memory module.
The NI 5411/5431 can correct for slight deviations in the flatness of the frequency characteristic of the analog lowpass filter in its passband. The response of the filter is stored in an onboard EEPROM in 1 MHz increments up to 16 MHz for NI 5411 and up to 8 MHz for NI 5431. If you want to generate a sine wave at a particular frequency with filter correction applied, you have to specify that frequency through your software.

## Slide 6: Hardware Features (cont.)
Minimum buffer length is 256 samples
Buffer should be integer multiples of 8 samples
Maximum buffers in memory – 5,000
Maximum loops on same buffer – 65,535
Number of links – 256
– Infinite in later versions
Notes: • Buffer number – Specifies the buffer number to be generated.
• Buffer size – Specifies the total count of the buffer to be generated. This count may not be the actual size of that buffer. If the count is less than the actual size of the buffer, only a part of that buffer is used for that stage. If the count is more than the actual size of that buffer, part of the next sequential buffer is also used. If the buffer size is set to zero, the software automatically uses the true size of that buffer.
• Buffer loops – Specifies the number of times that buffer has to be looped. The maximum number of loops possible is 65,535.
• Marker offset – Specifies where the marker must be generated within that buffer. For more information on markers, see the Marker Output Signal section later in this chapter.

## Slide 7: Hardware Features (cont.)
Direct digital synthesis (DDS) for repetitive standard waveforms
  Sine, TTL, triangular, and so on
  Phase continuous switching
  High frequency resolution
  ~10 MHz to 16 MHz in 10 MHz steps.
  Dedicated lookup memory of 16K points
  Instruction buffer for multiple frequencies
Notes: Direct digital synthesis (DDS) is a technique for deriving, under digital control, an analog frequency source from a single reference clock frequency. This technique produces high-frequency accuracy and resolution, temperature stability, wideband tuning, and rapid and phase-continuous frequency switching. You should use DDS mode for standard function generation rather than for arbitrary waveform generation. The NI 5411/5431 uses a 32-bit, high-speed accumulator with a lookup memory and a 12-bit DAC for DDS-based waveform generation.
The lookup memory is dedicated to the DDS mode only and cannot be used in arb mode. You can store one cycle of a repetitive waveform – a sine, triangular, square, or arbitrary wave – in the lookup memory. Then, you can change the frequency of that waveform by sending just one instruction. You can use DDS mode for fine frequency resolution function generation. You can generate sine waves of up to 16 MHz for NI 5411 and 8 MHz for NI 5431 with a frequency resolution of 10.0 mHz. Because this mode uses an accumulator, waveform generation loops back to the beginning of the lookup memory after passing through the end of the lookup memory.

## Slide 8: Hardware Features (cont.)
Output attenuation from 0 dB to 73 dB in 0.001 dB steps
  1dB attenuation = (-) 20*log(V desired/ V in)
V in = 5 V, V desired is in volts
Pattern generation 16-bits data and clock
Selectable output impedance
  50 W or 75 W
Software adjustable variable duty cycle on SYNC output
Notes: The output attenuators, made of resistor networks, can be switched in any combination. The maximum attenuation possible on the NI 5411/5431 is 73 dB.
By attenuating the output signal, you keep the dynamic range of the DAC. You do not lose any bits from the digital representation of the signal because the attenuation is done after the DAC and not before it.
attenuation (in decibels) = -20 log 10 (Vo /Vi )
Vo = desired voltage level for the output signal
Vi = input voltage level
The NI 5411/5431 has 16-bit digital pattern generation outputs at the digital connector. This digital data is first synchronized to the sample clock then buffered and sent to the connector through a 68 series resistor. The sample clock is also buffered and sent to the digital connector to latch the data externally. Figure 2-23 shows the data path for digital pattern generation. Since the digital pattern data does not go through the digital filter, it is available directly from the memory. This means that there is a fixed delay of 26 sample clocks between the analog waveform (which lags behind the digital waveform) and the digital patterns. Although the digital filter can be disabled in software, there will still be a 26 sample clock delay.
You can enable or disable digital pattern generation through software. All linking and looping capabilities also are available for digital pattern generation. If you select DDS mode, the DDS data appears at the digital I/O connector. You can use digital pattern generation to test digital devices such as serial and parallel DACs and to emulate protocols.
Before the signal reaches the output connector, you can select an output impedance of 50 Ωor 75 Ω. If the load impedance is 50 Ωand all the attenuators are off (an output attenuation of 0 dB), the output levels are ±5 V.
Most applications use a load impedance of 50 Ω, but applications such as video device testers require 75 Ω If the load is a very high-input impedance load (~1 Ω), you will see output levels up to ±10 V.

## Slide 9: Hardware Features (cont.)
Internal 40 MHz updates
Internal 16-bit clock divider
PLL locking to external references – 1 MHz, 5 MHz thru 20 MHz
Multi-board synchronization using RTSI Bus using PLL, RTSI clock, and RTSI trigger lines
Internal reference also available

## Slide 10: Hardware Features (cont.)
Software
  NI-5411 supported in NIFGEN 1.0.1
    NI-DAQ™ required for installation
  LabVIEW 6i, C API, and Visual Basic support
  VirtualBench™ – arb and VirtualBench – function generator
    Free

## Slide 11: Hardware Features (cont.)
Triggering
  External (I/O connector)
  Over the RTSI Line
  Software generated (internal)
Triggering Modes
  Single, continuous, stepped, burst
Notes: Trigger sources are software selectable. By default, the software produces the trigger sources. You can also use an external trigger from a pin on the digital I/O connector, the RTSI trigger lines on the RTSI bus, or the TTL trigger lines on the PXI trigger bus on the backplane.
If you need to automatically trigger the waveform generation, use software to generate the triggers. For external triggering, you need a rising TTL edge. For external triggering, apply a rising edge TTL signal to the EXT_TRIG input. This signal should remain deasserted (logic low) until after the waveform generation has been initiated in software.
For more information on triggering using RTSI lines, see the RTSI/PXI trigger lines section of this presentation.
The NI 5411/5431 has four triggering modes – single, continuous, stepped, and burst. These trigger modes are available for both arb and DDS modes.
Single trigger mode – The waveform you define in the staging list is generated only once by going through the entire staging list. Only one trigger is required to start the waveform generation.
Continuous trigger mode – The waveform you define in the staging list is generated infinitely by continually cycling through the staging list. After a trigger is received, the waveform generation starts at the first stage, continues through the last stage, and loops back to the start of the first stage, continuing until you stop the waveform generation. Only one trigger is required to start the waveform generation.
Stepped trigger mode – After a start trigger is received, the waveform defined by the first stage is generated. Then, the device waits for the next trigger signal. On the next trigger, the waveform described by the second stage is generated, and so on. Once the staging list is exhausted, the waveform generation returns to the first stage and continues in a cyclic fashion.
Burst trigger mode – After a start trigger is received, the waveform defined by the first stage buffer is generated until another trigger is received. At the next trigger, the buffer of the previous stage is completed before the waveform defined by the second stage buffer is generated. Once the staging list is exhausted, the waveform generation returns to the first stage and continues in a cyclic fashion.

## Slide 12: Three Products To Choose From
NI-5401
  One channel 12-bit function generator
NI-5411
  One channel arbitrary waveform generator and function generator
NI-5431
  One channel arbitrary waveform generator specifically for video signals

## Slide 13: NI 5401
One channel 12-bit function generator (AFG)
Waveforms generated
  Sine 16 MHz
  SYNC(TTL) 16 MHz
  Triangle 1 MHz
  Ramp 1 MHz
  Ramp 1 MHz

## Slide 14: NI 5411
Update rate of 40 MS/s
On board memory options of 4 or 16 Mbytes.
Features
Linking and looping
Four trigger modes, digital pattern output, markers, filters, and TTL output

## Slide 15: Arbitrary Waveforms
Single arbitrary waveform
  Waveform array issues
    Minimum buffer length is 256 samples
    Buffer should be integer multiples of 8 samples
    Minimum and maximum values -1.0 to +1.0
    Continuous should start and end at same value
Notes: The NI 5411 device memory architecture imposes certain restrictions on the buffer size and resolution. The minimum buffer size for arb mode is 256 samples, and the buffers must be in multiples of eight samples. If these buffer requirements are not met, NI-FGEN returns an error. For example, if you request the NI 5411/5431 to load a buffer of 257 samples, NI-FGEN will return an error.
The values in the waveform array should be no larger than 1.0, and no smaller than -1.0

## Slide 16: Arbitrary Waveforms
Single arbitrary waveform
  Hardware issues
    Maximum Update Rate 40 MSamples/sec
    Update Rate 40 MSamples/sec/n
      Where n = 1 to 65,535
    Memory 2 million samples default
    8 million samples optional
Notes: On the NI 5411/5431, the high-speed DAC is always updated at 80 MHz, but the maximum update clock for waveform memory is 40 MHz. You can further divide the update clock for the waveform memory by a 16-bit counter, as shown in Figure 2-2. Therefore, the slowest update rate is 40 MHz divided by 65,536, which is 610.35 Hz.
The NI 5411/5431 uses a waveform memory that is 16 bits wide. The standard memory size for the NI 5411 is 2,000,000 samples, and for the NI 5431 is 8,000,000 samples. With a minimum standard memory size of 2,000,000 samples, you can store long waveforms on the board and obtain reliable waveform generation even at full speed. You can upgrade the NI 5411 to an 8,000,000-sample waveform memory by installing the 16 MB memory module.

## Slide 17: Arbitrary Waveforms
Example 1 Arbitrary Waveform Generation.vi

## Slide 18: Multiple Arbitrary Waveforms
Looping and linking
1. Define segments
2. Loop segments to create stages
3. Link stages to create waveform – create the staging list

## Slide 19: Multiple Arbitrary Waveforms
Notes: Buffer loops – Specifies the number of times that buffer has to be looped. The maximum number of loops possible is 65,535.

## Slide 20: Multiple Arbitrary Waveforms
Two sets of memory on the instruments
  Waveform memory
  Instruction first in, first out (FIFO) memory
1. Load segment into the waveform memory
2. Load linking and looping instructions into the instruction FIFO memory
3. The instrument looks up the instructions, puts together the waveform from the segments, and outputs the signal

## Slide 21: Multiple Arbitrary Waveforms
Up to 5,000 different waveform segments
  Minimum size of segment has to be 256 (arb mode)
  Segment size has to be a multiple of eight.
Each segment can be looped up to 65,535 times to form a waveform stage.
Up to 292 stages to define the output or waveform (arb mode).

## Slide 22: Linking and Looping
www.ni.com
Waveform Points Form
A Wavefrom Buffer(1)
Loops of Waveform Buffer(1)
Form A Stage
Waveform Buffer(2)
Waveform Sequence (or Staging List)
Stage3
Stage4
Stage5
Stage6
Stage2
Stage1
Waveform Buffer(3)
Buffer loop count = 3
Buffer loop count = 2
Buffer loop count = 4
Buffer loop
count = 1
Notes: Before you begin generating waveforms, you must load the buffers on your NI 5411/5431. Load each signal to generate into the memory in the form of 16-bit digital samples. A finite number of these samples makes a waveform buffer, also called a waveform segment. You can load multiple buffers in the memory of the NI 5411/5431. To generate these buffers, you prepare a staging list, or a sequence list, which contains a sequence of stages. Each stage specifies the buffer, its number of loops, and its marker offset.

## Slide 23: Multiple Arbitrary Waveforms
Example 2 – Arbitrary Waveform Staging.vi

## Slide 24: Modulation
Complex waveform
  Multiple frequency components
    Carrier frequency M
    Signal/frequency representing information
  Signal frequency modulates carrier frequency

## Slide 25: Modulated Waveforms
AM – amplitude modulation

## Slide 26: Modulated Waveforms
FM – frequency modulation

## Slide 27: Modulated Waveforms
Example 3 – AM/FM modulation waveform.vi

## Slide 28: DDS
Direct digital synthesis (DDS) characteristics
  Generate standard waveforms
  16 MHz maximum sine wave generation

## Slide 29: DDS
Digital technique to create an analog frequency source from a single reference clock frequency
The instrument downloads a single cycle of a standard waveform into a lookup 16 k memory
    16,384 samples represent a cycle
Notes: Direct digital synthesis (DDS) is a technique for deriving, under digital control, an analog frequency source from a single reference clock frequency. This technique produces high-frequency accuracy and resolution, temperature stability, wideband tuning, and rapid and phase-continuous frequency switching. You should use DDS mode for standard function generation rather than for arbitrary waveform generation.
The NI 5411 uses a 32-bit, high-speed accumulator with a lookup memory and a 12-bit DAC for DDS-based waveform generation.
The lookup memory is dedicated to the DDS mode only and cannot be used in arb mode. You can store one cycle of a repetitive waveform – a sine, triangular, square, or arbitrary wave – in the lookup memory. Then, you can change the frequency of that waveform by sending just one instruction. You can use DDS mode for very fine frequency resolution function generation. You can generate sine waves of up to 16 MHz for NI 5411 with a frequency resolution of 10.0 MHz. Because this mode uses an accumulator, waveform generation loops back to the beginning of the lookup memory after passing through the end of the lookup memory.

## Slide 30: DDS
Only one standard waveform at a time
Controls frequency and duration
Phase continuous frequency switching
 Frequency sweep and hop applications
f1, t1 f2, t2 f3, t3 f4, t4
Stage 1
Stage 2
Stage 3
Stage 4
Notes: You can define a staging list in DDS mode for performing frequency hops and sweeps. The entire staging list uses the same buffer loaded into the lookup memory. All stages differ in the frequency to be generated.
Note – The minimum time that a frequency should be generated is 2 µs. Therefore, the maximum hop rate from frequency to frequency is 500 kHz.
The maximum number of stages the instruction FIFO for DDS mode can store is 512. For more information on the waveform generation process, refer to your software documentation.

## Slide 31: DDS
Example 4 – Frequency Sweeping and Hopping.vi

## Slide 32: Stimulus Response (Bode Plot)
Frequency response of a system
  Magnitude
  Phase
Stimulus
  Sinusoidal source
  Covering a range of frequencies
Notes: Bode plots used to illustrate the frequency response of a system to a sinusoidal excitation. Bode diagrams plot the magnitude and phase of system response against the frequency (generally on a log scale) at which a system is excited.

## Slide 33: Stimulus Response (Bode Plot)

## Slide 34: Stimulus Response (Bode Plot)
Example 5 – Stimulus Response.vi