## Slide 1: Hybrid VXI/PXI System Architecture
November, 2002

## Slide 2: Agenda
Overview
Basic VXI/PXI System Integration
  PC-based control
  VXI-based control
  PXI-based control
  Distributed control

## Slide 3: Why Build Hybrid VXI/PXI Systems
Protect investments in existing systems
Aid transition to modern architectures
Optimize system resources and minimize costs

## Slide 4: Test Platforms
 | GPIB | VXI | Standard PCs | PXI / CompactPCI
Transfer Width (bits) | 8 | 8, 16, 32 | 8, 16 (ISA) 8, 16, 32, 64 (PCI) | 8, 16, 32, 64
Throughput (Mbytes/s) | 1 8 (HS488) | 40 80 (VME64) | 1-2 (ISA) 132-264 (PCI) | 132-264
Timing and Synchronization | None | Defined | Proprietary | Defined
Product Availability | >10,000 | >1000 | >10,000 | ~1000
Form Factor | Large | Medium | Small-Medium | Small-Medium
Standard Software Frameworks | None | VXIplug&play Defined | None | Defined
Modular | No | Yes | No | Yes
EMI Shielding | Optional | Defined | Board Specific | Module Specific
System Cost | High | Medium-High | Low | Low-Medium

## Slide 5: The Tools : Overview of Controllers

## Slide 6: Transparent Remote Control Interfaces
PXI
VXI

## Slide 7: Ethernet Remote Control
Typically requires an Embedded Controller combined with software
Many implementation choices - depends on HW & SW used
  VXI common examples
    VISA Server (NI-VISA 2.6 or later) : transparent VXI control over ethernet
    Client-Server systems (Datasockets, TCP/IP, LabVIEW Remote Front Panels, VI Server)
  PXI common examples
    LabVIEW Real-Time
    Client-Server systems (Datasockets, TCP/IP, LabVIEW Remote Front Panels, VI Server)

## Slide 8: PC Control of Hybrid VXI/PXI Systems
PXI
VXI
Desktop or
Rack-Mount PC
MXI-3 (copper or fiber-optic) or
Ethernet Link to PXI System Controller
MXI-2, MXI-3, 1394 (firewire), GPIB, or Ethernet Link
to VXI Slot 0 Controller
VXI Slot 0 Controller Options:
VXI-MXI-2, VXI-MXI-3, VXI-1394 (Firewire), GPIB,
or VXIpc Embedded Controller (for Ethernet)
PXI System Controller Options:
PXI-MXI-3 or PXI Embedded Controller (for Ethernet)

## Slide 9: PC Control of Hybrid VXI/PXI Systems (cont’d)
Primary Use Cases:
  Adding a new PXI or VXI chassis to an existing system already based on remote PC control
  You want the combination of high processor performance and low-cost of the latest PCs on the market
  Minimize large-system cost
    Use VXI chassis for existing, custom, unique, or high-performance VXI modules
    Use PXI chassis for general-purpose instrumentation

## Slide 10: VXI Control of Hybrid VXI/PXI Systems
VXI
MXI-3 (copper or fiber-optic) or
Ethernet Link to PXI System Controller
PXI
VXI Slot 0 Controller Options:
VXIpc Embedded Controller
(for MXI-3 link, use a VXIpc Embedded Controller
That contains a built-in PCI slot for PCI MXI-3,
such as the NI VXIpc-872B)
PXI System Controller Options:
PXI-MXI-3 or PXI Embedded Controller (for Ethernet)

## Slide 11: VXI Control of Hybrid VXI/PXI Systems (cont’d)
Primary Use Cases:
  Adding a PXI chassis to an existing VXI embedded controller-based system
  Minimize large-system cost
    Use VXI chassis for existing, custom, unique, or high-performance VXI modules
    Use PXI chassis for general-purpose instrumentation

## Slide 12: PXI Control of Hybrid VXI/PXI Systems
VXI
VXI Slot 0 Controller Options:
VXI-MXI-2, VXI-MXI-3, VXI-1394 (Firewire), GPIB
or VXIpc Embedded Controller (for Ethernet)
PXI
MXI-2, MXI-3, 1394 (FireWire), GPIB, or Ethernet Link
to VXI Slot 0 Controller
PXI System Controller
Options:
PXI Embedded Controller
Install MXI-2, MXI-3, 1394 (FireWire), GPIB, or
Ethernet Peripherals in PXI chassis

## Slide 13: PXI Control of Hybrid VXI/PXI Systems (cont’d)
Primary Use Cases:
  Adding a VXI chassis to an existing PXI system
  Expansion or phased transition of system to high performance and lower-cost PXI technology

## Slide 14: Distributed Control of Hybrid VXI/PXI Systems
Ethernet
Use Embedded Controllers in all systems
Transfer data between systems using software
  ie Datasocket (LabVIEW, LabWindows/CVI, C/C++, Visual Basic, etc)

## Slide 15: Mix and Match as Needed

## Slide 16: Summary of Hybrid VXI/PXI System Control Options
Many Options Exist for the following Topologies
  VXI as the hybrid system controller
  PXI as the hybrid system controller
  PC as the hybrid system controller
  Distributed control (ie multiple independent PXI and VXI controllers with inter-communication)
Combinations and permutations (ie multiple chassis systems, etc) are nearly endless. Mix or match to your needs.