## Slide 1: Acquiring Dynamic Signals using (Multiplexed) MIO Boards
Alain Moriat
Senior Architect
Wed Aug 16
12:00-1:15 p.m., 3:30-4:45 p.m.
Colorado (6A)

## Slide 2: Presentation Outline
Dynamic signal acquisition requirements
  Anti-aliasing protection – why and how
  Multi-channel simultaneous sampling – why
Artifacts of dynamic signal acquisition with a multiplexed data acquisition board
Techniques for post-processing correction
Demos
Notes: This presentation presents a new processing technique that corrects for some artifacts of acquiring multi-channel dynamic signals using a data acquisition board with a multiplexed front end such as National Instruments line of multifunction I/O (MIO) data acquisition boards.
The presentation begins with a discussion of these artifacts, which include aliasing and inter-channel delay.
Next, we discuss a post-processing algorithm that can correct these artifacts.
The presentation concludes with a real-world demonstration involving a multi-channel MIO board and LabVIEW™ 6i.

## Slide 3: What is Aliasing? (Time Domain)
Samples acquired at 1 kHz
  150 Hz sine tone ?
  850 Hz sine tone ? (1000 Hz – 150 Hz)
  1150 Hz sine tone ? (1000 Hz + 150 Hz)
Notes: The solid yellow trace on the graph shows samples of a 1 k sample/sec acquisition that did not include any anti-aliasing protection.
These samples look like the result of the acquisition of a pure sine tone of 150 Hz, but they could also correspond to the acquisition of an 850 Hz, 1150 Hz, or any frequency that fulfill the relation
Signal frequency = n * sample rate +/- 150 Hz
where n is an arbitrary positive integer value. The only valid frequency value in this case is 150 Hz. All other values are the result of aliasing.
In the general case, only values below half the sample rate (also called the Nyquist frequency) are acceptable. In this example, the Nyquist frequency is 500 Hz.

## Slide 4: Aliasing (Frequency Domain)
n * Fsampling 150 Hz
150, 850, and 1150 Hz
Notes: This graph shows the frequency-domain representation of the time-domain samples from the previous slide.
The plot shows a number of frequency components, but only the component below 500 Hz is valid.

## Slide 5: Anti-Aliasing Filter (Analog Only)
Analog anti-aliasing filter
  Passband – DC to 400 Hz
  Stopband – 600 Hz 
Notes: To avoid aliasing, you must block signals that exceed the Nyquist frequency before they are digitized. One common method of doing so is to pass the input signal through an analog filter (green trace) with the shown frequency characteristic. The filter lets any signal below 400 Hz pass through and attenuates signals above 600 Hz.
The region between 0 Hz (DC) and 400 Hz is called the passband of the filter.
The region between 400 Hz and 600 Hz is called the transition band of the filter.
The region above 600 Hz is called the stopband of the filter.

## Slide 6: Anti-Aliasing Filter (Analog+Digital)
Analog filter
  Passband – DC to 400 Hz
  Stopband – 1600 Hz 
Digital filter (2X decimation)
  Passband – DC to 400 Hz
  Stopband – 600 to 1400 Hz
Notes: Another method of aliasing protection involves the combination of analog filtering, over-sampling, digital filtering, and decimation. To understand how this works, consider that if you increase the sampling rate for the acquisition, the Nyquist rate also increases. This increased Nyquist rate means that you can increase the stopband frequency of your analog anti-aliasing filter.
In the example, we increased the sampling rate by a factor of two. To be specific, we over-sampled the signal with an over-sampling ratio (D) of 2. When we do so, the first aliasing occurs at
2 * 1000 Hz – 400 Hz = 1600 Hz
instead of 600 Hz.
Next, we apply a digital filter to our 2-kHz sampled signal. With a filter that blocks a band of frequencies between 600 and 1000 Hz (the Nyquist rate for our samples), we can block frequency components above 400 Hz.
As a final step, let’s decimate the data by D=2, that is, create a new signal by keeping only every other sample of the digital filtered data. This results in an output sample rate of 1000 Hz and a fully alias-protected acquisition system.

## Slide 7: Is the Analog Filter Always Necessary?
Measuring a pure sine tone?
Measuring a bandwidth limited signal?
  Example – vibration signal from power plant
Notes: The role of the analog filter in a combined analog-and-digital filter and over-sampled acquisition system is to remove the signal components above a frequency Fstop. Fstop is defined by the output sample rate and the decimation factor D of the system
Fstop = ( D – k ) * output sample rate,
where k represents the normalized passband frequency, ideally 0.5 (Nyquist), but in a practical situation, you use a lower value to allow the filter to have a transition band. In all previous and following examples, we use a value of k = 0.4.
The analog filter is not necessary if the spectral contain of our signal above Fstop is negligible and therefore does not affect our signal of interest when it is aliased into our passband. One case where this situation occurs is when the measured signal is a pure sine tone or a combination of tones (multitone signal) with no frequency components above Fstop. Another example is if the spectral contain of the signal rolls off at high frequencies, as the graph depicts.

## Slide 8: Is the Analog Filter Always Necessary?
Notes continued from previous slide.
Notes: Getting back to the example from the last slide, if you are only interested in signal information up to 2 kHz, you might consider an acquisition system with an A/D sample rate of 20 kHz and a decimation factor of D = 4. This results in an output sample rate of 20 kHz/D = 5 kHz, a passband frequency of 5 kHz * k = 2 kHz and an Fstop value of
Fstop = ( 4 – 0.4 ) * 5 kHz = 18 kHz
At a sample rate of 20 kHz, if you apply a digital filter with a passband frequency of 2 kHz, the first aliasing would appear above 20 – 2 = 18 kHz, where the measured signal appears to be attenuated more than 50 dB.

## Slide 9: Is the Analog Filter Always Necessary?
Combined responses for piezoelectric accelerometer and second order LP-filter
Notes: In certain cases, the transducer frequency characteristic might help avoid the need of any additional analog anti-aliasing filtering.
Let’s consider a real-world example involving the acquisition of a vibration signal using a piezo-accelerometer with a resonance frequency of about 5 kHz. As you can see from the graph, the frequency roll-off of the transducer (brown trace) above the resonance frequency follows a second-order (40 dB/decade) characteristic. If we combine this with the second order Butterworth low-pass filter that is typically built in the transducer conditioning unit, we get a transducer system gain below –60 dB at Fstop = 18 kHz.
A physical signal above 18 kHz that is 50 dB below the signal in the passband acquired with a transducer that attenuates frequency above 18 kHz more than 60 dB results in an overall signal level above Fstop that is more than 50 + 60 = 110 dB below the passband signal level. At this signal level, additional analog filtering is unnecessary.

## Slide 10: Classic Multiplexed MIO
Low cost/flexible
  No anti-aliasing filters
  Only one A/D converter for all channels
Conflicts with some common requirements of many applications that require dynamic signal acquisition
  Aliasing protection
  Simultaneous sampling
Notes: Let’s next consider how this relates to a general-purpose MIO (Multipurpose Input/Output) data acquisition board.
The conflicts shown on the second point apply to most general-purpose data acquisition boards (National Instruments and other vendors).

## Slide 11: Multiplexing MIO Board
Some definitions
  Channels – the actual number of input channels scanned by the board
  Scan clock – the output data rate for each channel
Notes: Channels – the actual number of input channels scanned by the board. Note that this does not represent the maximum number of channels supported by the board (common numbers are 16 or 64), but the number of channels defined by the user for a particular acquisition.
Scan clock – the output data rate for each channel. For each scan, the board returns one sample for every acquiring channel for a total number of samples equal to the number of channels.

## Slide 12: Multiplexing MIO Board
Some definitions
  Decimation factor (D) – the acquisition over-sampling factor for each channel
  A/D clock – the actual sample rate of the multiplexing A/D converter
A/D clock = channels * decimation * scan clock
Notes: Decimation factor – the acquisition over-sampling factor for each channel, that is the ratio between the actual acquired scan clock and the output scan clock exported after digital filtering and decimation.
A/D clock – the actual sample rate of the multiplexing A/D converter
A/D clock = Channels * Decimation * Scan clock
Note – to simplify the concept, the above definitions assume that the board can physically generate the actual A/D. In certain cases, limitations in clock resolution do not allow this. In these cases, the software tools demonstrated in this presentation automatically adjust the parameters to find an optimum combination.

## Slide 13: Multiplexed Input
4 channels (same input signal on all channels)
Scan clock = 1 kHz
A/D clock = 4 kHz
Notes: Sampling multiple channels with a multiplexed MIO board results in an undesirable artifact called inter-channel delay.
What does this inter-channel delay do to our data? To answer this question, consider what happens when we use four channels of a multiplexed MIO board to sample the same signal. The graph above shows the real waveform, along with the sampling points for each channel.
The acquisition of 4 channels at a scan rate of 1 kHz requires a 4 kHz A/D clock.
Since the A/D clock is 4 kHz corresponding to a sample period of 250 μs, the delay between 2 adjacent channels is also 250 μs.

## Slide 14: Resulting Delayed Acquisitions
Our four channels appear to have different phases even though we input the same signal to each
Scan clock = 1 kHz
A/D clock = 4 kHz
Notes: If we examine the samples from the four channels of our example, we see what appears to be four separate waveforms, each with a different phase. Because our inter-channel scan rate is 4 kHz, the apparent delay between two consecutive channels will be the sample period, 250 microseconds.
This assumes that your are considering all of the samples acquired during a scan to be acquired at the same time, that is having the same time stamp t0.

## Slide 15: Relative Phase Responses
4 channels
Scan clock = 1 kHz
A/D clock = 16 kHz (over-sampled 4X)
Notes: The constant inter-channel delays caused by the multiplexed acquisition is equivalent to a frequency-dependant inter-channel phase errors.
Now, by increasing the decimation (over-sampling) factor, you can reduce the phase error, because at higher rates less time is spent between samples. Nonetheless, even using a decimation factor of 4 as the example shows, the worst relative phase error is as high as 27 degrees.

## Slide 16: What If
We over-sample each channel and decimate them using appropriate digital filters?
  We fix the aliasing problems (assuming the transducers and the measured signals are “fairly” bandwidth limited)
Notes: You can partially correct for the lack of aliasing protection with an MIO board by following these steps.
Acquire the data at a higher rate than the needed scan rate (over-sampling by a factor of D)
Apply appropriate digital filtering
Decimate by the decimation factor D
Recall that with a high enough decimation factor, your system might not require extra analog filtering. Instead, you can rely on the inherent low-pass characteristics of your transducer.

## Slide 17: What If
The digital filters individually correct for each channel’s relative delay?
  We fix the inter-channel delay due to non-simultaneous sampling that is inherent with a multiplexed front end.
Notes: Digital filters can also correct for the inter-channel errors we discussed.

## Slide 18: Data Post-Processing (Host PC)
One common process for
  Anti-alias filtering (digital only!)
  Inter-channel delay correction
Notes: The same filters that perform digital filtering and decimation can also perform the correction of the inter-channel delay errors, thus allowing us to perform both operations in a single process. This is important because the host PC must supply processing power for these computations.

## Slide 19: Multiplexed Filter Coefficients (FIR)
1 channels, decimated 20X
4 channels, each decimated 5X
Notes: The process of designing the appropriate filters to perform digital anti-alias protection, decimation, and delay error correction is not trivial. But the concept is similar to the generation of poly-phase filters, such as those used for interpolation.
A classic FIR decimation filter of order A*D, where A and D are integer numbers, could have an impulse response of length A*C taps.
If the impulse response of this filter is split into A interleaved impulse responses of each length C, it results in a bank of A filters, each performing the filtering needed for a decimation by D and with a group-delay increasing more or less accurately with 1 / (A/D clock) from filter to filter. This group-delay variation compensates exactly for inter-channel delay errors.
In the example above, an original filter designed to perform decimation by 20 is split into 4 filters, each decimating 5 times. In other words A = 4, C = 17, and D = 5.

## Slide 20: MIO Solution (NI-DAQ™)
New pre- and post-processing functions
Notes: How does this fit into LabVIEW and NI-DAQ?
Performing the digital anti-alias protection and the inter-channel delays correction with multi-channel data using an MIO board, require two additional VIs.
A pre-configuration VI configures the MIO board to acquire the different channels with the specified over-sampling factor D. The VI also generates the appropriate filters used by the post-processing VI.
A post-processing VI inputs the multi-channel’s over-sampled and multiplexed data delivered by the AI read VI and performs the digital decimation and inter-channel delay correction.

## Slide 21: MIO Solution
Performance (real-time)
  Depends on the PC processor
  Typical (Pentium III, 500 MHz)
  (Channels. * Decimation * Fscan) > 500 kSamples/s
  for example
  16 channels, 4X, Fscan = 10 kSamples/s
  64 channels, 2X, Fscan = 5 kSamples/s
Notes: Since the host PC performs the post-processing of the acquired data is, the PC performance sets an upper limit for the real-time capability of the system.
The operations needed to process data is almost proportional to the A/D clock rate.
Channels * Decimation * Scan clock
On a typical Pentium III processor, real-time acquisition and processing can be sustained for an A/D clock of more than 500 kHz, so you can sustain real-time acquisition and processing on setups like
4 channels, D = 4, scan clock = 40 kHz, signal bandwidth = 16 kHz or
16 channels, D = 4, scan clock = 10 kHz, signal bandwidth = 4 kHz or
64 channels, D = 2, scan clock = 5 kHz, signal bandwidth = 2 kHz

## Slide 22: MIO Solution – Tradeoffs
Specifications
  Different options trade-off between
  Computation speed (limited by the PC)
  Filter quality (passband ripple and alias rejection)
  Global delay (impulse response length)

## Slide 23: MIO Solution – Specs
Typical specifications, with the corrections we discussed
  Alias rejection > 90 dB, ripple < 0.02 dB
  Inter-channel amplitude mismatch < 0.0005 dB
  Inter-channel phase mismatch < 0.005 deg
Notes: The different filter options used to process the multi-channel data are designed to give high quality measurement results. They have a high level of aliasing rejection, typically more than 90 dB, thus supporting a dynamic range expected with a 16-bit board. In addition, with a passband amplitude ripple typically less than 0.02 dB, we can perform accurate absolute amplitude measurements.
Furthermore, due to the fact that all channels physically share most of their analog path and the same A/D converter, the match between two arbitrary channels is extremely good. The amplitude mismatch between two arbitrary chosen channels is typically under 0.0005 dB, while the phase mismatch is less than 0.005 degree, making relative calibration unnecessary.

## Slide 24: Recap
Your application requires
  Multi-channel acquisition
  Measurement bandwidth of bw Hz corresponding to a scan clock of bw / 0.4 Hz
Notes: To summarize, you might consider the over-sampling, decimation, and phase correction technique if:
Your input signals have a bandwidth corresponding to a scan clock = bandwidth / 0.4. This constraint does not eliminate the need for some type of anti-aliasing. Instead, it implies that you might not need an extra anti-aliasing filter(s) if other component(s) of your measurement system provide the filtering.
Multi-channel acquisition

## Slide 25: Checklist for Applicability
Single channel check
Choose an over-sampling factor D
A/D clock support
Select an appropriate digital filter
Check PC performance
Notes: You can go through the checklist to ensure the applicability of this technique to your application.
Step 1 – single channel check
As stated on the previous slide, it is important that you do not introduce aliases in any of the channels of interest. To make sure that this does not happen, you can check (one by one) each channel for aliases. You can do this by sampling your signal at a high enough rate to define a safe fstop frequency.
Step 2 – choose an appropriate over-sampling factor D using the formula
D >= Fstop / scan rate – k
Step 3 – A/D clock support?
After running check of step 1, you have an idea of the bandwidth of your measurement. Divide the bandwidth by 0.4 to determine the scan clock requirement. Determine the A/D clock requirement by finding the product of the scan clock, the decimating factor (D), and the number of channels (N).
Step 4 – selection of a digital filter
An appropriate digital filter trades off ripple, attenuation, and impulse response length versus computational requirements.
Step 5 – check PC performance
Use the real-time performance evaluation VI to check your PC performance. (If your processor cannot process data at the A/D clock rate during acquisition, you might stream the data to disk and perform these corrections offline.)

## Slide 26: Demos
PCI-6071E, 64 channels, 1.25 MHz, 12 bit

## Slide 27: Demo Setup