## Slide 1: Using PXI™ to Build Measurement and Automation Systems
CompactPCI

## Slide 2: Overview
PXI™/CompactPCI architecture
  Mechanical
  Electrical
  Software
Third party support
  PCI Industrial Computer Manufacturing Group (PICMG)
  PXI Systems Alliance
Configuring a PXI system
  Products and technology
  Applications
Notes: PR2B1-<number>
Using PXI to Build Measurement and Automation Systems
<header>
<date/time>
PXI and CompactPCI technology revolutionize how engineers and scientists build systems for test and measurement, data acquisition, and industrial automation. This seminar covers the basics of the PXI and CompactPCI specifications including the mechanical, electrical, and software features of these open standards. The seminar also takes you through the steps necessary to configure a PXI system. This information helps you get started in building PXI/CompactPCI systems that deliver faster system performance, lower costs, better integration, and faster system development.
More details on PXI/CompactPCI specifications and products can be found at
ni.com/pxi
A full copy of the PXI specification is available online at no charge.

## Slide 3
Price
Performance
Desktop PC
VXI
Modular instrumentation
for mainstream users
Filling the Gap
Notes: Customers using desktop PCs for a variety of applications face limitations resulting in the need for more:
Users leverage the latest desktop PC technology in industrial environments where ruggedness maximizes performance and value.
While desktop PCs are “closing up” (the number of expansion slots is decreasing), industrial users need more PCI slots as their applications scale up (move from the laboratory onto the production floor).
Users also need advanced timing and synchronization features for instrumentation and automation applications. These capabilities are not built in to the desktop PC architecture.
There has been a relatively large gap between low-cost desktop PC solutions and high-end VXI and GPIB solutions.
Today, users have another option. By extending CompactPCI, PXI fills the gap between PCs and high-end test platforms such as VXI. PXI effectively brings modular instrumentation to the mainstream user.

## Slide 4
CompactPCI puts PC technology in a small, rugged package by combining three standards
PCI bus
Eurocard packaging
Better IEC connectors
CompactPCI
PXI Starts with CompactPCI
Notes: PXI (PCI eXtensions for Instrumentation) is a system-level specification based on the CompactPCI core specification. CompactPCI brings together three standard technologies:
PCI (from the desktop PC industry)
Eurocard mechanics (from the VME and VXI industries)
Standard IEC, high-performance connectors (from the telecommunications industry)
CompactPCI gives PC technology a small, rugged form factor. The primary strength of this architecture is that it is based on the popular PCI bus that is used in today’s desktop PCs.

## Slide 5
Wide industry support
Plug and play capability
Thousands of software products
32-bit data transfers at 33 MHz (132 Mbytes/sec)
PCI is a de facto standard
Why use the PCI bus?
Notes: PCI is the de facto computer bus standard because it is driven by the multi-billion dollar desktop PC market. As a result, PCI offers the latest technology with state-of-the-art performance.
In addition to the state-of-the-art technical features of PCI, the key benefit of implementing PCI is its broad hardware and software support in the computer industry.
Because extensions of PCI for 3.3 V operation, 64-bit data transfers, and 66 MHz operation are already defined, PCI has a long life as a standard bus architecture. Linking multiple PCI buses is also already defined with standard PCI-to-PCI bridging technology. CompactPCI and PXI systems use the same extensions and technology.

## Slide 6: Why Use Different Connectors?
Higher pin count
Rugged and reliable
Integrated shield
Impedance matched
Most importantly: MORE SLOTS!
J1
J2
Notes: The pin-and-socket connectors of CompactPCI are much more rugged than the card-edge connection scheme in desktop PCs. The IEC standardized the connectors, so numerous connector suppliers can offer interchangeable products.
The smallest CompactPCI package, 3U, consists of two connectors that carry the 32-bit and 64-bit PCI signals with room to spare. J1 and J2 each have 110 signal pins, as well as 22 pins worth of ground connections through the integrated metal connector shield. The high density of connections takes up less space than the card-edge connector scheme.
The CompactPCI connectors were designed with electrical characteristics in mind from the start. Designers used extensive modeling and simulation for the final design, which has since seen wide use in numerous industries, particularly telecommunications. Integrated guides, keying, and gas-tight connections make these connectors well-suited to industrial environments.
Extra slots are the most important benefit that these connectors deliver. Due to their excellent electrical characteristics, the connectors used in CompactPCI yield more available slots per single PCI bus segment. A CompactPCI system can have seven peripherals on one bus segment, while desktop PC systems can have a maximum of four peripheral slots.

## Slide 7: Why Use Eurocard Packaging?
Proven through decades of use in industrial applications (VME, VXI, and so on)
Defined by IEEE 1101 standard
Readily available mechanical components
Excellent thermal properties
Modular, rugged, and compact
Notes: For years, a variety of industrial bus architectures have used the modular Eurocard platform successfully. IEEE standardized the dimensions and mechanical components resulting in numerous vendors that supply mechanically interchangeable components such as card cages, card guides, ejector/injector mechanisms, and front panels.
Eurocard systems deliver excellent support of plug-in cards with card guides on the top and bottom and two screws on each front panel. Users can easily insert and remove modules with the standardized injector/ejector mechanisms.
The cards in a Eurocard system stays cooler than in a desktop PC. Cards and the backplane in a Eurocard system line up vertically for an unobstructed airflow path from bottom to top.

## Slide 8: PC Motherboard = Controller + Backplane
1 2 3 4 5 6 7 8
1 2 3 4
PC motherboard
with four PCI slots
CompactPCI
embedded controller
CompactPCI
8-slot backplane
Notes: In a PXI/CompactPCI system, the embedded controller takes on the active functions of a PC motherboard while the backplane provides the connectivity between the controller and the peripheral modules. With the active PC components on a controller module, upgrades to a system are easy.

## Slide 9: Peripheral Sizes in PCI and CompactPCI
PCI
PXI/CompactPCI
Half
size
3U
Full
size
6U
Users can redesign PCI boards to fit in PXI/CompactPCI with little or no electrical changes
Notes: PXI/CompactPCI offers two size options. 3U is the most compact and the most popular for measurement and automation I/O modules. 6U is available for larger systems. The 6U scheme roughly translates to the half-size and full size PCI cards used in desktop PCs. With 3U modules, users have room to accommodate dual circuit cards in a single slot width so users can convert many full-size PCI designs into the compact 3U size.

## Slide 10
PXI Combines Standard Technologies
CompactPCI
PCI
Notes: The major strength of PXI is that it uses proven industry-standard technology. PXI leverages off of CompactPCI and adds the best features of Windows and VXI.
The open PXI specification governed by the PXI Systems Alliance defines the details of PXI. PXI builds on the CompactPCI core specification and defines additional mechanical, software, and electrical requirements for building components for modular instrumentation systems in measurement and automation applications.
Because PXI and CompactPCI are completely interoperable, users can use any core CompactPCI product in a PXI system and vice-versa.

## Slide 11
Software
Electrical
Mechanical
The PXI system specification
extends CompactPCI for
instrumentation.
S
Specification
Revision 2.0
The Open PXI System Specification
Notes: The PXI specification, governed by the PXI Systems Alliance, defines specific software, mechanical, and electrical features and is divided into respective sections.
Users can download copies of the specification from ni.com/pxi or pxisa.org

## Slide 12: Mechanical Extensions
PXI improves reliability and interoperability with
Mandatory active cooling
Temperature testing and ratings
System-level environmental specifications for EMC, shock, vibration, and humidity
Defined embedded controller location
Notes: PXI offers all of the mechanical features defined by CompactPCI.
In addition to the CompactPCI mechanical specification, PXI requires integrated active chassis cooling. PXI also requires temperature ratings for all components of a system. The PXI specifications include recommendations for shock, vibration, and humidity testing and documentation. The specification also lists requirements for EMC.
PXI specifies that the system controller is on the left end of the chassis and that it should not expand into peripheral slots. Instead, controllers requiring more space should expand to the left into space set aside for the controller only. This definition clarifies interoperability and maximizes the expandability of a PXI system by protecting valuable I/O slots.

## Slide 13: PXI/CompactPCI Form Factors
All 3U cards work in 6U systems
64-bit PCI and
PXI Features
3U
PXI/CPCI
J1
J2
32-bit PCI
6U
PXI/CPCI
J5
J4
J3
J2
J1
64-bit PCI and
PXI Features
32-bit PCI
PXI Reserved
PXI Reserved
PXI Reserved
6U Adapter Panel
Notes: CompactPCI defines two basic sizes for modules – 3U and 6U. PXI uses the same form factor to ensure interoperability with CompactPCI. A unique benefit of the 3U/6U architecture is that you can use any 3U module in a 6U system since the connectors build from the bottom up. A simple adapter panel (NI part number 777763-01) fills the space above a 3U module in a 6U system.
The five connectors used by PXI and CompactPCI systems are:
J1 – defined for 32-bit CompactPCI
J2 – partly defined by CompactPCI for 64-bit PCI, the remaining pins defined by PXI for timing and triggering features
J3, J4, and J5 – reserved by PXI, some CompactPCI systems use these connectors for a variety of other functions including rear panel I/O and application specific features such as telecom buses
J1 and J4 – integrated connector keys. Keying is also provided with the IEEE 1101.10 defined ejector handles and card guides. Keying safeguards against potential interoperability issues.

## Slide 14
Software Extensions
PXI speeds application development because
PXI Controllers MUST support a standard software framework by including a pre-loaded operating system (OS)
Peripherals modules MUST be supplied with a WIN32 device driver
Notes: PXI defines two standard software frameworks – Windows NT and 9x. Any PXI compliant product must support at least one of the frameworks. A PXI embedded controller must come with a Windows operating system installed. Also, peripheral module manufacturers must supply device driver software so that users can quickly develop applications. An implementation of the industry standard Virtual Instrument Software Architecture (VISA) is also defined for PXI.

## Slide 15
PXI and PC Software Is Identical
Operating systems and application software run unchanged on PXI systems
Configuration tools recognize PXI modules as PCI devices
“We ported from a desktop PC to PXI without changing any code” – Len Burczyk, Los Alamos National Laboratories
Notes: Los Alamos National Laboratories chose PXI for its NTVision application. They ported from a development PC to PXI in about an hour without any code changes. This was possible because the image acquisition board used in the PCI desktop computer was also available in the PXI form factor. Thus, the software was exactly the same. This example shows the power of leveraging PC technology.

## Slide 16
Trigger bus
System
controller
Star trigger
controller
Peripheral
Peripheral
Peripheral
10 MHz
CLK
132 MB/s, 33 MHz, 32-bit computer bus
Star trigger
Local
bus
Electrical Extensions
PXI timing and triggering improves performance
Notes: Electrically, PXI adds a trigger bus (eight lines), a star trigger (13 slot-specific signals), a 10 MHz reference clock (equally distributed), and local buses (13 lines wide). The star trigger control function runs separate from the main system controller for standard CompactPCI controllers and to require users to only pay for the feature if they need it.
System reference clock – a highly accurate 10 MHz clock is independently distributed to each PXI peripheral slot through equal-length traces. Multiple cards can use this common time base for synchronization of events.
Trigger bus – eight bused trigger lines link all PXI slots in a bus segment so that multiple boards can interact and control each other’s events through hardware.
Star trigger – a special star trigger slot provides an independent dedicated line for each of up to 13 peripheral slots on a single backplane. When you plug a star trigger controller into this slot, it can be used to control, monitor, or route triggers among peripheral slots with low skew (within 1 nanosecond).
Local bus – a right and left local bus composed of 13 lines is available to each peripheral slot for private communication between adjacent slots. Pairs of peripherals can pass analog or digital signals back and forth using the local bus.

## Slide 17
Local Bus Triggers Clock Star Bus
VXI 12 lines 8 TTL, 2 ECL 10 MHz ECL D-size only
PXI 13 lines 8 TTL 10 MHz TTL 1 per slot
PXI Leverages VXI Features
Timing and triggering extensions in PXI and VXI
Notes: PXI leverages the instrumentation extensions in VXI. Note that ECL logic signals are not implemented in PXI as they require an additional power supply which adds significant cost to the system.

## Slide 18: Trigger Bus Synchronizes Modules
PXI-6070E
NI 5102
PXI-6533
Start scan
RTSI oscillator
Multifunction I/O
Oscilloscope
Digital I/O
Notes: Many National Instruments data acquisition, instrumentation, and communication modules take advantage of the PXI trigger bus by implementing real-time system integration (RTSI) timing signals. Users can easily program multiple modules for simultaneous data acquisition. This type of function is an important benefit of PXI.
Similarly, switch multiplexers from National Instruments take advantage of the PXI trigger bus by handshaking with a PXI digital multimeter module across the trigger bus. By using these timing signals, a switching system can guarantee the maximum scan speed because the hardware does scan time and routes it directly across the backplane.

## Slide 19: PXI Architecture Summary
Electrical
Mechanical
Software
CompactPCI
PCIbus
EuroCard
High performance connectors
Components
Systems
PXI builds on CompactPCI to define complete measurement and automation systems
Notes: PR2B1-<number>
Using PXI to Build Measurement and Automation Systems
<header>
<date/time>
PXI is based on the CompactPCI core specification that brings together three standard technologies:
PCI (from the desktop PC world)
Eurocard mechanics (from the VME and VXI worlds)
Standard IEC, high-performance connectors (from the telecommunications industry)
The open PXI specification governed by the PXI Systems Alliance defined the details of PXI. PXI builds on the CompactPCI core specification and defines additional mechanical, software, and electrical requirements for building components for modular instrumentation systems for measurement and automation applications.
All of the PXI extensions maintain complete interoperability with the core CompactPCI specification. The major strength of PXI comes from its dependence on proven, industry-standard technology. PXI leverages off of CompactPCI and adds the best features of Windows and VXI.
Download copies of the specification from ni.com/pxi or pxisa.org

## Slide 20: The CompactPCI Specification Body
PCI Industrial Computer Manufacturers Group
Governs the core CompactPCI specification
Focus on telecommunications infrastructure
Online directory contains hundreds of products
More than 500 company members
www.picmg.org
Notes: The PICMG specifications development group governs the core CompactPCI specification (PICMG 2.0). The three membership levels of PICMG are executive, associate, and affiliate. National Instruments is an executive level member and has participated in CompactPCI specification development since 1996.
Due to telecommunications companies’ strong interest in the CompactPCI architecture, PICMG subcommittees developed extensions of CompactPCI that target telecom infrastructure equipment. Other extensions of the core CompactPCI specification include PXI, which is targeted at measurement and automation applications. The PXI Systems Alliance governs the PXI specification and is referenced by the PICMG through an instrumentation subcommittee designated as PICMG 2.8.
The PICMG Web site includes member listings and a searchable product directory. You can purchase CompactPCI specifications from the PICMG.

## Slide 21: The PXI Specification Body
Charter of this group is to - Promote PXI - Ensure interoperability - Control the PXI specification
Focus is on user success in measurement and automation
More than 50 company members
Systems Alliance
www.pxisa.org
Notes: The PXI Systems Alliance, a group of 50 companies determined to ensure user success controls the PXI specification. The PXI Systems Alliance formed around a charter document that defines the goals of the organization, which include:
promotion of PXI and CompactPCI for measurement and automation
maintenance and revision control of the PXI specification
ensuring the highest levels of multi-vendor interoperability at the mechanical, electrical, and software levels
The PXI Systems Alliance Web site contains:
white papers
free soft copy of the PXI specification
list of all members
upcoming event information
product directory arranged by category

## Slide 22: PXI Systems Alliance Members
GenRad
Gespac
Goepel Electronic
GDE Systems (Marconi)
GTE-ERS
Innovative Integration
KineticSystems
LeCroy
MAC Panel
MEN Mikro Elektronik
Talon Instruments
Acqiris
Advanced Power Designs
Advanced Test Methods
Alphi Technology
AMP
Analogic
ASCOR
ARVOO Engineering
ATEME
A&T Engineering Tech. Ctr.
B&B Technologies
BittWare Research Systems
Bode Enterprises
BRIME
C&H Technologies
CHROMA ATE
Data Patterns
Dateppli
Datum
Dolch Computer Systems
National Instruments
Pickering Interfaces
PX Instrument Tech.
Racal Instruments
Rohde and Schwarz
Quantum Controls
SBS GreenSpring
Shaanxi HiTech
SRC
TestWare
Tracewell Systems
TTI Testron
Vero Electronics (APW)
Virginia Panel Corp.
ZYNX
Schroff
Marek Micro
ERNI
Ballard Technology
Notes: PXI Systems Alliance members include instrument suppliers, platform suppliers (chassis/controllers), component suppliers, and system integrators. For more information on the alliance and links to member companies refer to
pxisa.org

## Slide 23: Configuring Your PXI System
1. Choose your OS and application development software
2. Select your PXI chassis
3. Select a remote or embedded controller
4. Select your PXI modules
5. Select your accessories
6. Select NI Installation Services
Notes: You can configure a complete PXI system by following five basic steps.

## Slide 24
Step 1 – Choose Your Software
Select a Windows operating system
Select a standard application programming interface
  Graphical programming – LabVIEW™
  Text-based programming – Measurement Studio™
  Real-time program development – LabVIEW RT
Notes: Software is one of the most important considerations for your PXI system. PXI compliant products come with driver software for Microsoft Windows operating systems. This ensures that you can quickly integrate your plug-in modules into your system when you work with Windows.
With any Windows operating system, you can choose from any off-the-shelf application software to speed development. For the most powerful tools for test and measurement and data acquisition, use LabVIEW™ graphical programming or Measurement Studio™ text-based programming in C and C++.
For applications that require hard real-time program execution, use LabVIEW RT with RT series hardware available in the PXI format.

## Slide 25: Step 2 – Select Your Chassis
Choose enough slots for room to grow
  four or eight slot chassis
  use MXI™-3 for multichassis configurations
Consider application requirements
  portable or embedded
  bench-top or rack-mount
Integrated signal conditioning (SCXI™)
  combination chassis with PXI and SCXI
Notes: The requirements of your application tie directly to your choice of PXI chassis. National Instruments offers several choices that span a variety of applications. Choose from four or eight slot chassis that house one controller and three or seven peripheral modules respectively. All National Instruments PXI chassis feature filtered, forced-air cooling and conformance to strict requirements for reliable operation in noisy, humid, vibration, and shock-prone environment. Chassis with integrated screens are available for portable applications and some chassis come with integrated SCXI™ slots for signal conditioning. All chassis are available with rack-mounting kits.

## Slide 26: Flexible Installation Options
PXI/CompactPCI
For applications in
Test and measurement
Data acquisition
Industrial automation
Notes: The modular, compact, and rugged form factor of CompactPCI makes PXI systems suitable for a wide range of installations including bench-top, rack-mount, portable, and embedded. With a variety of chassis options, PXI is an ideal system platform for applications in test and measurement, data acquisition, and industrial automation. PXI is an attractive standard for a variety of applications because it is based on familiar PC technology and specifies all components of a system (including software, mechanical, and electrical components), for quick and easy integration.

## Slide 27: Step 3 – Select Your Control Scheme
Embedded controllers
  Most compact solution
  Modular = easy maintenance
  Pentium and Pentium III class
  Windows or LabVIEW RT (8156B/333)
Remote MXI-3 controllers
  Short or long distance
  Fully transparent
  Low-cost
Notes: You need to choose whether you want to control your system from a standard PC with MXI-3 or with a fully integrated embedded computer.
If your application precludes the use of an external PC and you require a completely integrated, compact solution, then you need an embedded computer. You need an embedded computer to control portable chassis with integrated displays. National Instruments offers a variety of embedded computers with your choice of processor speed, I/O configuration, operating system, and application development software. All embedded computers include a built-in hard drive, floppy drive, and video. Peripheral ports such as USB, serial, parallel, mouse and keyboard are standard. GPIB and Ethernet ports are available as options on some embedded computers.
For real-time control applications, use an embedded PXI controller as a LabVIEW RT target that controls a chassis full of data acquisition modules.
MXI™-3 kits give you a fully transparent link from your PC to your PXI system. You plug a MXI-3 interface card into your PC and connect it with a copper or fiber optic cable to a MXI-3 module in slot 1 of your PXI system. MXI-3 extends the PCI bus at full speed so that the processor in your PC transparently configures and controls the PXI modules – whether the PXI system is 2 m or 200 m away. With MXI-3, a low-cost control solution, you to benefit from the latest in PC technology.

## Slide 28
Host
LabVIEW RT on
PC running Windows NT/9x
Ethernet
Target
RT Engine on
PXI controller
Target RT Controller
LabVIEW RT on Embedded Controller
Notes: With the LabVIEW RT engine running on an embedded controller, you can control a chassis full of data acquisition modules in real-time. You program in a standard LabVIEW development environment on a host PC running a Windows operating system. Then, you download the application to the target controller in the PXI system through Ethernet to control the PXI modules in real-time.

## Slide 29
Host
LabVIEW RT on Embedded controller running Windows NT/9x
Target
RT engine on PCI or PXI/CompactPCI DAQ board
Target RT peripheral
LabVIEW RT on DAQ Modules
Notes: Another way to use LabVIEW RT in a PXI system is through the intelligent peripheral modules. RT peripherals are PXI (or PCI) modules with their own on-board processor. Develop your LabVIEW application from a Windows environment on the host embedded controller. Then download the application to the RT peripheral across the backplane to run a data acquisition application in real-time, independent of the host embedded controller. Currently, National Instruments offers three RT peripherals for PCI and PXI for analog and digital data acquisition and control.

## Slide 30: Multisystem Extension Interface MXI-3 architecture
MXI-3
Serial I/F
MXI-3
High speed serial link
1.5 GHz
Primary
Secondary
PCI to PCI Bridge
Still software transparent!
Serial I/F
PCI
CompactPCI
PCI 1
PCI 0
Notes: PCI to PCI bridging is designed for monolithic (single box) systems and implemented as a single chip mounted on a PC motherboard. Therefore, standard PCI to PCI bridges do not provide a ready link to external systems.
MXI-3 improves on the standard PCI to PCI bridge architecture by splitting the bridge and connecting the halves through a 1.5 Gigabit/s serial link. The system interfaces with the MXI-3 link exactly as it would with a standard PCI to PCI bridge. Transfers between the two PCI buses occur according to the windowing rules provided in the PCI to PCI Bridge Architecture Specification. MXI-3 benefits fully from all of the benefits of PCI to PCI bridging because it complies with the detailed requirements defined in the specification.
The circuitry required for each half of the PCI to PCI bridge is in the National Instruments MXI-3 ASIC. The high-speed serial link implements with standard driver components, connectors, and cables.
Every MXI-3 link consists of a primary and a secondary MXI-3 card and a copper or fiber optic cable connecting the two cards. A primary MXI-3 card corresponds to the primary PCI interface of a monolithic PCI to PCI bridge. It acts as a PCI initiator and target, but does not provide any additional PCI functions. A secondary MXI-3 card provides clock distribution and arbitration for its PCI bus segment.

## Slide 31: MXI-3 Configurations
PCI to PXI/CompactPCI
PXI/CompactPCI
to PXI/CompactPCI
Cabling - Copper (up to 10 m) - Fiber Optic (up to 200 m)
Notes: By plugging a PCI MXI-3 card into a desktop PC and a PXI MXI-3 card into slot 1 of a CompactPCI/PXI chassis and connecting the two cards with a cable, you transparently link the CompactPCI/PXI chassis to the PC. The application and driver software run unmodified on the PC as if the PC were embedded in slot 1 of the CompactPCI/PXI chassis. By plugging a PXI MXI-3 card into any of slots 2 to 8 of a CompactPCI/PXI chassis and a PXI MXI-3 card into slot 1 of an additional CompactPCI/PXI chassis and connecting the cards with a cable, one CompactPCI/PXI chassis can be expanded to two. The PXI MXI-3 cards detect the slot they reside in and automatically configure as primary or secondary.
Since MXI-3 complies with the standard PCI to PCI bridging architecture, it can take advantage of a variety of other extension configurations. MXI-3 can be used to daisy-chain multiple systems in row, all under the control of the primary system – in this case a PC. You also can use star configurations with multiple MXI-3 cards in the primary system.
MXI-3 works with both copper and fiber optic interconnects. Copper links are the least expensive interconnect and can span up to 10 m between systems. Fiber optic links offer connections between systems separated by up to 200 m without repeaters.

## Slide 32: MXI-3 Benefits
More slots for PCs and PXI/CompactPCI
High-performance serial link
Easy to integrate – software transparent
Short and long distances
Low-cost
Notes: MXI-3 expands the range of possible applications for PCs and PXI/CompactPCI by giving the user more I/O slots. With both copper and fiber optic interfaces, the distance between systems can be as great as 200 m. By providing a high-speed, software transparent link MXI-3 gives users an easy way to expand their test, measurement, data acquisition, and industrial automation systems.

## Slide 33: Step 4 – Select Your I/O Modules
Interfaces
GPIB
SCSI and Ethernet
PCMCIA
CAN and DeviceNET
RS-232
RS-485
DAQ and control
RT series
Multifunction I/O
Digital I/O
Analog output
Image acquisition
Motion control
Counter/timers
Instruments
Oscilloscopes
Digital multimeters
Arbitrary waveform generators
Temperature meters
Switch multiplexers
Serial data analyzers
MXI-2 kits for VXI/VME
Choose from more
than 50 NI modules:
Notes: National Instruments offers the widest variety of modules for PXI and CompactPCI systems. Whether your application is in test, measurement, data acquisition, or industrial automation you can choose modules to meet your specific application’s requirements. National Instruments offers more than 50 peripheral modules, including instrumentation, data acquisition, motion control, image acquisition, and industrial communications modules. Interfaces to other buses such as PCMCIA, SCSI, Ethernet, RS-232, RS-485, CAN, VXI, VME, and GPIB are also available.

## Slide 34
PXI Modules Link to Other Standards
MXI
GPIB
VXI or VME
Stand-alone
instrument
PXI system
modules
modules
CompactPCI
MXI-3
Notes: Just as PXI leverages off of other standards, these other instrumentation architectures can leverage off of PXI technology. With readily available interface modules, you can use a PXI system as the main controller for VXI, VME, and GPIB systems. PXI and CompactPCI modules and systems are completely interoperable. Because PXI also works with these other standards, you preserve your investment in other technologies.
Until recently, you could not use standard desktop PCs to directly control a PXI system without an embedded controller. With the introduction of the new MXI-3 technology, direct PC control of PXI/CompactPCI systems is now possible.

## Slide 35: Third-Party Modules
Use online directories
  ni.com/pxi
  pxisa.org
  picmg.org
Look for PXI module compliance data
  Environmental specifications
  Regulatory specifications
  Driver software support
Notes: Since numerous companies around the world support PXI and CompactPCI open standards, you can complete your system with special purpose modules from other vendors.
Several online resources are available for locating PXI/CompactPCI products. When selecting modules, look for compliance to the PXI specifications to ensure that you can quickly develop a reliable system.

## Slide 36: Step 5 – Select Your Accessories
Connectivity scheme
  Cabling
  Terminal blocks
  Direct front panel connection
  Third party receivers
Signal conditioning
  Chassis-based – SCXI
  External
Mounting kits
Notes: Connectivity is an important part of every application. Complete high-performance systems require superior, easy-to-use connections. National Instruments broad selection of cable assemblies, terminal blocks, and other accessories connect your PXI system to test fixtures, field sensors, and process control points. Signal conditioning options range from low-cost, single-channel conditioning (SCC) to high-performance signal conditioning for large signal counts (SCXI).

## Slide 37: Step 6 – Select NI Installation Services
Receive a ready-to-use PXI system
Configure online with PXI configurator
National Instruments installs
  PXI modules in your chassis
  Memory upgrades
  National Instruments application software
  All required driver software
Notes: For the fastest, easiest way to get your PXI system up and running, use Installation Services. National Instruments technicians configure your PXI system to your specifications, and deliver the system ready to use. We can set up the system for you, so you can get started with your new system straight out of the box.
Installation Services delivers you a ready-to-use PXI measurement and automation system. We install any PXI modules you select in your chassis, any memory upgrades and application software you select, and any required driver software. We test your system and verify that its configuration meets your specifications. You can add Installation Services to any PXI order containing an embedded controller or a PXI chassis.

## Slide 38: PXI Configurator – Online!
Point
  ni.com/pxiconfig
Click
  choose chassis
  choose controller and software
  choose modules
Configure
  place modules in slots
  select accessories
  review the online price quote
  order complete systems online!
ni.com/pxiconfig
Notes: PXI Configurator online utility takes you through the five steps of configuring a PXI system. PXI Configurator
Customizes and orders a PXI measurement and automation system
Eliminates the need to visit 10 to 15 different online store pages to order a complete PXI system
Eliminates errors in orders
Provides a visual representation of your custom PXI system for presentations, proposals, and more
PXI Configurator leads you through a series of questions such as chassis type, processor type, processor speed, and memory. Then you select and place PXI modules into your chassis graphically. You can use PXI Configurator to print a graphical image of your system, export an image for use in other documents, or generate an online order for the system. The part numbers representing your PXI system configuration are automatically added to a shopping cart. You can continue shopping for other products in the online store or simply check out. Every unique PXI configuration is assigned an ID so you can retrieve your configuration later. In addition, because you can save your unique ID, a sales engineer can start a configuration for you and you can view it by entering your assigned ID. You can also use the ID to create a new configuration from a saved configuration by modifying which modules you choose.

## Slide 39: Summary
PXI is a complete system specification that extends CompactPCI for measurement and automation
PXI and CompactPCI are open specifications supported by numerous vendors
Hundreds of products are available today for configuring PXI/CompactPCI systems
Notes: Use PXI with National Instruments software tools to realize
Faster system performance
Lower costs
Better integration
More productive development