## Slide 1: The Basics of Switching
Jim Reimund
Design Engineer
Thurs Aug 17
12:00-1:15 p.m., 3:30-4:45 p.m.
San Jacinto (5B)
Notes: Choosing the right switching topology is an essential step when developing production test systems. This session reviews signal routing fundamentals and basic configurations using off-the-shelf switching multiplexers, scanners, and matrices. We will explore signal routing challenges that confront computer-based test systems such as PXI, and example configurations are presented that simplify development, increase measurement accuracy, and optimize test performance.

## Slide 2: Agenda
Why are switches important
Basic switch configurations
  Discussions
  Examples
Signal connection considerations
Available switching solutions
Notes: We start with a basic discussion of relays and their associated terminology. Then, we’ll look some example switching applications. The discussion then moves into signal connection considerations. In closing, we’ll look at National Instruments and third party switch solutions.

## Slide 3: Importance of Switching
Every ATE (automated test equipment) needs switching
  One scope, 20 channels -> switches
  One DMM, 200 RTDs -> switches
Switches help keep firm control on cost
Notes: In every ATE the switching section allows the user to get a very high number of channels without having to invest more money in purchasing more instruments.
The use of switches is so common that the switching part of the system can be a major percentage of the overall system cost, which is why it is important to evaluate switching solutions with great attention.

## Slide 4: Basic Configurations
General purpose switch
  Single pole single throw
  Single pole double throw
Multiplexers and scanners
Cascades and trees
Matrices
Notes: A general purpose switch card consists of a group of independent relays. capable of switching high (relative to the multiplexer or matrix) current loads. Examples of use are in power switching (connecting power to a DUT) and load switching.
National Instruments Products in this category are:
Form A – NI 2565; FP-RLY-420; SSR-OAC; SSR-ODC
Form C – ER-8; ER-16; SCXI-1160; SCXI-1161; PXI-2565; FP-RLY-422
A multiplexer is a set of switches that reduce multiple inputs to a single output (or expand a single input to multiple outputs). A scanner is a multiplexer that can select only one input (output) to route to the lone output (input) at a time. A multiplexer allows multiple relays to be closed at once. Finally, trees are built by having a group of multiplexers feeding a single multiplexer.
A Matrix can be viewed as a group of multiplexers, where any of the inputs can be routed to any output (or, any other input).
National Instruments products in these categories are:
  Combination – NI 2501; NI 2503; SCXI-1127
  Multiplexer Only – AMUX-64T; SCXI-1163R; NI 2590 for PXI; SCXI-1190; SCXI-1191

## Slide 5: The Ideal Relay
Resistance of 0 when closed
When opened
  Infinite resistance
  No capacitive coupling
  No inductance
Isolation
  Complete from all other switches
  Complete from itself
Notes: Before moving into switch card discussion, we need to present the ideal relay. This model will be used in the remaining discussion. An ideal relay is identical to the ideal switch. No contact resistance when on and infinite contact resistance when off. The models presented are good for f < 1 MHz.

## Slide 6: Relay Equivalent Circuit
Coil
Rcc
Rcc
Ccc
Ccc
Ccl
Vc
Rc
Rcl
IN
OUT
Notes: Typical values:
Rcc 1010 – the coil to contact leakage resistance
Ccc 1pF – the coil to contact leakage capacitance
When Closed:
Rc 1 – Contact (on) resistance. Can vary greatly. For example, a FET switch could have 2k  on resistance.
Vc 30V – a thermally induced contact potential (due to dissimilar materials in the contacts). Thermal EMF on SCXI™-1127, NI 2503 is ~4 V.
When Opened:
Ccl 2pF – leakage capacitance across the contacts.
Rcl 109 – leakage resistance across the contacts.

## Slide 7: General Purpose Relays – Formats
Form A
  Single pole single throw
  Normally open
Form B
  Single pole single throw
  Normally closed
Notes: Before discussing a general purpose switch card, let’s review some relay terminology. We’ll start with various Contact Forms.
The Form A or Form B switches are single pole single throw – (SPST) switch types. The differentiation is in the rest (no power applied to the coil) state. In this state, Form A switches are open. Form B switches are closed at rest.
Double pole single throw (DPST) relays are known as “2 Form A” or “2 Form B”.
Contact Forms – identifies the contact mechanism and the number of contacts.
Form A – SPST, normally open; also referred to as normally open contacts or make contacts
Form B – SPST, normally closed; also referred to as normally closed contacts or break contacts
Pole – number of common terminals in a switch
Throw – number of positions available per pole

## Slide 8: General Purpose Relays – Formats
Form C
  Single pole double throw
  Break before make contacts
Form D
  Single pole double throw
  Make before break contacts
Notes: Moving on to the double pole formats, we’ll look at Form C and Form D switches.
Form C or Form D switch operation is based on a single pole double throw (SPDT) switch. The operational difference is the Form C switch opens the connection before it closes the other connection (this is also known as break before make – BBM). The Form D switch closes both connections before opening the original connection (this is also known as make before break – MBB).
Double pole double throw (DPDT) switches are known as “2 Form C” or “2 Form D”
Form C – SPDT, BBM – also known as transfer contacts or changeover contacts (Form CO)
Form D – SPDT, MBB
BBM – the normally closed contacts open before the normally open contacts close (and vice-versa).
MBB – the normally open contacts close before the normally closed contacts open (and vice-versa).

## Slide 9: General Purpose Switch
Com 1
Com 2
Com 3
Com 4
Com 5
NO 1
NO 2
NO 3
NO 4
NO 5
They are used to let current pass or to provide a voltage level from point A to point B
(NO= normally open)
Notes: A general purpose switch card is composed by a series of form A or form C switches (form A shown in slide)
In this way, the user decides which circuits to close and which “A” input connect at any given moment with the correspondent “B” terminal.
The working mechanism in this case is fairly simple – the user issues a command to open or close the single switches to accommodate his needs.
You can use this architecture to allow a current to flow through a circuit or to “copy” a voltage from the input to the output.

## Slide 10: Some Other Formats
2 Form A/B
2 Form C/D
Sometimes the user needs to switch more that one signal at the same time
  2-wire and 4-wire switch modules
  differential measurements, 4- wire measurements…
Notes: 2 Form A – DPSP One operation closes or opens both contacts
2 Form C – DPDT

## Slide 11: General Purpose Switch Cards
Usually populated with form A or form C relays
Discrete, independent switching
Applications
  Power to DUT
  Loads
  Sharing a current shunt (form D)
Notes: General purpose switch cards are used in ATE systems for a variety of generic tasks.
You may use form A switches to apply power to a device under test (DUT) or to control higher power relays that are external to the instrument chassis.
You may use form C switches to share a load between two locations or apply the same load to two different pins.
You may use form D switches to switch a current shunt between multiple power lines.

## Slide 12: Example – General Purpose Switches
+5V Supply
-5V Supply
+5V UUT
-5V UUT
0.01
DMM
K1a
K2a
K2b
K1b
K1 Control
K2 Control
Notes: In this example, we measure the voltage across a shunt resistor to determine the current draw from the power supply. If power is to be applied to the DUT continuously during the test you should use, Form D switches.
Current is simply computed from Ohm’s law. At a full 10 A, the voltage across the resistor is 0.1 V. You must be careful when selecting a shunt for two reasons.
1) Voltage drop. If the shunt resistor is too large, it will drop enough voltage to prevent the UUT from receiving its operating voltage. For example, if the shunt resistor is 0.5 , a full 5 V will be dropped across the resistor.
2) Power rating. Power is determined from (I^2)*R. In the above example, the power dissipated by the resistor is 1 W.

## Slide 13: Some Terminology
Number of wires switched (1-wire, 2-wire…)
Latching and non-latching relays
Shielding
Wet and dry switching
Hot and cold switching
Notes: The number of wires switched also refers to the number of poles switched. For example, a two wire switch is a double pole (2 Form A or 2 Form C).
Shielding – The case of a relay may be designed as a signal shield. This protects the signal inside the relay, especially when switched, from interfering with external circuitry (emission). It also eliminates susceptibility. Shielding may also exist between the contact and the coil.
Single Side Stable – Relay that turns on when energized and off when de-energized. A non-latching relay.
1 Coil Latching – Relay with a latching construction that can maintain the on or off state with a pulse input. With one coil, the relay is set or reset by applying signals of opposite polarities.
2 Coil Latching – Relay with a latching construction composed of two coils a set coil and reset coil. The relay is set or reset by alternately applying pulse signals of the same polarity.
Wet Switching – The relay contacts are covered with a thin film of mercury. Design for hot switching large currents or voltages.
Dry Switching – Switching at low voltage and current levels.
Hot Switching – Contacts make or break current.
Cold Switching – Contacts do not make or break current. Close relay contacts BEFORE applying the signal. Open AFTER removing the signal.

## Slide 14: Multiplexer
Connect one “source” to multiple test points (1:N)
Connect multiple test points to one “measure” (N:1)
May close more than one connection at a time
Operation may be non-sequential
Source
DUT Pin
DUT Pin
DUT Pin
Notes: Getting back to the other switch card topology, let’s look at a multiplexer.
You can use this topology to connect a signal source or an instrument to multiple test points, either multiple DUTs or multiple points on a single DUT. You can synchronize source and measure connections using a pair of multiplexers. One multiplexer would be configured as 1:N, the other as N:1.
Channel – The switched signal route through a multiplexer
Example of Usage – Scope measures four different signals one at a time. Function generator sourcing a signal to different points.

## Slide 15: Scanner/Multiplexer
They are used to route one signal from several sources to one common terminal
1B
2B
3B
4B
5B
Common
Notes: A multiplexer is built from general purpose (Form A or B) relays – one terminal of the switches is connected together, so they have one side independent, while the other side is common to all the devices.
The user can connect the common terminal to any output point by closing the correct switch. This architecture is very commonly used in automated test equipment (ATE), where normally the user has, for instance, one scope and many channel to acquire or one arbitrary waveform generator and many terminals where to distribute the signal.
In many cases (SCXI-1127, NI 2501 for PXI, NI 2503 for PXI), the user can download a scan sequence on the switch module. So the user will be able to download something like “common connected to 1B, then 3B, then 2B etc”, then the operation can be initiated and executed with or without exchanging handshaking signals with an external measurement device such a DMM.

## Slide 16: Scanner
Identical to a multiplexer
Except:
  Sequential signal connections only
  Only one signal selected at a time
Source
DUT Pin
DUT Pin
DUT Pin
Notes: Think of a scanner as a multi-position selector switch. Only one signal path can be connected (unlike a multiplexer) at a time, and these connections must be sequential (random access is not allowed).
Note, these are usually available as MBB (shorting) or BBM (non-shorting) versions. In addition, there may always be one connection (you can’t open all relays).

## Slide 17: Example – Scanning with the NI 4060
SCXI
1140
Mux
SCXI
1140
Mux
SCXI Slot 0
High Voltage Analog Buss
Trigger and Control Buss
NI 4060 5.5 Digit DMM
SCXI Digital Buss
SCXI-1127 Multiplexers
Notes: The NI 4060 is a 5.5 digit DMM. The SCXI-1127 is a 32:1 2-wire scanning multiplexer (with other modes available).
In this configuration, the pair of SCXI-1127 acts as a single 64:1 scanning multiplexer. The selected signal is routed through the scanner and places on the high voltage analog bus (4 2-wire pairs, only one is used). This bus is connected to the input of the DMM, where the measurement is made. When the measurement is complete, the DMM sends a signal (SCAN_ADVANCE) to the SCXI-1127s and they advance to the next channel in the scan list. The DMM waits for the switches to settle and then performs the measurement on the new channel. This process continues through the scan list.

## Slide 18: Tree
Built from cascaded multiplexers
Requires more relays than a multiplexer
Advantages
  Signal isolation
  Reduces
    Signal coupling
    DC leakage
Source
DUT Pin
DUT Pin
DUT Pin
DUT Pin
DUT Pin
DUT Pin
DUT Pin
DUT Pin
DUT Pin
Notes: In a voltage source application with a large output distribution, a tree may be desirable over a multiplexer. This is because open channels are equivalent to a 1 G resistor. Many of these in parallel can create significant leakage currents (we’ll review the relay model later in the presentation).

## Slide 19: Binary Tree
Source
DUT Pin
DUT Pin
DUT Pin
DUT Pin
Notes: A special version of the tree, a binary tree, is built from Form C switches. The binary tree is primarily used for high frequency applications.

## Slide 20: Matrix
8 Rows
4 Columns
They are used to connect any signal from any row to any column
Example: DUT Connected to Rows, Instruments to Columns
Single Switch
Single Scanner/MUX
Notes:   To create the matrix, we start from several multiplexers and we connect the free side of the switches together.
  In this way we obtain a grid that allows us to connect any row to any column.
  This architecture is very flexible and is widely used in ATE to connect several instruments (scopes, DMMs, arbitrary waveform generators) to the different test point on the unit under test.

## Slide 21: Sparse Matrix
Built from multiplexers
Reduced relay count
Signal connection limitations
DUT Pin 1
DUT Pin 2
DUT Pin 3
Source 1
Source 2
Source 3
Notes: A sparse matrix saves money in implementation and removes a lot of flexibility from the end application.
You need to carefully plan the switch to DUT interface so there are no signal conflicts that may make a test impossible. This planning also lengthens the test development cycle (construction on a fixture can’t begin until the switching interconnects are defined). Finally, there is a risk that future products can not be tested with this architecture.

## Slide 22: Single Wire Switching
Only one side is switched
The other side is common to all channels
Applications
  Voltage (referenced to ground)
  Current (single ended measurements)
  High frequency
Signal 1
Signal 2
Signal 3
GND
Notes: Single wire (Form A is most common) switching is used to switch the high side of the signal. The low side of all signals are common on the switch card. Typical applications are single ended voltage and current measurements. A single wire switch is also used in high frequency switching where low impedance returns and/or coaxial cabling are required. This has an advantage over two wire high frequency switching because the characteristic impedance of the switch changes when both high and low are switched.
If you are switching a signal that must be terminated when not in use, a Form C switch would be better. The signal would enter at the common pin, the Normally Closed contact would be connected to the termination resistor, and the Normally Open contact would connect to the destination.
You can build matrices and multiplexers from single wire switches.

## Slide 23: Two Wire Switching
Both high and low sides are switched
Increases common mode noise rejection
Applications
  Differential measurements
  Low voltage
  Resistance
Notes: In a two wire switch, both the high and low sides are switched. This provides for great common mode noise rejection, as you are now able to switch a differential signal. Two wire switching is typically done with either a 2 Form A or a 2 Form C switch.
We’ll examine the low voltage and high current applications later. A two wire switch is fine for measuring resistance between 100  and 10 M. Different switching is required for values outside of this range.

## Slide 24: Three Wire Switching
Switches high, low, and a guard signal
Guard shunts noise and leakage currents
Applications
  High impedance voltage measurements
  Large resistance (> 10 M) measurements
  Low leakage
Notes: Three wire switching is used in applications where isolation must be maximized. Rather than using very expensive relays, a three wire configuration is used to switch. The third wire switched is the guard signal. This signal is routed through the entire system to reduce the leakage and capacitance effects.

## Slide 25: Large Resistance Guard Circuit
R
x
M
1
R
l
R
g
Notes: This is a guarded resistance measurement. Rx is the resistor to be measured and its value is 50 M. There is leakage between the HI test lead and the LO lead, represented by Rl (1 G). Without the guard circuit, Rx and Rl would be in parallel. The measured value without a guard would be 47.6 M, an error of almost 5 percent! By adding the guard circuit (a low impedance source) connected to the cable shield, the voltage of the shield is the same level as the voltage being measured, so no current flows from the signal being measured, all leakage currents are sourced by the guard circuit.

## Slide 26: Four Wire Switching
Switches
  Measure high and low
  Sense or excitation high and low
Applications
  Kelvin measurements
  Strain gauges
Notes: A Kelvin measurement is a four wire resistance measurement. Two of the four wires are used to source the current. The other two are used to sense the resistance.

## Slide 27: Applications of Switching
Voltage switching
  Low voltage (< 100 mV)
  High voltage (>100 V)
Current switching
  Low current ( 1A)
  High current ( 1 A)
Switching for resistance measurements
  Low resistance ( 100)
  High resistance ( 10M)

## Slide 28: Low Voltage (< 100 mV) DC Switching
Contact potential
  Due to dissimilar materials
  Use two wire switching to reduce the effect
Contact resistance
  Increases due to formation of contamination with use
V
Notes: One of the key relay characteristics that affects a low voltage signal is the contact potential in the relay. In a copper-to-copper contact, the thermal EMF is less than 0.2 V/C and oxidized contact can have a thermal EMF of up to 1mV/C. Use a symmetrical switching approach (2-wire) as shown in the diagram above. The thermal EMF factors, while not completely canceling each other will be greatly reduced. National Instruments uses relays with a very low thermal EMF in its SCXI-1127 and NI 2503 switch card products.
Things to keep in mind:
Minimize the number of switches in the signal path
Use copper connections
Allow temperatures (ambient/system) to stabilize
Minimize temperature difference between leads (2 Wire)
Try to dry switch or switch very low powers
Measure as soon as possible after closing the relay (prevents heat generated by the coil from affecting the measurement)
Measure ground or a shorted channel to establish a reference voltage

## Slide 29: High Voltage (> 100 V) DC Switching
Switching capacitive loads
  Initial “turn on” load is a short
  Reduced relay life expectancy
Switching inductive loads
  Flyback voltage causes arcing
  Reduced relay life expectancy
Shielding
Notes: An uncharged capacitor is a short, when switching a capacitive load the only current limiting that is available is the contact resistance (switch closure issue). When a switch is opened, especially with an inductive load, there is contact arcing. A relay’s life will be greatly diminished switching under these conditions. Usually a relays maximum switching capability is specified with a resistive load. This specification must be derated for inductive and capacitive loads. Fusing the input side (assuming the load is on the output side) of the relay is a good idea.
When switching an inductive load, you may want to examine adding a diode or back to back zener diodes to clamp the flyback voltage.
Use shielded relays (and cables) to prevent contact noise from interfering with external circuitry. Both the case and the coil should be shielded.

## Slide 30: General Current Considerations
General current 1 A to 100 mA
Notes: A current source should be switched with a Form D relay configuration (MBB). The source is tied to the common, the normally open contact to the target, and the normally closed contact is shorted to the return. The operation is desired to prevent transients when switching. Note this configuration is only required if the compliance voltage of the source exceeds the maximum voltage rating of the switch. You can avoid the MBB operation if the source allows a standby mode (entered while switching). Because Form D relay cards are not very common, a Form D relay is usually built from a pair of Form A relays (switched separately).
Typically, on the source (high) side is switched.
Switching currents in a system to a meter to measure them could be tricky because the meter must be placed in series with the current. Recommended alternatives are to use a shunt resistor and measure the voltage drop. Make sure the voltage drop across the shunt resistor won’t affect normal operation of the DUT. For AC currents, a current probe with a oscilloscope is an option. One other thing worth mentioning is a LEM sensor (lem.com).

## Slide 31: Current Switching
Low current (< 1 A )
  Offset current generated by the switch card
  High isolation
    Channel to channel
    Contact to coil
  Shielding
High current (> 1 A)
  Contact resistance
  Shielding
Notes: Offset current will likely be the largest source of error. This current is generated by the switch card (due to coil to contact impedance, piezoelectric sources, and so on). By having a high degree of isolation, extraneous current paths are reduced or eliminated. Shielding eliminated susceptibility to radiated noise. The mechanical operation (open or close) of the relay contacts cause a current pulse in the circuit (due to a charge transfer between the contacts) – this may affect measurement. If the current measurement is made with a shunt, minimize the input capacitance of the switch and the cables (signal settling time is related to the RC time constant).
For high current switching, the ideal situation is to dry switch the signal and then increase the current. This greatly increases switch life. Power dissipation is related to I2R, so contact resistance is a critical parameter. Shielding is recommended to reduce emissions.
When switching a current source, use a Form D switch (or two Form A switches). The source should be tied to the common, the normally open contact goes to the target, and the normally closed contact is tied to the return.
For very high current (> 5A) relays are typically outside the relay chassis and controlled by digital outputs (SCXI-1163 is a good module for this).

## Slide 32: Resistance Switching
Low resistance
  4-wire (Kelvin) switching
    One pair sources current
    One pair measures voltage
  Shielding usually not necessary
High resistance
  Shielded cables
  Relays with high off resistance
Notes: Measuring single ended resistance can be done with a single wire relay between the high side of the meter and the value to be measured. There must be a very low impedance return to the meter to prevent errors.
Measuring normal resistance levels can also be done with a two wire relay (recommended). The drawback with the two wire approach is we’ve doubled the amount of resistance we’ve added (not the DUT’s value, but contact resistance, etc.) to the measurement circuit. For very small values, a four wire measurement approach is necessary.
High resistance measurements require shielded cables to achieve stable readings, the resistor being measured should also be in a shielded enclosure.

## Slide 33: Example
NI 4060 and SCXI-1127
Resistance measurements
  Two wire
  Four wire
Notes: We are going to measure a group of low value resistors using a single wire configuration, two wire configuration, and a four wire configuration. We’ll compare the results of these measurements.
Note this example also demonstrates the random scanning capability of the SCXI-1127.

## Slide 34: NI Switching Solutions
General purpose switches
  Form A – NI 2565; FP-RLY-420; SSR-OAC; SSR-ODC
  Form C – ER-8; ER-16; SCXI-1160; SCXI-1161; FP-RLY-422
Multiplexers and matrices
  Combination: NI 2501; NI 2503; SCXI-1127
  Multiplexer Only: AMUX-64T; NI 2590; SCXI-1163R; SCXI-1190 (1.3GHz) (new!); SCXI-1191 (4GHz) (new!)
Notes: True switching only. Other products with switching (multiplexers) are the SCXI-1100, SCXI-1102, SCXI-1104, SCXI-1122 and VXI-SC-1102.
Descriptions:
NI-2565: PXI – non-latching; Form A; 16 channels
FP-RLY-420 – FieldPoint; non-latching; Form A; 8 channels
FP-RLY-422 – FieldPoint; non-latching; Form C; 4 channels
SSR-O(A|D)C – non-latching; Form A; Discrete
ER-8, ER-16 – Standalone (control via digital card or multifunction DAQ board); non-latching; Form A; 8, 16 or 32 (combined) channels
SCXI-1160 – latching; Form C; 16 channels
SCXI-1161 – non-latching; Form C; 8 channels
NI-2501, NI 2503: PXI – non-latching; 1, 2, or 4 wire; matrix with TB-2606
SCXI-1127 – non-latching; 1, 2, or 4 wire; matrix with SCXI-1332
AMUX-64T – Standalone (MIO control); non-latching; 1 or 2 wire (up to 256:1 one-wire).
NI-2590 – 4:1, 1.3GHz Multiplexer
SCXI-1163R – Use as eight 4:1 Multiplexers
SCXI-1190 – quad 4:1 RF muxes (1.3 GHz)
SCXI 1191 – quad 4:1 RF muxes (4 GHz)