## Slide 1: Telecom Audio Testing
Gerardo Leal
DSA Engineer
Thurs Aug 17
12:00-1:15 p.m.
Lavaca (6B)
Notes: Learn how to perform audio testing on telecom equipment such as cell phones using the swept sine mode of the NI-4551 Dynamic Signal Analyzer (DSA) from National Instruments.

## Slide 2: Audio Tester Application
The goal
Perform single or multiple point frequency response measurements to obtain important characteristics of a unit under test (UUT)
Today’s application
Production line audio testing of phones, cell phones, and other telecommunications equipment
Notes: This presentation describes a basic setup to perform audio testing on cell phones. We cover the hardware and software used, as well as the advantages of using a computer based instrument to perform such a test.
We also cover the basics on swept sine, which is the technique used to perform the frequency response measurements that we need to characterize the units under test (UUT).
A Frequency Response measurement is used to obtain the transfer function of the UUT, which can be a speaker, an electronic circuit , a cell phone, a regular phone, or any system that needs to be analyzed on the frequency domain. In this presentation, we cover the audio testing of cell phones on production lines.

## Slide 3: Highlights
Great measurement quality
Test time – dramatic reduction per unit (PCI/PXI bus)
Flexibility – a single DSA board can replace - FFT analyzer + octave analyzer + swept sine analyzer (SINAD, distortion)
Integration – can work with other NI instruments (DMMs, scopes, arbs) and other products (FieldPoint™, IMAQ™...)
Floor plant space – multiple instruments in a single and compact PXI chassis or PC
Cost – lower compared to traditional instruments with same capabilities
Development tools – Measurement Studio™, LabVIEW™
Notes: Highlights
The NI Dynamic Signal Analyzer (DSA) computer-based instruments hardware and software can deliver high dynamic range measurements capabilities. They offer a high measurement quality comparable with any desktop traditional instrument but offering a lot of other advantages as well.
By leveraging the speed and flexibility of the PCI or PXI computer buses the test time in production lines can be dramatically reduced compared with the time required by a traditional GPIB controlled instrument to perform the same measurements.
The NI DSA offers the flexibility to perform a variety of measurements with interchangeable software running on the same hardware. As an example, the NI-4551 DSA instrument can behave as a real time FFT analyzer, octave analyzer of swept sine analyzer allowing real time measurements of different types to be performed on the same UUT. The measurement flexibility does not stop there, since by having other types of instruments on the same computer, a complete test station can be replaced by a single computer with scopes, generators, RF testers, DSA analyzers, image analyzers etc. All the systems can be integrated into a single test software and by using a single PXI chassis, for instance, a lot of floor space on the production line can be saved.

## Slide 4: System Configuration (PCI)
1. Hardware –
  Real time dynamic signal analyzer (NI-4551)
  Cable (SHC68-68C-A1)
  Connector block (BNC-2140)
2. Software –
  Measurement Studio or LabVIEW
  NI-DSA
3. Test Unit –
Notes: These are the system requirements for the setup of the audio tester.
Software
A computer with Windows 95/98 or NT.
NI-DAQ 6.6 or later installed. (Generic hardware driver.)
NI-DSA 1.1 or later installed. (DSA instrument driver.)
LabVIEW 5.1 or LabWindows/CVI 5.0.1 or later versions.
The NI-DSA swept sine audio tester library or project.
Hardware
An NI-4551 DSA instrument board installed and working.
A BNC-2140 or BNC-2142 connector block .
An SHC68-C68-A1 cable connecting the BNC block and the board.

## Slide 5: The Hardware (NI-4551)
High fidelity dynamic signal analyzer
  95 dB dynamic range
  95 kHz alias free span16-bit
  Delta/Sigma, 128 times over-sampling
  Simultaneous sampling (I/O) (2 inputs/2 outputs)
  AC/DC coupling
  Input range (10 mV;+-42 V) (output 0.1,+-10 V)
  Over-range detection (analog and digital)
  Calibration
  Synchronization bus (RTSI™)
  Integrated circuit piezoelectric (ICP®) conditioning
  Real-time DSP on board processing
Notes: Dynamic signal acquisition hardware from National Instruments is designed to make the same quality of measurements you expect to find in a stand-alone instrument but at a fraction of the cost. Our DSA products leverage quality board design and manufacturing to make audio and vibration measurements with a 95 dB dynamic range. Because of the nature of the Delta/Sigma ADCs, the products are alias free, which is very important in sound and vibration applications. The range of signal attenuation/gains combined with the extended input range reduces the need for expensive signal conditioning devices. Over-range detection lets you know when you have exceeded the instrument’s input limits and lost data, allowing you to adjust the gains on the fly to get the proper voltage input range. If you need to synchronize multiple instruments to the same sample clock, or if you want to share trigger signal, you can use the real-time system integration (RTSI). ICP (Integrated Circuit Piezoelectric) signal conditioning is included in our BNC-2140 connector and provides a constant-current source to power internally amplified sensors without the need for an external excitation power supply. The major advantage of the ICP accelerometer is that it is simple and easy to use due to the built-in microelectronics.
National Instruments analysis software for sound and vibration applications offers a wide variety of real time on board and on host processing capabilities such as FFT analysis, averaging, sweep sine analysis, fractional octave analysis and much more.

## Slide 6: The Analysis Software (NI-DSA)
Measurement level real time instrument driver
- Swept sine analyzer mode
- FFT analyzer mode
- Octave analyzer mode
Notes: The NI-4551 and NI-4552 Instruments have two main personalities, NI-DAQ compatibility and NI-DSA instrument mode. The Instrument mode performs analysis in real time on the on board processing. There are three main instrument modes implemented so far for the DSA instrument – FFT Analyzer mode*, Sweep Sine Analyzer mode*, and Octave Analyzer Mode*.
* The DSA FFT Analyzer Engine runs in the on board processor and provides four real time dual channel analyzers as a powerful and flexible analysis tool.
* Octave Analysis – In many cases, when dealing with sound measurement and analysis, the final “customer” is the human ear. And, like most human senses, the ear exhibits a response based on a logarithmic scale for both the level and the frequency. So, to produce results that are somewhere related to this human perception, sound levels are expressed in decibels and frequency content measured with a logarithmic scale. A lot of research efforts are currently focused on this field of psycho-acoustics, but octave band analysis remains a technique of prime choice. Octave band analysis is performed through a parallel bank of bandpass filters. The real-time output of each filter is then averaged to compute the power in each band and displayed as a bar graph.
* A swept sine measurement is a sine sweep that steps through a specified sequence of frequency points. At each frequency, the source tone is maintained and the inputs to the analyzer are measured only at that source frequency. Both the magnitude and phase of each component is measured so that the frequency response of input/output systems can be characterized.

## Slide 7: Phone Audio Tester Application
Use the NI-4551 in its swept sine mode to measure RMS, frequency response, sinad, THD, THD+noise, harmonic amplitudes and noise on cellular phones
Example tests –
- Feedback test (speaker and mic)
- Input test (mic)
- Buzzer test (buzzer)
Notes: A common application of the swept sine capabilities of the NI DSA instruments is phone audio testing. The phone is tested on the frequency domain to obtain the frequency response or transfer function of its microphone, speaker and buzzer for instance. Several other tests can be performed but these are some of the more common ones.
The phone will be submitted to a single frequency tone sweeping a defined range or frequency points though an artificial mouth or high fidelity speaker. The electric output of the microphone or feedback audio of the cell phone speaker will be analyzed at each frequency point to obtain amplitude and phase variations as well as harmonic distortion and signal to noise characteristics. All this determines the quality of the microphone and speaker components on the phone.
Other tests such as the buzzer test will not take any excitation from the instrument. They basically measure the frequency domain characteristics of the output generated by the phone buzzer.
These are just a couple of possible tests that can be done to cell phones, several other audio tests such as octave analysis can be performed to cell phones and other audio equipment using similar setup, hardware and software.

## Slide 8: What is Swept Sine?
Measures the frequency response/transfer function of a device, a single frequency at the time, on each step of the sweep
Provides methods for optimization of the measurements performed
at each frequency point offering higher dynamic range than broad
band methods such as FFT
Notes:     Why swept sine analysis for transfer function measurements ?
Various techniques are used to perform a frequency response measurement to calculate transfer functions of systems under test. Obtaining the transfer function of a system using FFT based measurements has the advantage of speed. The whole frequency response over the desired frequency span is calculated simultaneously. However, to do this it is necessary to generate a stimulus signal with energy in all the signals within the span. To acquire the response, a large input range is set to acquire the larger peaks or components on the frequency contents of the signal. For signals with large variations over the test span the dynamic range achieved by an FFT method is very low. Swept sine measurements provide a method of optimization of the measurements performed at each frequency point, extending their dynamic range beyond 140 dB (a much better dynamic rage). Some applications require a single tone stimulus per measurement (such as the one provided by the swept sine method) instead of the one used in broadband methods based on chirp or random excitations with energy all over the spectrum and a single measurement. Swept sine analysis is the most accurate technique for transfer function calculation of devices or systems subject to non-linearities and/or very noisy environments. By exciting just one frequency band at the time within the whole test span, this method offers a better measurement control and accuracy for characterization of some systems than an FFT-based method. The per point optimizations available on the swept sine technique improve dynamic range, speed, and much more. For instance, the source amplitude can be optimized and changed depending on the frequency measured to compensate or maintain response levels at critical spans. An other example is the auto resolution feature; this feature will skip measurement points in regions where the response is not changing significantly, therefore speeding the overall sweep time. Other optimizations such as auto ranging automatically detect the response level and change the amplification or attenuation settings on the input stages to achieve better conversion resolution and dynamic range.

## Slide 9: Basic Swept Sine Demo
Obtain transfer function of a UUT using the basic swept sine analyzer example
Notes: The basic swept sine analyzer example is a LabVIEW example also available in C (for Measurement Studio and LabWindows™/CVI) – a very simple application that shows the basics on how to create a virtual swept sine audio tester. It exposes you to a limited amount of controls and features of this mode.
We show a more complex example later – the NI-DSA Swept Sine Audio Tester. It is a simple LabVIEW or LabWindows/CVI example that uses the NI-DSA instrument driver to control an NI-4551 DSA Instrument PCI plug-in board. You can use it as a template and can customize it for any swept sine application development. This program uses the NI-4551 in its swept sine mode. The NI-455X series boards offer two additional modes of operation accessible through the NI-DSA driver – the FFT analyzer mode and the fractional octave analyzer mode. All three modes perform DSA measurements in real time on the on-board processor of the instrument.
Although this program is called swept sine audio tester, it is a generic example of a swept sine analyzer and its use is not restricted to audio applications. You can use it to perform several frequency response measurements on any system. A basic swept sine analysis application generates single tones through the output of the board, performing a sweep of specified or calculated frequencies. At the same time, the inputs acquire the system response and the stimulus sent and the on-board processor will perform several measurements at each step frequency of the sweep. Unlike an FFT measurement, the swept sine measurement measures the frequency response at a single frequency at the time on each step of the sweep. The next section explains some of the features of this particular example and some of the features of the NI-DSA swept sine mode.

## Slide 10: Phone Audio Tester Demo
Basic swept sine audio tester setup
Example test
– Feedback test
Signals go to the DSA board through the BNC-2140 connector block
Notes: Connecting Signals
Input and output connections to your NI 4551 instrument are available via the two front panel analog I/O and digital I/O connectors of the instrument.
The analog I/O port for the NI 4551 is a 68-position VHDCI 0.8 mm connector. It connects to the BNC-2140 or BNC-2142 DSA accessory through the SHC68-C68-A1 shielded cable. You can access the analog I/O of the NI 4551 using standard BNC connectors on the BNC-214X. If you are not using a BNC-214X, you can access the analog I/O lines through a standard DB-25 female connector with the SHC68-D25-A1 shielded cable.
Make sure the following connections are correct :
  The cable connecting the NI 4551 DSA instrument with the BNC 214X terminal block is correctly connected.
  The analog output channel 0 (DAC0 OUT on the BNC-214X) is connected to the input of the system to test.
  The same analog output channel 0 (DAC0 OUT) is connected through a BNC T to the analog input channel 0 (ACH0).
  The analog input channel 1 (ACH1) is connected to the output of the system to test.

## Slide 11: Phone Audio Tester Demo
Features
auto resolution
auto ranging
source ramping
source autolevel
averaging
Measurements
total RMS
FRF, cross power
THD, SINAD
stimulus, response
and harmonic levels
Notes:   Features Summary
This program offers a basic interface for interactive swept sine analysis. It performs all the following measurements in all points of the sweep – spectrum or level of stimulus and response at the specific frequency, cross spectrum, transfer function or frequency response, total harmonic distortion, total harmonic distortion + noise, sinad, harmonics spectrum, total response power, and noise.
It offers settling and integration averaging settings defined in either cycles or time units. With the program you define the sweep frequencies in three different ways – by defining start frequency, stop frequencies, and number of steps (in either linear or log increments); by setting threshold parameters and maximum number of steps to skip for auto resolution mode; and finally, by defining each of the desired test frequencies on the sweep.
You can also specify the initial voltage range and auto range enabling or disabling in a per channel basis. You can set the source amplitude and ramping enable with the specified ramping slope. You can also enable source auto level locked on either input channel and setup from the front panel of this program. All this can dramatically increase the dynamic range of the measurement to more than 140d#B and the signal to noise ratio.
With the program you can specify the number of times a swept sine test is performed; and if you enable data logging, all results are written to a file. The user can specify the file and the measurements he/she wants to log to disk. The program generates an excel file that can also be opened in a text editor. The file contains the date and time of each test and the desired measurements results per frequency of the sweep.

## Slide 12: Questions?