## Slide 1: Computer-based Test and Measurement for Semiconductor Components
Eric Mills
Software Engineer
Thurs Aug 17
1:45-3:00 p.m.
Lavaca (6B)

## Slide 2: I-V and C-V Measurements
Measurement overview
I-V measurements
C-V measurements
Notes: First, we become familiar with the reason behind the measurements, then we go into specific details about each measurement and show how it is made.

## Slide 3: Measurement Overview
I-V measurement (current vs voltage)
  Typical ranges
    +-10 Volts
    1 pA to 1A
C-V measurement (capacitance vs voltage)
  Typical ranges
    +-10 volts
    1 pF to 100 nF
    100 Hz to 1 Mhz
Notes: I-V and C-V measurements have been used extensively for the past 35 years.
The term “I-V measurement” refers to the current versus voltage characteristics of a component.
The term “C-V measurement” refers to the capacitance versus voltage characteristics of a component.

## Slide 4: I-V Measurement
Notes: The I-V measurement has many different uses. At a glance, it provides enough information to determine if the component has serious flaws. Upon closer inspection, this measurement answers questions like:
At what temperatures is the V/I relationship of this resistor linear?
What is the breakdown region of this zener diode?
What is the maximum current that can move through the source of this MOSFET with a 1 V bias on the gate?

## Slide 5: C-V Measurement
Notes: The C-V measurement is possibly the most common test in the semiconductor industry. This measurement provides a wealth of information to the user about p-n or schottky junctions, and metal oxide semiconductor (MOS) and metal insulated semiconductor (MIS) devices. This information is highly correlated with the performance parameters and reliability of the devices under test.

## Slide 6: Questions
Who makes these measurements?
Why are these measurements made?
How are these measurements made?
Notes: These key questions provide a basis for understanding the overall picture. By knowing the answers to these questions, we can further our understanding and communicate intelligently with our customers.

## Slide 7: Who?
Semiconductor industry
  Fabrication plants
  Cleanrooms
Research facilities
  Characterization labs
  Cleanrooms
Educational institutions
  Electronics labs
  Physics labs
Notes: I-V and C-V measurements are made by many various types of customers. This market is much larger than just semiconductor fabrication facilities. Many engineering and physics research facilities use these these systems regularly to further measurement theory and component theory. Almost all educational facilities use these systems as a tool to further understanding and research.

## Slide 8: Why?
Component characterization
Wafer testing
Doping research
Quality assurance
Reliability studies
Notes: The applications of these tests are almost limitless. It may seem strange that something seemingly insignificant like the slope of a curve can lead to such incredible conclusions, but the interpretation of the test results is the key to relating this information to something useful.

## Slide 9: How?
Traditional measurements
  HP 4140 series
  Techtronics curve tracer
  Extremely high cost
  Little flexibility
A better solution
  National Instruments hardware
  National Instruments software
Notes: Traditionally, these measurements have been dominated by large rack-mounted measurement systems. Popular instruments were the HP 4140 series systems, and the early Tektronix I-V curve tracers. These instruments came with a $50,000-250,000 price tag. The output was usually to a four inch monochrome display and 3.5” floppy disk. The measurement parameters were not nearly as flexible as a PC-based solution. The measurement industry sold these high cost systems year after year while providing only small upgrades, but still raising the price with each new version. We believe that there should be an alternative. Today, it is common to see LabVIEW used to automate all these expensive systems. With the combination of National Instruments hardware and software, a PC-based measurement system could be built that could compete with many of the traditional measurement systems.

## Slide 10: I-V / C-V Startup Kit Hardware
Notes: This is the National Instruments hardware for use with the I-V/C-V Startup Kit.

## Slide 11: Measurement Requirements
I-V measurement
  Two analog outputs
  One analog input
  16 bit resolution
C-V measurement
  333 kS/s update rate and sampling frequency
  One analog output
  One analog input
  16 bit resolution
Notes: The hardware chosen for the I-V/C-V Startup Kit had to meet the above requirements. Future hardware improvements are expected to increase the accuracy and range of the specifications.

## Slide 12: I-V/C-V Startup Kit Hardware
PXI-1000B chassis
PXI 8156B controller card
PXI 6052E MIO DAQ card
Fixtures
Notes: This is the recommended hardware configuration for use with the I-V/C-V Startup Kit.

## Slide 13: PXI-1000B Chassis
8 Slot backplane
High output power supply
Rugged design
Portable
Small footprint
Notes: The PXI-1000 chassis combines a high-performance 8-slot PXI-compatible backplane with a high-output power supply, integrated cooling, and a compact structural design optimized for maximum usability in a wide range of applications. The modular design ensures the highest level of maintainability, resulting in a very low mean-time-to-repair (MTTR). The PXI-1000 complies fully with the PXI Specification, Revision 1.0, offering advanced timing and synchronization features. Because of the interoperability with CompactPCI defined by the PXI Specification, the PXI-1000 chassis can accept both 3U PXI and 3U CompactPCI modules.

## Slide 14: PXI 8156B Controller Card
Controls up to seven PXI modules
333 Mhz CPU
6.4 GB hard drive
Ethernet connector
GPIB connector
32 MB RAM
Notes: The PXI-8150B Series embedded computers are high-performance system controllers compatible with CompactPCI and with the PXI (PCI eXtensions for Instrumentation) specification. PXI defines a compact, rugged PC platform for modular instrumentation. The industrial packaging of CompactPCI is ideal for harsh environments. In addition, the high level of expandability makes it desirable for a wide range of applications.

## Slide 15: PXI 6052E MIO DAQ Card
Two output channels
16 input channels
+/-10 V output
333 Ks/s sampling rate
333 Ks/s output rate
16-bit input and output resolution
Notes: The PXI-6052E uses E Series technology to deliver high-performance and reliable data acquisition (DAQ) capabilities to meet a wide range of application requirements. You get up to 333 kS/s, 16-bit performance on 16 singled-end analog inputs. Depending on your type of hard drive, this PXI DAQ module can stream to disk at rates up to 333 kS/s. The PXI-6052E E Series module features analog and digital triggering capability, as well as two 24-bit, 20 MHz counter/timers, and 8 digital I/O lines. The PXI-6052 also features two 16-bit analog outputs.

## Slide 16: Fixtures
BNC 2110 terminal block
NI SH68-68-EP cable
Demo
  BNC/minigrabber cables
Probe station
  BNC/triaxial converter
Notes: The BNC-2110 device is a shielded connector block with signal-labeled BNC connectors. You can use these connector blocks with E-series multifunction I/O DAQ devices and 671x analog output devices. See the PDF catalog page or compatibility table for more detailed compatibility information.
Extended performance version of the SH6868 cable assembly; features individually-shielded analog twisted pairs for reduced crosstalk with high-speed boards. This cable works with all 68-pin E Series or 67xx devices. It does not work with 435x devices.
For demonstration purposes, we use BNC-to-minigrabber cables to connect from the BNC 2110 Terminal Block to the test circuit built on a breadboard. A probe station typically uses triaxial cables to connect to the probes. These cables provide far more noise protection than a typical BNC cable. For this reason, a converter is needed to connect the BNC 2110 Terminal Block to the probe station.

## Slide 17: National Instruments Software
I-V/C-V startup kit (available upon request)
LabVIEW™ Full Development System
NI-DAQ™ 6.6
Notes: The I-V/C-V software is a Startup Kit designed to make basic I-V and C-V measurements.
LabVIEW is a highly productive graphical programming environment that combines easy-to-use graphical development with the flexibility of a powerful programming language. It offers an intuitive environment, tightly integrated with measurement hardware, for engineers and scientists to quickly produce solutions for data acquisition, data analysis, and data presentation. Requires LabVIEW 5.1 or greater.
Requires NI-DAQ 6.6 or greater.

## Slide 18: I-V Measurement
Two terminal components
  Resistors
  Diodes
  Wafers
Three terminal components
  Bipolar junction transistors
  MOSFETs
  JFETs
  Wafers
Notes: The IV measurement can be run on any two terminal or three terminal component. Some sample components are:
Two terminal components
  Resistors
  Diodes
  Wafers
Three terminal components
  Bipolar junction transistors
  MOSFETs
  JFETs
  Wafers

## Slide 19: Two Terminal I-V Circuit
Notes: This schematic shows the test circuit for the two terminal I-V measurement. There are only two components, the device under test (DUT) and the load resistor.

## Slide 20: I-V Block Diagram
Notes: This diagram shows the signal flow for the I-V measurement. The DAQ card sends a voltage into the circuit. The voltage comes back out of the circuit and into the DAQ card. This voltage level is divided by the value of the load resistor. The value of the load resistor should be known as accurately as possible for accurate measurements. From Ohm’s law, we know that voltage/resistance yields current. This current value is sent to the graph, and the I-V plot displayed.

## Slide 21: Two Terminal I-V Screenshot
Notes: This is a screenshot of the I-V/C-V Startup Kit. The DUT is a Zener diode (two terminal I-V measurement). The vertical (y) axis is current, and the horizontal (x) axis is voltage. The diode turns on at approximately .7 V, and the breakdown region is approximately -6 V. It is possible to determine these values more accurately by zooming on the graph, which is shown in the demonstration.

## Slide 22: Three Terminal I-V Circuit
Notes: This schematic shows the test circuit for the three terminal I-V measurement. The DUT is a MOSFET.

## Slide 23: Three Terminal I-V Screenshot
Notes: This is a screenshot of the I-V/C-V Startup Kit. This is an I-V measurement. The DUT is a MOSFET (three terminal I-V measurement). The vertical (y) axis is current, and the horizontal (x) axis is gate voltage. Each color represents a different voltage on the drain of the MOSFET.

## Slide 24: C-V Circuit
Notes: This schematic shows the test circuit for the C-V measurement. There are only two components – the DUT and the load resistor. This is a classic low-pass filter.

## Slide 25: C-V Bock Diagram
Notes: This diagram shows the signal flow for the C-V measurement. The DAQ card sends a voltage into the circuit. The voltage comes back out of the circuit and into the DAQ card. This voltage level is sent into a Hanning window function which cleans up the data so that the harmonics can be withdrawn.
From the window function, the signal goes into the auto power spectrum where a Fast Fourier Transform is applied. This changes the signal from the time domain to the frequency domain. The signal then goes to a total harmonic distortion analyzer. There, the harmonic frequencies and amplitudes are captured and sent into a formula node. The formula node contains equations for converting this data into a value for capacitance. This capacitance value is sent to the graph, and the C-V plot displayed.

## Slide 26: C-V Measurement Theory
To arrive at capacitance values:
  DC bias voltage + sin wave
  Differentiate this function
  Multiply by expansion of capacitance function
  This generates a function for the current I(t)
Comparing terms in the current function:
  The formulas for capacitance become evident
Notes: Input signal: V = Vo + Vac Cos wt
I(t)= -A1 sin wt - A2 sin wt - A3 Sin wt
Also, I(t) = C(V) dV/dt
C(V) = C0 + C1(V-Vo) + ½ C2 (V-Vo)^2
By comparing terms, we see that
  C0 = (A1/wVac) – 1/8 * Vac^2 * C2
  C1 = dC/dV = (2*A2)/w(Vac)^2
  C2 = d^2C/dV^2 = (8*A3)/w*(Vac)^3

## Slide 27: C-V Screenshot
Notes: This is a screenshot of the I-V/C-V Startup Kit. This is C-V measurement. The DUT is a MOSFET. The top graph is the Auto Power Spectrum of the measurement, which shows the harmonics of the signal while the test is running. The harmonic frequencies and amplitudes are shown in the upper right. The two lower graphs are the C-V curves. The vertical (y) axis is Capacitance, and the horizontal (x) axis is Voltage. The graph in the lower left is the actual data, while the graph in the lower right is the same data passed through a smoothing algorithm before it is sent to the display.

## Slide 28: Demonstration
I-V/C-V startup kit