## Slide 1: Structural Test and Measurement
Doug Shepard
Business Development Manager
Fri Aug 18
1:45-3:00 p.m.
San Jacinto (5B)

## Slide 2: Applications
Purpose – verify structural design
Static
  Load bearing test
Fatigue
  Long-term effects of slowly varying loads (<4 Hz)
Dynamic
  Vibration & modal test (~10KHz BW)

## Slide 3: Applications
Aircraft airframes
  Static and fatigue
Spacecraft & missiles
  Dynamic
Buildings (earthquake resistance studies)
  Low-frequency dynamic

## Slide 4: System Components
Mechanical stimulus
  Servo hydraulic jacks (low frequency)
  Shaker table (dynamic)
Sensor signal conditioning hardware
Data acquisition hardware
Data monitor software
Data recorder
Data analysis & presentation tools

## Slide 5: Example – Static and Fatigue Test
Predictive failure test and analysis
Hydraulics simulate in-flight mechanical loads
Monitor strain and mechanical changes
Periodic tear down and inspection
Hydraulics control and
SC/DAQ systems

## Slide 6: Traditional Measurement Technologies
UNIX
  Turnkey control, monitor, and analysis software
Monolithic signal conditioning
  Neff
VME or rack & stack data acquisition
Data logging
  Log to disk
  Analog or digital tape recorder

## Slide 7: PC-based Technologies
MIO E-Series 16-bit A/D
  PXI or PCI
SCXI signal conditioning
  SCXI-1530 accelerometer
  SCXI-1540 displacement/angle
  SCXI-1520 strain

## Slide 8: SCXI-1520
Eight-channel strain conditioning module
  Two to 10-wire interface
Fully software programmable per channel
  Excitation
  Bridge configuration
  Offset null
  Shunt calibration
  Filter
  Gain
  Simultaneous sample & hold

## Slide 9: 512 Channel Static/Fatigue Measurement System
  Integrator – Stress Engineering
  PXI/SCXI-based
  PXI-8170 700 MHz PC (NT)
  PXI-8210 100 baseT LAN
  PXI-6030E A/D
  SCXI-1520 strain conditioning (64)
  Dimensions
  24U (42” or 1.1m high)
  Standard 19” rack mount
SCXI 12-slot
Chassis (5)
PXI/SCXI Chassis

## Slide 10: Distributed Systems
Multiple stations
  Up to 512 channels each
  Controlled via network
  Synchronized by common clock signal
Station
Station
Station
Station

## Slide 11: Distributed System Advantages
Minimizes cabling requirements
  Short analog signal leads
Station
Station
Station
Station

## Slide 12: Distributed System Advantages
Independent system setup
  Each station can operate in local mode
Station
Station
Station
Station

## Slide 13: Distributed System Advantages
Reconfigurable
  Stations can be grouped into smaller subsystems
Station
Station
Station
Station

## Slide 14: PXI/SCXI System Architecture
Station
Excitation outputs signal inputs
Local
area
network
Station
Station
Station
Servo-
control
system
Acquisition control/monitor console
Timing
Notes: Mkt1– <number>
Expanding Your Business with Marketing
<header>
<date/time>

## Slide 15: SCXI Timing Distribution Module
One clock input, four clock outputs (coax)
Opto-isolated 50 Ohm line drivers
Any station can generate the master clock
Any station can be slaved to any other station
Master station
Slave station
Slave station
Slave station
Timing network

## Slide 16: Servo Control System Interface
Input signals from servo system
  Load cell output voltages
  Limit check event/alarm signals
    Data log command
Output signals to servo system
  Data log command acknowledge

## Slide 17: Software Structure
Stress Engineering StrainDAQ
  User interface
  System configuration and control
LabVIEW™
  Graphical programming
  Data acquisition
  Distributed system tools
NI-DAQ™
  SCXI and A/D driver
  Remote measurements

## Slide 18: Data Monitor and Log
Monitor
  Low-speed (~2-5 samples/sec) continuous
  Local and remote viewing
  Alarms
Event driven data log
  Signal threshold
  Periodic
  User command
Pre/post trigger data recording
  High-speed (~50 samples/sec)
  Deep history (~10 minutes)