## Slide 1: Understanding Timing, Triggering and Synchronization
Waqar Mahmood
Software Engineer
Wed Aug 16
10:15 a.m.-1:15 p.m., 1:45-4:45 p.m.
Exhibit (3G)

## Slide 2: Agenda
Timing
  Role of timing in measurements
  Timing error in measurements
    theory about sources of timing error
    examples
  How to minimize timing error
    measurement techniques
    phase-lock loop

## Slide 3: Agenda (cont.)
Triggering
  Role of triggering in measurements
  Triggering examples
Synchronizing
  Combine timing and triggering techniques
  Example
  Synchronize distributed measurements
Conclusion

## Slide 4: Role of Timing in Measurements
Timing is a fundamental component of measurements
Examples of timing in measurement
  Sampling clock for input
  Update clock for output
  Counter clock for frequency measurement
Timing error results in measurement error
Notes: You encounter different types of errors while making measurements: errors in sensors and errors at the measurement device. The errors in the measurement include voltage errors and timing errors. This discussion focuses on timing errors.
To acquire an analog waveform, the data acquisition device provides a scan clock to the analog to digital converter (ADC) to indicate when the conversion should occur. This acquisition results in a series of datapoints that represent the analog voltage sampled by the ADC. Adjacent samples are assumed to be separated in time as determined by the scan clock. The timing information implied by the scan clock depends on the quality of the scan clock.
For analog output, an update clock indicates to the digital to analog converter (DAC) when to change the voltage at its output. To generate a waveform, the output of the DAC is changed at regular intervals using the update clock. The frequency generated is related to the frequency of the scan clock.
Many applications require measuring a frequency. The reliability of your measurement is determined by the reliability of your timing. To measure distance, you must have a good yardstick. To measure time or frequency, you must have a good clock.

## Slide 5: Sources of Timing Error
Timing accuracy in measurements determined by
  Quantization error
    related to measurement duration and resolution for measuring time
  Clock error
    related to accuracy and stability of clock
Notes: Timing error describes how well you can measure or provide time. For example, if you measure a perfect 1 ms pulse and your measurement result is 1.003 ms, the 0.003 ms is your timing error in this measurement. When you provide timing, timing error describes how well you provide the desired frequency. For instance, if you want to sample an analog voltage, you will need a scan clock. Suppose you specify a 1 MHz scan clock while the actual scan clock frequency is 1.001 MHz. In this case, the 0.001 MHz frequency error represents the timing error.
Quantization error in timing is similar to quantization errors you encounter when converting an analog voltage to a digital voltage. For measuring analog voltage, you take a continuous parameter (voltage) and compare it against specified bins. The closer together these bins, the less the error. Similarly, for measuring time you take a continuous parameter (time) and sample it at specified instances. The closer together these instances, the less the error. The amount of time for quantization error is fixed (the separation between the instances at which you sample). Longer measurement times result in smaller quantization error because the error becomes a smaller fraction of the measurement time. Clock error represents the error in the frequency you use as a reference for measurements. When you generate a frequency, your error comes from the clock error. When you measure time, your error comes from the clock error as well as quantization error.

## Slide 6: Quantization Error Explained
GATE
SOURCE
1 2 3 4 5
GATE
SOURCE
1 2 3 4
Notes: The top half of the figure above shows that the falling edge of Gate is aligned with an edge of Source. You see that Gate duration is not equal to an integer number of cycles of Source. Gate goes high shortly before a rising edge on Source. Based on when Gate goes high and low with respect to Source as shown in the top half of the diagram, the counter returns a count of 5.
The bottom half shows Gate of the same width, and Source of the same frequency. The only difference in the bottom half is that Gate goes high shortly after a rising edge on Source. In this case the count returned is 4.
Gate occurs asynchronously to Source. That is, Gate can change independently of Source. How many counts are returned by the counter depends on when Gate changes with respect to Source. Because you have no control over when Gate changes with respect to Source, you must allow for the count to vary by one. This variation is referred to as quantization error. If the Gate duration lies between x and x+1 cycles of Source, the count returned will be x or x-1. In the example above, the Gate duration is between 5 and 6 cycles of Source. The count returned is 5 or 4.

## Slide 7: Quantization Error Explained (cont.)
GATE
SOURCE
Miss both edges
0 1 2 3 3
0 1 2 3 4
Miss one, catch one
1 2 3 4 5
Catch both edges
Notes: Suppose that Gate is equal to an integer number of cycles of Source. In the figure above, Gate is equal to 4 cycles of Source. Thus, our measurement should return a count of 4. When an edge on Gate occurs at the same time as an edge on Source, it is uncertain whether the counter will detect the Gate edge at that instance, or until the following edge on Source. Properties of the circuit and laws of probability determine on which edge the counter would detect the change in Gate. Because a Gate edge occurs the same time as the Source edge twice, let’s consider all possible cases. In the first case, the counter ignores both Source edges that occur at the same time as edge on Gate and returns a count of 3. In the second case, it ignores the Source edge at the beginning of the measurement and catches the one at the end of the measurement. The count returned is 4. The same count would be returned if it counts the edge at the beginning and ignores the edge at the end. In the last case, it sees each Source edge while the Gate is high and returns a count of 5.
Notice that you expect a count of 4 but may receive a count of 3 or 5. Thus, you have a quantization error of +/- 1 count.
More often, the quantization error will result because the Gate is NOT equal to an integer number of cycles of Source and occurs asynchronously to Source (as shown in previous slide). The situation described in the diagram above has low probability of occurrence.

## Slide 8: Clock Accuracy
Clock provided by oscillator
Accuracy describes how well actual frequency matches specified frequency
Factors that affect oscillator accuracy
  Quality of crystal
  Assembly of oscillator
  Calibration, if applicable
Described in parts per million (ppm)
Notes: Accuracy describes how well the actual frequency of the clock matches the specified frequency. Clock accuracy is described in parts per million (ppm). An accuracy of 5 ppm indicates that the mean frequency of the clock may be off by 5 Hz for every 1 MHz of its specified value. To find the error in Hz from a specification given as x ppm, multiply the frequency of the clock by x * 10-6. A 1 MHz clock with an accuracy of 5 ppm may have a frequency error of up to 5 Hz. A 2 MHz clock with the same accuracy may have a frequency error of up to 10 Hz.
A clock is provided by an oscillator. This clock may be used to provide secondary clocks. These secondary clocks still are related to the oscillator. The ppm error of the secondary clock remains the same as that of the primary clock (or oscillator). For example, if you take a 5 ppm 1 MHz clock provided by an oscillator and derive a 100 kHz clock from this 1 MHz clock, the 100 kHz clock will also have a 5 ppm error.

## Slide 9: Clock Stability
Stability describes how well the frequency resists fluctuations
Main factors that affect clock stability
  Temperature
  Aging
  Supply voltage
Notes: Stability describes how well the clock frequency resists fluctuations. Factors that can cause the frequency to fluctuate include variations in temperature, time (aging), supply voltage, shock, vibration, and capacitive load that the clock must drive. Temperature is often the dominant factor that affects crystal oscillator stability. Some oscillators are designed to be less susceptible to temperature – temperature compensated crystal oscillators (TCXOs) and oven controlled crystal oscillators (OCXOs). OCXOs offer the best temperature stability in crystal oscillators. Regular crystal oscillators give stabilities with temperature on the order of tens of ppm. TCXOs offer stability with temperature better than 10 ppm. OCXOs can offer stability with temperature better than 10 ppb! Aging is the dominant factor that affects crystal oscillator stability in OCXOs.
For crystal oscillators, the stability is usually provided as a ppm figure that includes all variables. For OCXOs, the stability can be broken down with respect to the different variables.
OCXOs are the most stable form of crystal oscillators. To achieve better stability, you must use atomic frequency standards such as a Rubidium frequency standard.

## Slide 10: Inaccurate, Unstable 10 MHz Clock
Frequency
Time
10
MHz
Notes: The curve represents the frequency of a 10 MHz clock over time. This clock is neither accurate nor stable. If you take the average value of the frequency, it will be greater than 10 MHz. Furthermore, the value of the frequency fluctuates about a mean value over time.

## Slide 11: Accurate , Unstable 10 MHz Clock
Frequency
Time
10
MHz
Notes: The curve represents the frequency of a 10 MHz clock over time. If you take the average value of the frequency it will be very close to 10 MHz. Thus, this clock is accurate. However, the clock is not stable since the frequency fluctuates about its mean at approximately 10 MHz.

## Slide 12: Stable, Inaccurate 10 MHz Clock
Frequency
Time
10
MHz
Notes: The curve represents the frequency of a 10 MHz clock over time. This clock is stable since its frequency does not fluctuate. However, because its frequency is greater than 10 MHz, it is not accurate.

## Slide 13: Accurate and Stable 10 MHz Clock
Frequency
Time
10
MHz
Notes: The curve represents the frequency of a 10 MHz clock over time. Its value is very close to 10 MHz and it shows very little or no fluctuations. Thus, this clock is accurate and stable.

## Slide 14: Timing Error Example
Measure frequency of 1 kHz signal by measuring its period using PXI-6602 counter
  How to measure frequency
  Quantization error example
  Clock error example
  Quantization and clock error example
Notes: Use a PXI-6602 to measure a 1 kHz frequency. A PXI-6602 has an 80 MHz clock with +/-50 ppm accuracy over a temperature range from 0 to 50 °C. This specification indicates that as the temperature varies from 0 to 50 °C, the frequency of the clock varies from 79,996,000 Hz (80 MHz minus 4000 Hz) to 80,004,000 Hz (80 MHz plus 4000 Hz).

## Slide 15: Counter Overview
32 bits
SOURCE
UPDOWN
GATE
A counter can be used to measure frequency
Notes: A counter has three basic inputs – Source, Gate, and UPDOWN. Pulses on Source cause the counter to increment or decrement. The counting direction (whether the counter increments or decrements) is determined by UPDOWN. Gate is a control signal that controls the counting process in different ways. A counter may be configured so that it starts counting when it sees a low to high transition on Gate and continues to count until it sees another low to high transition on Gate. Alternatively, the counter may be configured so it counts only when Gate is at a high level.

## Slide 16: Frequency Measurement via Measuring Period
GATE
SOURCE
0
1 2 3 4 5 5
5
Count
HW save register
Measure time between rising to rising or falling to falling edge to measure period
  Source is of known frequency
Notes: Frequency represents the number of cycles of a signal that occur in one second. Period represents the time between the start and finish of one cycle.
Frequency = 1/(period)
To measure frequency of the 1 kHz signal, we will measure its period and take the inverse of what we measure to get frequency. To measure period, we configure the counter so it begins to count when it sees a low to high transition on Gate and stops counting when it sees a second low to high transition on Gate. We use a clock of known frequency as Source. Thus, the counter returns to us how many cycles of the known frequency occur in one period of the signal we are measuring, which is connected to Gate. We multiply the count returned by the period of the known frequency on Source. The result is the duration of one period of the signal we are measuring. The inverse of the period gives the frequency.
For example, suppose the Source in the above diagram is a frequency of 1 kHz and, thus, has a period of 1 ms. We receive a count of 5. This count indicates that the period of the signal we are measuring is
5 * 1 ms = 5 ms
The frequency of the signal is
=1/ (5 ms)
=200 Hz

## Slide 17: Quantization Error Example
Measure 1 kHz frequency using perfect 80 MHz clock
Expected count is 80,000
Assume quantization error of -1, count returned is 79,999
Notes: Suppose you measure the frequency of a 1 kHz signal. You use a counter on a PXI-6602 to measure the period of this signal and take the inverse of period to calculate frequency. The counter uses an 80 MHz clock to measure the period. It does so by counting how many cycles of the 80 MHz clock occur in one period of the signal being measured. Assume that the frequency of external signal is a perfect 1 kHz and that 80 MHz clock has no clock error. The expected count of number of cycles is
= period of 1 kHz signal/period of 80 MHz signal
= frequency of 80 MHz signal/frequency of 1 kHz signal
= 80 MHz/1 kHz
= 80,000
Assume that quantization error is -1 count. The count returned is
= 79,999
We assumed that there is no clock error in order to focus on quantization error.

## Slide 18: Quantization Error Example (cont.)
Measured frequency is
= 1/period of signal measured
= 1/(measured count * period of perfect 80 MHz)
= 1/(79,999 * 12.5 ns)
= 1000.0125 Hz (actual frequency is 1000 Hz)
= 12.5 ppm error
Notes: We are assuming that there is no clock error and the 80 MHz clock has a perfect 80 MHz frequency. The frequency measured is:
= 1/period of signal measured
= 1/(measured count * period of 80 MHz)
= 1/(79,999 * 12.5 ns)
= 1000.0125 Hz
= 12.5 ppm error
We measured a perfect 1,000 Hz frequency and our measurement returned 1,000.0125 Hz, which represent a 12.5 ppm error. A quantization error of 1 results in an error of 12.5 ppm in this example.

## Slide 19: Clock Error Example
Measure 1 kHz frequency using 80 MHz clock that has -25 ppm error (2,000 Hz)
Expected count is 79,998
Assume no quantization error
Notes: Suppose you measure frequency of 1 kHz signal. You use a counter on a PXI-6602 to measure the period of this signal and take the inverse of period to calculate frequency. The counter uses an 80 MHz clock to measure the period. It does so by counting how many cycles of the 80 MHz clock occur in one period of the signal being measured. Assume that frequency of external signal is a perfect 1 kHz and that under ambient conditions the actual frequency of 80 MHz clock is 80 MHz - 25 ppm, which is 80 MHz - 2000 Hz . Expected count of number of cycles is
= period of 1 kHz signal/period of 80 MHz signal
= frequency of 80 MHz signal/frequency of 1 kHz signal
= (80 MHz - 2000 Hz)/1 kHz
= 79,998
In this measurement we assume that we have no quantization error so we can focus on the effect of the clock error.

## Slide 20: Clock Error Example (cont.)
Measured frequency is
= 1/period of signal measured
= 1/(measured count * period of perfect 80 MHz)
= 1/(79,998 * 12.5 ns)
= 1000.025 Hz (actual frequency is 1000 Hz)
= 25 ppm error
Notes: You do not know your exact frequency for the 80 MHz clock at the time of measurement. What you do know is that it will be 80 MHz +/- 50 ppm over the operating temperature range. Since you do not know what the actual frequency is at any given time, for calculations you treat the 80 MHz clock as a perfect clock. Thus, the frequency calculated by the above measurement is:
= 1/period of signal measured
= 1/(measured count * period of 80 MHz)
= 1/(79,998 * 12.5 ns)
= 1000.025 Hz
= 25 ppm error
We measured a perfect 1,000 Hz frequency and our measurement returned 1,000.025 Hz, which represent a 25 ppm error. A clock error of 25 ppm gives a measurement error of 25 ppm.

## Slide 21: Quantization and Clock Error Example
Measure 1 kHz frequency using 80 MHz clock that has -25 ppm error
Expected count is 79,998
Assume quantization error of -1, count returned is 79,997
Notes: Suppose you measure frequency of 1 kHz signal. You use a counter on a PXI-6602 to measure the period of this signal and take the inverse of period to calculate frequency. The counter uses an 80 MHz clock to measure the period. It counts how many cycles of the 80 MHz clock occur in one period of the signal being measured. Assume that the frequency of external signal is a perfect 1 kHz, and under ambient conditions the actual frequency of 80 MHz clock is 80 MHz - 25 ppm, which is 80 MHz - 2000 Hz . Expected count of number of cycles is
= period of 1 kHz signal/period of 80 MHz signal
= frequency of 80 MHz signal/frequency of 1 kHz signal
= (80 MHz - 2000 Hz)/1 kHz
= 79,998
Assume that quantization error is -1 count. The count returned is
= 79,997
.

## Slide 22: Quantization and Clock Error Example (cont.)
Measured frequency is
= 1/period of signal measured
= 1/(measured count * period of perfect 80 MHz)
= 1/(79,997 * 12.5 ns)
= 1,000.0375 Hz (actual frequency is 1,000 Hz)
= 37.5 ppm error
= 25 ppm (clock error) + 12.5 ppm (quantization error)
In this example, neither error dominates
Notes: You do not know your exact frequency for the 80 MHz clock at the time of measurement. What you do know is that it will be 80 MHz +/- 50 ppm over the operating temperature range. Since you do not know what the actual frequency is at any given time, for calculations you treat the 80 MHz clock as a perfect clock. Thus, the frequency calculated by the above measurement is:
= 1/period of signal measured
= 1/(measured count * period of 80 MHz)
= 1/(79,997 * 12.5 ns)
= 1000.0375 Hz
= 37.5 ppm error
We measured a perfect 1,000 Hz frequency and our measurement returned 1,000.0375 Hz, representing a 37.5 ppm error. Notice that 37.5 ppm total error is the sum of 25 ppm clock error and 12.5 ppm quantization error. Also note that in this example, the errors were cumulative. If the clock error had been +25 ppm (instead of -25 ppm), the two would have worked to balance each other out and resulted in a total error of
= (25 ppm - 12.5 ppm)
= 12.5 ppm
Which error dominates depends on how much clock error you have, and your parameters that determine the quantization error.

## Slide 23: Summary of Quantization and Clock Error Examples
Notes: For the case in which we had quantization and clock error, we measured a perfect 1,000 Hz frequency and our measurement returned 1,000.0375 Hz, which represents a 37.5 ppm error. Notice that 37.5 ppm total error is sum of 25 ppm clock error and 12.5 ppm quantization error. Also note that in this example, the errors were cumulative. If the clock error had been +25 ppm (instead of -25 ppm), the two would have worked to balance each other out and resulted in a total error of
= (25 ppm - 12.5 ppm)
= 12.5 ppm
Which error dominates depends on how much clock error you have, and your parameters that determine the quantization error.

## Slide 24: How to Reduce Quantization Error
Quantization Error is
  (resolution for time measurement)/(measurement duration)
In our example,
  resolution for measuring time = 12.5 ns
  measurement duration = 1 ms
  12.5 ns/ 1 ms = 12.5 ppm
To reduce quantization error increase
  measurement duration
  resolution for measuring time
Notes: In our example, a sampling error of 1 count gave a 12.5 ppm error. Where does this number come from?
The quantization error results in an error of one count in the value returned by the counter. Since the count represents cycles of the 80 MHz clock in our example, the corresponding time error is 12.5 ns. Our measurement duration is one period of the signal being measured, which is 1 ms. Thus, the error is 12.5 ns out of 1 ms, which is 12.5 ppm.
To reduce quantization error, increase the measurement duration (measure over two periods), or increase the frequency of Source.

## Slide 25: Reducing Quantization Error for Frequency Measurement
Low frequencies
  Measure period, take inverse
High frequencies
  Gate counter for known time, count edges
Out
Gate
Source
Signal being measured
Ctr 0:
Pulse train
generation
Ctr 1:
Period
measurement
Ctr(0)
Ctr(1)
Source
Internal timebase
Notes: Assume clock error is zero so we can focus on quantization error. Period measurement is the best way to measure for low frequencies. Only one counter is required.
Frequency = 1/Period
Quantization error can cause period measurement to be off by +/- 1 cycle of Source. You can measure maximum frequency (F) through period measurement and stay within an error of 0.x
F = FSOURCE-(FSOURCE/(1+0.x))
FSOURCE is the frequency of Source.
As frequencies become higher, possible quantization error in single period measurement increases, so methods using two counters are recommended. For the two-counter method shown on this slide, gate one counter for known amount of time and count how many edges of the signal being measured occur in this time.
F = Count returned/Time in seconds
Quantization error can cause count to be off by +/- 1 cycle of Source. Measurement error is (+/- 1 Hz)/(Time in seconds). For example, suppose you are measuring the frequency of a 10 MHz signal and gating it for 1 ms. Expected count is 10,000.
F = 10000/1 ms = 10 MHz.
A quantization error of -1 would make expected count 999.
F = 9999/ 1 ms = 9,999,000 Hz (error = 100 ppm)
So a quantization error of 1 in 1 ms gating period results in an error of
1/( 1 ms) Hz, or 1000 Hz.
Increase gating period to 1 second and quantization error will reduce to 1 Hz.

## Slide 26: Reducing Quantization Error for Frequency Measurement
Higher frequencies
  Measure time over many cycles, derive single period, take inverse
–
Out
Gate
Source
Signal being measured
Ctr(0)
Ctr(1)
Internal timebase
Ctr 0:
Pulse train
generation
Ctr 1:
Period
measurement
Source
Notes: To measure frequency using the method shown above, measure time over many integer cycles of the external signal. Divide the time measured by number of cycles used and you obtain the period of one cycle. The inverse of the period gives frequency.
F=N/Time measured over N cycles of external signal
N is the number of cycles over which you make your measurement.
This method can be more accurate than the method of gating a counter for same measurement duration. Quantization error can be +/- 1 cycle of Source, which can be of higher frequency than the signal being measured, resulting in smaller error. As total time increases, the error of +/-1 clock cycle of Source becomes less significant. For example, if you measure a signal with a 10 MHz frequency over 10,000 cycles using an 80 MHz timebase, the error of +/- 1 clock cycle will amount to an error of 12.5 ns out of 1 ms (12.5 ppm). If you use 10 million cycles, the error will be 12.5 ns out of 1 s (12.5 ppb).
The advantage this method offers over gating a counter for known period and counting edges is that this is less sensitive to quantization error and makes the measurement over an integer number of cycles of the signal being measured. The disadvantage is that it is more likely to be effective if you know what the frequency should be of the signal you are measuring. Knowing the expected frequency will help determine how many cycles over which to make the measurement. For example, measuring time over 10 million cycles works out well for a 10 MHz signal, but would make the measurement too long for a 1 kHz signal.

## Slide 27: Reducing Clock Error
Use device with high stability clock
Use device with a phase-lock loop that uses a high stability clock as a reference clock

## Slide 28: Phase Lock Loop
Circuit that adjusts an oscillator so its frequency is locked to a reference clock
VCXO is voltage controlled crystal oscillator
  Frequency varies with input DC voltage
Compare
Phase
Filter
VCXO
Reference clock
Locked frequency
Notes: A phase lock loop is a circuit that compares two clocks: a reference clock and a voltage controlled crystal oscillator clock (VCXO). A VCXO is an oscillator that allows its frequency to be tuned by applying a DC voltage to it. The larger the voltage, the higher the frequency.
A phase lock loop has three main components: phase comparator, loop filter, and VCXO. The phase comparator accepts the reference frequency and the VCXO frequency and divides the two frequencies to attain a common frequency. For instance, if a phase lock loop is locking a 15 MHz VCXO clock to a 10 MHz reference clock, it may divide the frequency of the 15 MHz clock by 15, and that of the 10 MHz clock by 10. It then compares the phase and frequency of the two 1 MHz clocks thus derived and generates a 1 MHz pulse train. The better the two clocks have the same frequency and phase the narrower the pulsewidths. The pulses are fed to a loop filter that delivers a DC voltage at its output based on the width of the pulses at its input. This DC voltage is applied to the VCXO to control its frequency. The circuit works to adjust the DC voltage to minimize the width of the pulses, and to make the frequency and phase of 1 MHz clock derived from the VCXO (in this example) be the same as the 1 MHz clock derived from the reference clock. Thus, the frequency of the 15 MHz VCXO becomes exactly 1.5 times that of the 10 MHz clock.

## Slide 29: Phase Lock Loop Continued
80 MHz Locked
Frequency
10 MHz Reference Frequency
Using high-stability clock as reference reduces clock error of VCXO clock
Notes: The PXI-6602 counter has an 80 MHz VCXO that is phase-locked to a reference 10 MHz clock. Due to phase locking, the 80 MHz VCXO attains a frequency exactly 8 times that of the 10 MHz reference clock. In addition, the phase is adjusted so that every eighth edge of the 80 MHz VCXO clock aligns with an edge of the 10 MHz reference clock.
The stability of a VCXO driven by a phase-lock loop becomes the same as that of the reference clock. So if a VCXO is rated for 50 ppm stability but it is locked to a 10 MHz clock that has 25 ppm stability, the VCXO will now inherit the 25 ppm stability. It will follow fluctuations in the reference clock. Suppose that the 10 MHz reference clock has an error of -25 ppm. That amounts to -250 Hz error. An 80 MHz VCXO phase locked to this reference clock will have a frequency equal to exactly 8 times the reference clock. Its frequency will be
= (10 MHz - 250) * 8
= (10 MHz *8 - 250 *8)
= (80 MHz - 2000)
= 79,998,000 Hz
= 25 ppm error
So while the error in the 80 MHz clock is 2000 Hz compared to 250 Hz in the 10 MHz clock, the ppm value remains the same.

## Slide 30: Reduce Clock Error in PXI
PXI chassis has 10 MHz backplane clock
  Common reference clock provided to each slot
  Trace lengths matched to minimize skew
  May be replaced by another 10 MHz clock
Devices phase-locked to the backplane clock
  Clock edges occur simultaneously
  Frequencies are derived from common 10 MHz backplane clock
Replace backplane clock with high-stability 10 MHz clock to reduce clock error in PXI
Notes: We discussed that the PXI-6602 counter phase-locks its 80 MHz VCXO to a reference 10 MHz clock. This reference clock is the PXI backplane clock. This clock is provided to each module in the chassis. The clock distribution is designed to minimize clock skew across the modules. Modules that phase lock to the backplane clock include PXI-6608 counter, PXI-6602 counter, PXI-5411 Arbitrary Waveform Generator, and PXI-5911 100 MHz Scope.
Modules that are phase-locked to the backplane clock have their edges synchronized (more so if they have the same frequency VCXOs) and have their frequencies derived from the common 10 MHz clock. A 20 MHz frequency will be exactly 2 times the 10 MHz frequency. An 80 MHz frequency will be exactly 8 times the 10 MHz frequency. Because the frequencies are derived from the 10 MHz frequency, the 80 MHz frequency is now exactly 4 times the 20 MHz frequency. This type of synchronization of clocks across modules is particularly useful for a multi-device system that performs multiple measurements which must be compared to each other.

## Slide 31: Benefit of PXI-6608 Counter with High Accuracy Clock
Use 10 MHz OCXO on PXI-6608 to drive the PXI 10 MHz backplane clock
  75 ppb (10 MHz +/- 0.75 Hz) stability over 1 year
Improve timing of overall system
  devices phase-locked to backplane clock inherit stability of OCXO clock
Pass over PXI triggers (RTSI) for direct use
Notes: We discussed that a VCXO in a phase-lock loop inherits the stability of the reference clock to which it is phase-locked. And we know that a PXI chassis distributes a 10 MHz clock to all its modules. If this 10 MHz clock is highly stable, all modules phase-locked to this clock inherit the same stability. PXI chassis design allows this benefit. A module installed in start trigger controller slot (slot 2) of a PXI chassis can replace the chassis’ 10 MHz clock with its own 10 MHz clock.
The PXI-6608 counter has an oven controlled crystal oscillator (OCXO) on it that provides a 10 MHz clock with a stability of 75 ppb over one year--that’s 75 parts per billion! When the PXI-6608 is installed in the start trigger controller slot, NIDAQ software automatically uses the OCXO clock to drive the PXI backplane clock. Thus, all devices that have their clocks phase-locked to the PXI backplane clock now inherit the 75 ppb stability.
The OCXO clock can also be passed over PXI triggers (the equivalent of RTSI on PXI) to be used as scan clock, update clock, or general purpose counter clock for devices. RTSI is the real time system integration bus that allows boards within a system to share signals and timing information.

## Slide 32: PXI-6608 Improves Measurement by Reducing Clock Error
Frequency measurement with PXI-6608 counter and PXI-6602 counter
Notes: The benefit of the OCXO on the PXI-6608 counter is shown by the table above. We see that the error in frequency is about 0.1 Hz when measured by the PXI-6608 counter, compared to over 17 Hz when measured by the PXI-6602 counter. For the measured frequency, these errors amount to about 10 ppb and 1.7 ppm under the ambient conditions, respectively.
Note that if the PXI-6602 counter was used to make a measurement in a chassis that also had the PXI-6608 counter in the start trigger controller slot, the PXI-6602 measurement would have been as accurate as that of the PXI-6608 because of phase of phase locking.

## Slide 33: PXI-6608 Improves Generation by Reducing Clock Error
Sine wave generation with PXI-5411 arbitrary waveform generator and PXI-6608 counter
Notes: This slide shows how the OCXO of the PXI-6608 drives the backplane and provides benefit to the PXI-5411 Arbitrary Waveform Generator, which uses phase locking. Under the ambient conditions, the frequency error in generation when used with the PXI-6608 is about 6 ppb. With no PXI-6608 in the system, the error in frequency generation is 5.6 ppm.

## Slide 34: What is a Trigger?
A hardware or software event, prompts action
  Start acquisition
  Stop acquisition
  Sample/latch
Benefits of triggering
  Begin operation instantaneously when an event occurs
  Simultaneously start multiple operations
Notes: A trigger is a hardware or software event that usually prompts an action. For instance, a start trigger begins an acquisition. A device waiting for a hardware start trigger remains in the armed state but does not begin the acquisition or measurement until the start trigger occurs. When the trigger occurs, it begins the operation with deterministic timing. The trigger could also be provided by software. Software trigger involves software latency and some uncertainty about time between when the command is given and when the operation begins. A hardware trigger has deterministic and faster response. For example, a start trigger for a counter on a PXI-6602 counter arms the counter within 12.5 ns of when the trigger occurs. A trigger can also be used to stop an acquisition, or to latch a measurement value.
The deterministic and, often, instantaneous response to a trigger is one benefit of triggering. If you must begin an operation – measurement or generation – in response to an external event, a trigger allows you an effective way to do so. Also, a trigger is useful for starting multiple operations simultaneously. Each operation can be configured to use a common start trigger.

## Slide 35: Triggering Example
Simultaneous arm two counters in a PXI-6602 8-channel counter
  PXI-6602 has two NI-TIO ASICs
  Uses 80 MHz clock
Concepts of example using counters apply to other devices
Notes: To discuss timing concepts involved in triggering, we will consider an example application with the PXI-6602 counter. In this example, we will arm two counters simultaneously. The counters could be performing any function such as pulse generation or a time measurement.
The PXI-6602 counter has two NI-TIO ASICs. The NI-TIO ASIC is a sophisticated, state of the art ASIC developed by National Instruments for timing and digital I/O applications. Each ASIC has 4 counters in it. With two of these ASICs, a PXI-6602 counter offers 8 counters. NI-TIO ASIC 0 has counters 0 through 3. NI-TIO ASIC 1 has counters 4 through 7.

## Slide 36: Trigger Synchronized to Source
External start trigger
(asynchronous)
Ctr 0
Ctr 0 Source
PXI-6602 (8 channel counter device)
External start trigger synchronized to Source
ASIC 0-
Ctrs (0:3)
Ctr 4
ASIC 1-
Ctrs (4:7)
Notes: A digital circuit is more reliable when it is synchronous. In a synchronous circuit, signals change state on clock edges. The behavior of the circuit is more predictable and is more robust.
When a counter in a PXI-6602 counter is configured to use a start trigger, the start trigger is synchronized to Source of counter.

## Slide 37: Trigger Synchronized to Source
Source
External start trigger
External start trigger
synchronized to
Source
Armed counter awaiting trigger
Counter begins operation
Trigger seen by counter
1
2
3
4
External start trigger
External start trigger synchronized to Source
Counter
Notes: The figure above demonstrates how a start trigger is synchronized to Source of a counter. The external start trigger is sampled on each rising edge of Source and the value sampled is passed through. Because the value is sampled and passed through on rising edges of Source, transitions on start trigger at the input of the counter (at the output of the synchronizing circuit) occur only on rising edges of Source. Thus, the start trigger is said to be synchronized to Source of the counter.
In the diagram above, External Start Trigger is the input to the synchronizing circuit and goes high between the first and second rising edge of Source. The output of the synchronizing circuit is the External Start Trigger Synchronized to Source. The counter samples the External Start Trigger Synchronized to Source on each rising edge of Source. When it samples the start trigger at the output of the synchronizing circuit on the second clock edge, it sees a low voltage and does nothing. On the same clock edge, the synchronizing circuit samples its input (External Start Trigger) to be high and makes its output (External Start Trigger Synchronized to Source) high. Because the sampling on the second clock edge causes this output to change, the output becomes high a short delay after the second edge. Thus, the counter still sees the output to be low at the second clock edge. On the third edge the counter samples the start trigger at the output of the synchronizing circuit (External Start Trigger Synchronized to Source) to be high and, thus, begins its operation.

## Slide 38: Trigger Synchronized to Source of Two Counters in One ASIC
Ctr 0
Ctr 0 source
PXI-6602 (8 channel counter device)
Ctr 1
Ctr 1 source
ASIC 0
External start trigger
(asynchronous)
External start trigger synchronized to Source
I/O pad
3 ns
7 ns
Notes: Consider the case where you arm two counters simultaneously. As shown the diagram above, the start trigger is synchronized to Source at each counter. In order to start the two counters simultaneously, they must use the same Source. So Source of counters 0 and 1 must be the same clock, such as an internal clock.
The start trigger enters the ASIC through a pin. The circuit immediately inside the pin is called the I/O pad. The delay from the I/O pad to each counter is different. Assume a 3 ns delay from the I/O pad to counter 0, and a 7 ns delay from the I/O pad to counter 1.

## Slide 39: Trigger Synchronized to Source of Two Counters in One ASIC
Source
External start trigger
Start trigger synchronized at Ctr 0
Trigger seen by Ctrs 0 and 1
Start trigger at Ctr 0 after 3 ns delay
Start trigger at Ctr 1 after 7 ns delay
Start trigger synchronized at Ctr 1
1
2
3
4
Notes: The External Start Trigger goes high between clock edges 1 and 2. It reaches each counter before clock edge 3. Therefore, the synchronizing circuit at each counter see the External Start Trigger as high on clock edge 3 and passes it through. Each counter senses the Start Trigger Synchronized as high on clock edge 4 and begins its operation.

## Slide 40: Trigger Synchronized to Source of Two Counters in One ASIC
Source
External start trigger
Start trigger synchronized at Ctr 0
Trigger seen by Ctr 0
Start trigger at Ctr 0 after 3 ns delay
Start trigger at Ctr 1 after 7 ns delay
Start trigger synchronized at Ctr 1
Trigger seen by Ctr 1
1
2
3
4
Notes: The diagram above is similar to that on the previous slide. The External Start Trigger still goes high between clock edges 1 and 2. However, it occurs earlier than it did in the previous slide. Because it occurs earlier, External Start Trigger reaches counter 0 after a 3 ns delay just before clock edge 2. Thus, the synchronizing circuit at counter 0 samples External Start Trigger to be high at clock edge 2. Since the delay to counter 1 is longer 7 ns, the External Start Trigger is sampled low by the synchronizing circuit at counter 1 on clock edge 2. The synchronizing circuit at counter 1 does not see External Start Trigger as high until clock edge 3. The result is that counter 0 begins its operation on clock edge 3, while counter 1 begins its operation on clock edge 4.
We see this skew because the asynchronous External Start Trigger is synchronized to Source in two separate locations. Therefore, it is possible for it to be synchronized on two separate clock edges. A PXI-6602 counter has an 80 MHz internal clock. A skew of one cycle of this clock is 12.5 ns. For some applications, this skew is acceptable. If not, there is a way around this issue.

## Slide 41: Common Synchronized Trigger Passed to Each Counter
asynchronous trigger
Ctr 0
Ctr 0 Source
PXI-6602 (8 channel counter device)
Ctr 1
Ctr 1 Source
ASIC 0
Max
timebase
start trigger synchronized at pad
I/O pad
Notes: In the previous slide, we saw a skew between when counters 0 and 1 began their operations because the asynchronous External Start Trigger is synchronized to Source in two separate locations. To avoid this problem, the asynchronous External Start Trigger can be synchronized in one common location. This common location is the I/O pad. The Start Trigger Synchronized at Pad can then be passed to each counter. The Start Trigger Synchronized at Pad will be sampled by the synchronizing circuit at each counter and resynchronized to Source. Because it has already been made synchronous to Source at the pad, it now behaves more predictably and is more robust. By synchronizing at the pad, we ensure that both counters see the start trigger on the same clock edge.

## Slide 42: Common Synchronized Trigger Passed to Each Counter
Source
External start trigger
Start trigger at Ctr 0 after 3 ns delay
Start trigger synchronized at pad
Start trigger synchronized at Ctr 0
Start trigger at Ctr 1 after 7 ns delay
Start trigger synchronized at Ctr 1
Trigger seen by Ctrs 0 and 1
1
2
3
4
Notes: The External Start Trigger goes high at the same time it did for the case when the skew occurred. This time, we synchronize the external start trigger at the pad. The Start Trigger Synchronized at Pad goes high at clock edge 2. With the 3 ns delay to counter 0, and 7 ns delay to counter 1, Start Trigger Synchronized at Pad reaches counters 0 and 1 before clock edge 3. At clock edge 3, the synchronizing circuit of each counter samples the start trigger to be high at clock edge 3. Each counter sees the start trigger on clock edge 4.
At the pad of PXI-6602 counter, the External Start Trigger can be synchronized to the 80 MHz clock only. At the synchronizing circuit at each counter, the synchronization is done with respect to Source of the counter. Thus, for this technique to be effective, Source of counter must be the 80 MHz clock, or a timebase derived from this 80 MHz clock (internal 20 MHz or 100 kHz clock).

## Slide 43: Use Single Triggered Pulse to Synchronized to Source
Pulse is generated out output of counter upon receiving a pulse or trigger at Gate
Gate
Source
Out
Notes: To generate a pulse, the counter loads the value for inactive phase and counts down to zero. On reaching zero, its OUT toggles state and loads value for active phase. It then counts to zero and toggles state again. The duration of each phase is programmable.
Triggered pulse generation produces a single pulse with a programmable delay and pulse width on receiving a trigger on its Gate after the counter has been armed. The counter then disarms.

## Slide 44: Common Synchronized Trigger Passed to Counters in Two ASICs
asynchronous trigger
ASIC 0
G
O
Ctr 0
Ctr 1
Ctr 5
ASIC 1
Notes: Consider a case when you must arm counters in different ASICs. In this example we must arm counters 1 and 5. Pad synchronization is done at the I/O pad in each ASIC. If you use pad synchronization, again you are synchronizing in two separate locations and you may see a skew of one clock cycle. If you use opposite polarity for your start trigger in each ASIC (use positive polarity in one ASIC and negative polarity in the other ASIC), you reduce the skew to one-half clock cycle. In this case, you will always see a skew of one-half clock cycle. For a PXI-6602 counter, a one-half clock cycle delay of the 80 MHz clock amounts to 6.25 ns.
To avoid this skew, synchronize the external start trigger in one common location and pass that onto counters 1 and 5. You can use a triggered pulse generation with a counter to achieve this synchronization. For example, connect your external start trigger to Gate of counter 0, PFI_38. Configure counter 0 for triggered pulse generation and use the 80 MHz clock as its Source. When a pulse occurs on Gate of counter 0, it causes a pulse to be generated at the output of counter 0. The pulse at the output is synchronous to the 80 MHz clock. This output is driven to the I/O pin for the output of counter 0, PFI_36. PFI_36 is a bi-directional line. What is being driven as output can also be used as input. Now use this input as the start trigger and use pad synchronization. In software, select PFI_36 as the start trigger. Use the 80 MHz clock or a derived clock (internal 20 MHz or 100 kHZ) as Source for counters 1 and 5.

## Slide 45: Common Synchronized Trigger Passed to Each Counter
Source
External start trigger
Pulse generated at output of Ctr 0
Start trigger synchronized at Gate(0)
Pulse synchronized at pad for input
Pulse synchronized at Ctrs 1 and 5
Trigger seen by Ctrs 1 and 5
1
2
3
4
5
6
7
Notes: External start trigger is connected to PFI_38, Gate of counter 0. Counter 0 is configured for triggered pulse generation. The Gate of counter 0 is synchronized to Source of counter 0. Source of counter 0 is 80 MHz clock. The External Start Trigger goes high just before clock edge 1. It is synchronized to Source at Gate input of counter 0 on clock edge 1. We call this signal Start Trigger Synchronized at Gate(0) in the figure. The counter sees the transition on Start Trigger Synchronized at Gate(0) on clock edge 2 and begins its operation. A triggered pulse generation has an initial delay. The minimum value for this delay is two cycles. Thus, the Pulse Generated at Output of Ctr(0) does not go high until clock edge 4. This pulse is generated on PFI_36. PFI_36 is a bi-directional line on both ASICs. Each ASIC resynchronizes the pulse on PFI_36 to the 80 MHz clock at its I/O pad for PFI_36 on clock edge 5. Pulse Synchronized at Pad for Input in each ASIC is resynchronized to Source of counters 1 and 5 on clock edge 6. Source for each counter is the 80 MHz clock. The counters see the start trigger on clock edge 7 and begin their operation.
Note that the skew has now been eliminated. Although the trigger occurs before clock edge 1, the counters do not see it until clock edge 7. The process of synchronization adds a delay equal to an integer number of cycles of Source. In this example, a delay of 5 cycles is added and the trigger is shifted in time.

## Slide 46: Synchronizing Measurements
Make multiple measurements at the same time
To synchronize measurements
  Start all measurements simultaneously
  Use a common clock among all measurements
Synchronization required for correlating measurements
Notes: Synchronizing measurements means making multiple measurements at the same time. Most applications require making more than one type of measurement. Sometimes these measurements must be taken at the same time or sometimes they must be taken within strict timing constraints.
When synchronizing measurements (like making 100 temperature and speed measurements), you need to start all the measurements at the same time as well as share a common clock for latching data among all measurements.
For example, in control-loop applications, you need to make multiple measurements at the beginning of the loop, perform a calculation based on the new measurements, and finally do outputs based on this calculation. Control- loop applications require that you start all the measurements at the same time and synchronize them via a common clock signal. Similarly, if you want to correlate measurements, for example if you want to make a plot of speed and brake-pad temperature versus time, you will need to first synchronize both the speed and temperature measurements.

## Slide 47: Examples of Synchronizing Measurements
Anti-lock braking system (ABS) testing
  Synchronize analog input and counter measurements
Car engine ignition testing
  Synchronize digital input and counter measurements
Notes: In order to further understand the requirements for synchronization of measurements, let’s look at two real-life measurement systems.
The first example is an ABS testing system. This example demonstrates how to synchronize analog input and counter measurements using National Instruments products.
The second example is a car engine ignition testing system. This example demonstrates how to synchronize digital input and counter measurements using National Instruments products.

## Slide 48: Position Sensor Background
Magnetic pickup
sensor
Quadrature encoder outputs
Notes: Position sensors are used to track both rotational and linear movement. The following discussion applies to position sensors that track rotational movement.
The picture on the left shows a magnetic encoder. It consists of a rotating gear made of ferrous metal and a magnetic pick-up that contains a permanent magnet and the sensing element. The gear, which is mounted on the rotating shaft, has precisely machined teeth that provide the code pattern. As the disk rotates, these teeth disturb the magnetic flux emitted by the permanent magnet, causing the flux field to expand and collapse. These changes in the field are sensed by the sensing element, which generates a corresponding digital or pulse signal output.
The picture on the right shows the signals generated by a quadrature encoder. The most common type of quadrature encoder has two output channels (A and B). Using two code tracks with sectors positioned 90 degrees out of phase, the two output channels of the quadrature encoder indicate both position and direction of rotation. If A leads B, for example, the disk is rotating in a clockwise direction. If B leads A, then the disk is rotating in a counter-clockwise direction. By monitoring both the number of pulses and the relative phase of signals A and B, you can track both the position and direction of rotation.
Some quadrature encoders also include a third output channel, called a zero or index or reference signal, which supplies a single pulse per revolution. This single pulse is used for precise determination of a reference position.

## Slide 49: Test ABS
ABS controller
Pressure
Temperature
Speed
Brake
pad
Brake disc or rotor
Quadrature encoder
(speed sensor)
Apply brakes
Brake
pad
Notes: The above slide shows a simple disc brake ABS system. The ABS controller monitors the pressure applied to the rotating disc, the brake pad temperature and the speed of the rotor as it provides the brake pulses to the brake pad actuators. The pressure is monitored using a strain gauge, the temperature using a thermocouple and the speed using a quadrature encoder. A magenetic pickup sensor can also be used to track rotor speed.
The ABS testing system needs to plot a chart of the temperature and pressure of the brake pads and the speed of the rotor in revolutions per minute (rpm) versus time. To get this, first we need to synchronize the pressure, temperature and speed measurements. The pressure and temperature measurements are analog measurements while the speed measurement can be done by using a counter (counter measurement).

## Slide 50: Solution – Synchronize Analog Input and Counter Measurements
Analog measurements
Counter measurement
Both measurements
share a common clock
Pressure
Temperature
Speed
Common start trigger input
Notes: The brake pressure and the pad temperature signals are first conditioned via a signal conditioning accessory and then separately connected to two analog input channels. The channels are configured to be scanned at, say, 1 KHz. Also, the scan clock (1KHz) is shared via the real time system integration (RTSI) bus so that the counter can also receive it. To measure speed, the signals from the quadrature encoder are connected to the inputs of a counter. The counter is configured to count the number of input pulses between each edge of the shared scan clock (1KHz). The quadrature encoder produces 25 pulses per revolution. The speed of rotation can then be derived via
instantaneous speed in Hz = (number of quad. encoder pulses between the last
scan clock period / quad. encoder pulses per
revolution ) * freq. of the scan clock
= (n / 25) * 1000 Hz
The pressure and temperature measurements are read directly off the device and scaled based on the type of signal accessory used.
To synchronize the measurements, both the analog and the counter measurement are configured to start simultaneously via a hardware start trigger signal. This ensures a common start time for all measurements. Also, by sharing the scan clock, all the readings are synchronized to each other with respect to time.

## Slide 51: Plot of Pressure, Temperature, Speed Versus Time
Pressure
Temperature
Speed
Time
t0
t1
Measured value
Notes: The above chart shows one possible data set coming out of the ABS testing system. All three measurement are started at t0. Brakes are first applied at t1, which causes the the brake pad pressure to pulsate, speed of the rotor to fall and the temperature of the disc to rise. The dotted vertical line represents one instance in time(t1) when all the three measurements are synchronized to each other.

## Slide 52: Test the Timing of a Car Engine
1
3
4
2
1
2
3
4
Quadrature encoder
Distributer
Spark plugs
Fires spark plug 3
Shaft
Notes: The above slide shows a typical timing system of a car engine. The goal is to test the various parts of this timing system. It consists of a distributer, a shaft that rotates the distributer and four individual spark plugs (as shown above). As the shaft rotates, the distributer fires each of the spark plug at predefined angles. In the above picture, the distributer fires spark plugs 2, 4, 3, 1 in that order, each at 90 degrees angle to each other. The test system needs to make sure that:
1. The spark happens for a given minimum amount of time
2. The spark happens in the correct order
3. The angular position of the shaft when the spark occurs is correct
4. And the time difference between each spark remains constant.
The test system needs to be able to measure the items in bold above. The amount of time, angular position and time difference (items 1, 3 and 4) can each be measured using a counter. To ensure correct firing order (item 2), a digital change detection measurement needs to be made which records the entire port state whenever any of the lines in it changes state.

## Slide 53: Timestamping Explained
Record value of counter and lines being monitored each time a line changes state
Use counter as a timebase
Select individual lines their polarity
Counter
Line A
Line B
1
2 3 4 5 6 7 8 9 10
3
10
11
5
8
01
no
effect
Time data (counts)
Digital data (AB)
Notes: Timestamping is a way of acquiring digital data on a “as needed” basis. Data and time is read only when any of the monitored digital lines changes state. The time is kept using a counter as a timer. The measurement returns data in time/data pairs. Each data pair indicates the state of the lines and the time that change occurred. You can select which digital lines to timestamp and the active edges (rising, falling or both) at which to timestamp.
When using timestamping, less data is acquired overall since data points are read only when a change occurs. This can lead to less bus activity, better data throughput and a faster overall performance.
In the diagram above, both Line A and Line B can cause a timestamp. Line A causes a timestamp on both rising and falling edges. Line B causes a timestamp only on rising edges. The recorded data shows that AB were 10 at count equal to 3, 11 at count equal to 5 and 01 at count equal to 8.
Digital logic analyzers use the timestamping technique to record the states of the monitored digital lines. This technique can also be applied to any digital measurement acquiring lots of data.

## Slide 54: Solution – Synchronize Digital Input and Counter Measurements
counter 0
counter 1
80 MHz
PXI-6533
(32-bit high-speed digital I/O device)
PXI-6602 (8-channel counter device)
Quadrature
encoder
Both signals
shared via PXI
trigger (RTSI) lines
4
Common start-trigger
signal
“Change detected”
or latch signal
Notes: To measure the firing order, each of the spark plug firing signal is connected to a separate digital line of a PXI-6533 (32-bit high speed digital input/output) device. To track the position of the rotating shaft, a quadrature encoder connected to the shaft is wired to a counter (counter 0) of a PXI-6602 (8-channel counter device). To track the firing time and the time elapsed between each spark, another counter (counter 1) is used as a timer with its source connected to an internal 80 MHz timebase.
The PXI-6533 and the PXI-6602 are used together to perform timestamping of digital data. The PXI-6533 is configured to latch spark plug data whenever any of its input lines changes state. Also, a “change detected” signal is passed from the PXI-6533 to the counters on the PXI-6602 device. Whenever the PXI-6533 detects a change in any of its input lines state, it produces a pulse on its “change detected” line which causes both the counters to latch (store) their value onto their corresponding software buffers.
Two things need are required to synchronize the digital and counter measurements. First, both the devices share a common start trigger signal to start the digital and counter measurements at the same time (see above). Second, both the devices share a common signal (change detected) which causes them to store their current values into their corresponding software buffers.

## Slide 55: Analyze Data for the Car Engine Timing Test
Spark delayed
by 1 ms
and off by
90 degrees
Notes: The above table shows a possible data set for the car engine timing test. The first column shows which spark plugs are fired (‘1’ indicate spark plug is on, ‘0’ indicates spark plug is off). Column two shows actual encoder data read by the counter. It assumes that there are 1,000 encoder counts per revolution, so an increment of 250 indicates a 90 degree (360 degrees * 250/1000) revolution. Column three shows the same encoder data transformed to degrees (0 through 360). Column four shows the relative time in milliseconds between each of the spark plug event.
The bold data shows a possible problem with the timing system. It indicates that spark plug 3 was fired 2.05 ms after the previous firing instead of after 1.05 ms. In addition, the shaft position was at 90 degrees instead of being at 0 degrees when this happened.

## Slide 56: GPS Background
GPS stands for Global Positioning System
GPS satellites
  Periodically transmit their position and time
  Have multiple atomic-clocks which keep very accurate time (better than 10E-13 seconds accuracy!)
  Orbits the Earth once every 12 hours
There are more than 24 GPS satellites orbiting the Earth
Notes: The Global Positioning System (GPS) is a constellation of 24 satellites and a ground control system operated by the US Air Force. GPS provides precise position and time information throughout the world with very high accuracy.

## Slide 57: GPS Acronyms
GPS = Global Positioning System
IRIG-B = timing encoded signal generated by GPS receivers
PPS = pulse per second. Signal generated by GPS receivers
UTC = Universal Coordinated Time (almost equal to GMT)
Notes: IRIG stands for Inter Range Instrumentation Group.
UTC – a time system adopted by many countries in 1972. UTC is based on the weighted combination of atomic clocks located around the world. UTC occasionally changes by the addition of leap seconds.

## Slide 58: GPS Receiver Background
Provide very accurate global time and position
  All times are relative to UTC
Time accuracy improves with time!
  From 10E-6 seconds at power up to 40E-9 seconds at steady state
Can transmit time information via
  RS-232 (serial port)
  PPS timing signal
  IRIG-B timing signal

## Slide 59: GPS Timing Signals
1 Second
Sample PPS signal
Seconds Minutes Hours Days
Time of year
(repeats every second)
Sample IRIG-B signal (simplified)
Notes: IRIG-B time format consists of frames that repeat at exactly one second intervals. Each frame has detailed time information in terms of days, hours, minutes and seconds since the beginning of the current year.
There are multiple IRIG-B standards, but two of them stand out – Amplitude modulated analog IRIG-B and pulse-width modulated DC Level Shift (DCLS) IRIG-B.
PPS time format consists of pulses that repeat at exactly one second intervals. This format does not have the detailed time information which the IRIG-B format has. However, most GPS receivers have a serial port through which you can read this detailed time information.

## Slide 60: Selecting a GPS receiver
Need to measure time or both time and position?
Need occasional position data or something to allow accurate position steering?
Need accurate velocity measurements?
Consider cost versus accuracy tradeoff
Concerned about power consumption
Considerations for high acceleration, high velocity or extreme temperature
Notes: Commercial GPS receivers range from about $2000 for low-end devices having microsecond time accuracy to thousands of dollars for sophisticated, high-end, flexible, nanosecond time accuracy devices. You can buy a rack-mountable, GPS receiver with 100 ns accuracy for about $3000. See the reference section for links to GPS receiver vendors Web sites.

## Slide 61: Synchronizing Distributed Measurements
GPS receiver
(San Francisco)
GPS receiver
(Los Angeles)
Time data
PPS or
IRIG-B
timing
signal
PXI-6608
8-channel counter
device
PXI-6608
8-channel counter
device
GPS satellites
Notes: Sometimes you want to synchronize measurements in more than one location. These measurement systems can be located at opposite ends of a building, separate buildings in a city, different cities miles away or continents thousands of miles apart. You can use your computer clock when you need timing accuracy to within a second. For higher accuracy, up to 40 ns, you need to synchronize your measurements directly to a GPS satellite at each remote data-acquisition site.
For example, let’s say we are trying to track how an earthquake travels from San Francisco to Los Angeles. You can set up a distributed measurement system as shown above. At each location, set up a PXI-6608 (8-channel counter device with very accurate oscillator) device connected to a GPS receiver. The PXI-6608 has a real-time clock which synchronizes itself to an IRIG-B or PPS signal generated by the GPS receiver. The GPS receiver also locks itself to one or more satellites. Two different GPS receivers can be within 40 ns relative time to each other once locked to satellites. So, once the PXI-6608 is synchronized to the satellite time, you can use it as a master clock to timestamp any data recorded at that site. To do that, you can use the hardware signal indicating the start of the measurement at each location to latch the value of the real-time clock.
You can use distributed synchronization in a wide variety of applications including monitoring power distribution, distributed control, traffic monitoring, geological, environmental research, and more.

## Slide 62: GPS References
datum.com
truetime.com
nist.gov
“GPS – A Guide to the Next Utility”, Jeff Hurn, Trimble Navigation

## Slide 63: Conclusion
Timing is an important part of measurement systems and affects measurement accuracy
Triggering and clock sharing allow synchronized multiple measurements
Use timing, triggering, and synchronization to construct complex measurement systems