## Slide 1: Measurement Acquisition and Analysis Techniques with LabVIEW™
<date/time>
Michael Cerna
DSP Measurement Analysis
Fri Aug 18
10:15 a.m.-1:15 p.m., 1:45-4:45 p.m.
Exhibit (3I)

## Slide 2: Agenda
Acquisition for measurement
  Avoiding common pitfalls
Measurements in LabVIEW
  Waveform datatype, measurement analysis
Measurement analysis modules
  DC and RMS
  Frequency analysis, tone detection
  Distortion analysis, THD, SINAD
  Triggering, limit testing
  Signal generation
Notes:   Acquisition for Measurement – we will begin by covering important issues related to bringing the measurement data into the computer. We must acquire the signal appropriately to ensure measurement accuracy.
  Measurements in LabVIEW – most measurement applications include three distinct tasks, including acquisition, analysis, and presentation. As software for measurement, LabVIEW 6i includes tools for all three of these tasks with specifics for each.
  Measurement analysis modules – this is the bulk of the course – introducing fundamental theory behind common measurements like DC and RMS, averaged frequency domain analysis, tone detection, distortion analysis, software triggering, and limit testing. We will step through the process of making these measurements on real-world signals in LabVIEW, including tweaking parameters for speed or accuracy, and through the process uncover some interesting characteristics hidden in the signal.

## Slide 3: Acquisition for Measurement
Choosing the frequency range
  Sampling rate to avoid aliasing
Choose coupling
  AC coupling to remove offsets
  DC coupling for DC measurements
Choosing the amplitude range
  Voltage range to maximize dynamic range
Choosing the type of connection
  Differential or NRSE for grounded signals
  Differential, NRSE, or RSE for floating signals
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
You can minimize measurement errors, and thus obtain a better (more accurate) measurement, by properly choosing the frequency and amplitude ranges for your DAQ board, the type of coupling, and the type of connection.
Proper selection of the frequency range is necessary to avoid aliasing.
AC coupling is used to remove DC offsets.
DC coupling is used for DC (or low frequency) measurements.
Proper selection of the amplitude range can help to maximize dynamic range.
You choose the appropriate connection depending on whether your source is floating or grounded.

## Slide 4: Sampling
Samples of an analog signal
Discrete points in time
Sampling frequency – number of samples per second
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
An A/D converter takes samples of an analog signal at discrete points in time.
The number of samples taken per second is known as the sampling frequency.

## Slide 5: Aliasing
Minimum two samples per cycle
Alias – higher frequency appears as a lower one
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
You need to sample the analog signal at least twice per cycle.
In the figure above, the samples of both the higher frequency signal and the lower frequency signal occur at the same instants in time, and have the same amplitude values. Thus, looking at the samples alone, one cannot tell whether the samples came from the higher frequency signal or the lower frequency signal. You can avoid this by sampling the higher frequency signal at least twice per cycle, so that the samples obtained will not be confused with those of a lower frequency signal.

## Slide 6: Choosing the Sampling Frequency
Know the frequency (range) of your input signal
To avoid aliasing
  Sensor or system should be naturally band-limited
  Anti-aliasing filter
Sampling frequency > two times maximum input frequency
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
To choose the proper sampling frequency to avoid aliasing, it is useful to know the highest possible frequency present in the input signal. Knowing the frequency response of your sensor or system can give you a good idea of the highest frequency that can pass through it to the A/D converter. Some DAQ boards incorporate an anti-aliasing filter – a lowpass filter that attenuates signals beyond a certain frequency. Make sure you choose a sampling frequency at least two times the maximum frequency present in the signal.

## Slide 7: Dynamic Range
Range (in db) of signal levels that you can measure simultaneously
Ability to measure small signals in the presence of large signals
Determines the amplitude range over which you make measurements
Amplitude
Dynamic Range
Frequency
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Connect two signals to the input of your instrument (or a single signal composed of the addition of two tones)…one at the maximum allowable amplitude for the specified input range (full scale), another that is smaller. Reduce the amplitude of the smaller signal till it is just measurable. The ratio (in dB) of the two amplitudes will give you the dynamic range of your instrument.
Dynamic range depends on the following –
  Number of bits in the A/D converter
  Noise floor – due to the internal noise of the instrument
  Harmonic distortion of the larger amplitude signal due to any nonlinearities of the instrument

## Slide 8: Choosing the Amplitude Range
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
The figure on the top left shows the transfer function of a 3-bit analog-to-digital converter (ADC). If the amplitude range of your DAQ board is set from -4 to +4 Volts, and the input signal (bottom plot) also occupies the range from -4 to +4, you are using the full range (2^3 = 8 levels) of the ADC to represent your input signal. The output of the ADC is shown on the top right. Observe that all eight levels of the ADC are being utilized.

## Slide 9: Choosing the Amplitude Range
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Now the range of the DAQ board has been set from -8 to +8 Volts. The corresponding transfer function of the ADC is shown in the top left plot. However, the input signal (lower plot) only spans the range between -4 to +4 Volts. In this case, you are not using the full range of the ADC to represent your input signal. The output of the ADC is shown on the top right. Observe that only four levels of the ADC are being utilized.

## Slide 10: Choosing the Type of Connection
Measurement
System
Signal Source
Grounded
Floating
Differential
Referenced Single-Ended (RSE)
Non-Referenced Single-Ended (NRSE)
Vs
Vm
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Source
Vs is the source voltage. The source is either grounded or floating.
Grounded voltage signals are referenced by a system ground, such as earth or building ground. Thus they share a common ground with the DAQ board. Examples of grounded sources include devices such as signal generators and power supplies that plug into the building ground through wall outlets.
Floating voltage signals are not referenced to any common ground. Examples of floating sources include batteries, thermocouples, transformers, and isolation amplifiers.
Measurement System
Vm , the voltage measured by the measurement system, is either:
  Differential – neither input is connected to a fixed reference such as earth or building ground. The differential measurement system reads the potential difference between its two input terminals. Example - DAQ boards with instrumentation amplifiers.
  Referenced single-ended (RSE) – this is a grounded measurement system. The measurement is made with respect to the building ground.
  Nonreferenced single-ended (NRSE) – as for an RSE measurement system, all measurements are made with respect to a common reference. However, the common reference voltage is different from the earth or building ground.
If your signal source is grounded use a differential or NRSE measurement system. This is because for an RSE system, there can be a slight voltage difference between the system ground and the signal ground. This potential difference is called a ground loop and it results in erroneous data by introducing both AC and DC noise into the measured data in the form of a power-line frequency component (AC), and offset errors (DC).
If your signal source is floating you can use either differential, RSE, or NRSE measurement system.
There is much more to be said on this topic. For further details, see the LabVIEW Data Acquisition Course manual or the LabVIEW Data Acquisition Basics manual.

## Slide 11: Acquisition Summary
Choose a sampling frequency > twice the maximum frequency present in your signal
Select the amplitude limits to get the highest possible dynamic range
Make the proper connection depending on the source (grounded or floating)

## Slide 12: Measurements in LabVIEW
LabVIEW 6i addresses the three components of measurement:
  Acquisition
    Measurement and Automation eXplorer (MAX) integration
  Analysis
    Measurement analysis tools
  Presentation
    Internet, report generation, graphics
Waveform datatype touches all three
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Most measurement applications include three distinct tasks: acquisition, analysis, and presentation. As software for measurement, LabVIEW 6i includes tools for all three of these tasks with specifics for each. It also includes general-purpose measurement tools (such as the waveform datatype) that you can apply to all three tasks.
The following points list some highlights and the next several slides go into more detail.
Acquisition – Improved integration with the Measurement and Automation Explorer (MAX)
Analysis – A set of 33 new tools (VIs) for analysis related to measurement
Presentation – Internet support, report generation, 2D/3D/Digital/XY graphics

## Slide 13: LabVIEW 6i Acquisition
MAX integration
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
The best way to learn about some of LabVIEW 6i’s features for measurement is to discuss (and build) the LabVIEW visual program that defines a simple measurement system.
Let’s begin by assuming that you want to acquire a signal using a National Instruments data acquisition device, such as one of our PCI plug-in MIO (multi-purpose input/output) cards. After installing the NI-DAQ Driver Software and connecting the board to your PC, you should be able to open up the MAX configuration window (right, rear window on the slide) and see a listing for your board under “Devices and Interfaces” on the left pane. You can then configure one or more measurement channels on that board by right clicking the “Data Neighborhood” folder on the left and choosing “New Channel.” A series of dialog boxes will then guide you through the process of setting up the channel.
  After configuring one or more channels, you’ll next want set up LabVIEW to acquire data from that channel.
  Place an AI Acquire Waveform vi on a new diagram.
  To associate acquisition with your particular board, right click on the device input of AI Acquire Waveform with the wiring tool and choose Create Constant. Make sure that the number matches the device number of your device.

## Slide 14: LabVIEW 6i Acquisition (Notes)
Notes continued from previous slide
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
To associate a channel with your acquisition, right click on the Channels (0) input of AI Acquire Waveform with the wiring tool and choose Create Constant. Switch to the Finger tool and left click on this new constant. Notice that you can select from among the channels that you have configured with MAX.
Finally, specify an acquisition rate and number of samples using Constants.

## Slide 15: LabVIEW 6i Analysis
The Waveform datatype is the standard output of high-level Analog Input tools from NI-DAQ
Several palettes of tools are devoted to it, including:
  The Waveform Measurements palette contains common analysis tools
  Waveform Operations palette includes simple waveform operations such as Add, Select Subset, and Peak Values
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
After setting up an acquisition, consider how to analyze or process the waveform. You can find common tools for these jobs as sub-palettes of the Analyze palette. Examples include Power Spectrum, DC/RMS, tone detection, and digital filtering.
The Thick Brown wire that you used to connect the acquisition to the analysis VI in the last step represents the Waveform Datatype. This tool defines names, unit information, and timing characteristics of measured signals. By combining these and other elements, the Waveform Datatype simplifies your work (and your LabVIEW diagrams). Analysis and Visualization tools employ the elements of the Waveform Datatype such as timing information and units as input parameters for analysis and for graphics settings.

## Slide 16: LabVIEW 6i Analysis
Examine the power spectrum of a waveform
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Performing a power spectrum and graphing the results is easy:
Choose FFT Power Spectrum from AnalyzeWaveform Measurements and drop it to the right of the AI Acquire Waveform. Use the wiring tool to connect the waveform (upper right) output of AI Acquire waveform to the Time Signal input of FFT Power Spectrum.

## Slide 17: Waveform Analysis
33 new DSP Measurement Analysis tools work with the new waveform datatype
Highlights – DC/RMS, tone detection, spectral analysis, limit testing, digital filtering, waveform generation
Online examples accessible through Help and in a library – \LabVIEW\Examples\Measurement\maxmpl.llb
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
This presentation will discuss implementation details of some of LabVIEW 6i’s 33 new DSP Measurement analysis tools.

## Slide 18: LabVIEW 6i Presentation
LabVIEW 6i’s presentation tools work with the waveform datatype – 2D / 3D Graphs, Charts, Tables, and other indicators
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Go to the front panel of your LabVIEW VI and add a Graph. Return to the Diagram and wire the output of the FFT Power Spectrum to the icon of the graph.
When you run the VI, you should see a Power Spectrum of the input channel that you selected on the graph.

## Slide 19: Measurement Analysis Modules
Viewing the signal
DC/RMS
Frequency analysis
Tone detection
Distortion
Filtering
Triggering
Limit testing
Measurement recap
Signal generation

## Slide 20: Viewing the Signal
Using a graph to display the acquired waveform
  Continuous acquisition (in a loop)
  Time stamp t0
  Time resolution dt
  Visual evaluation of the signal

## Slide 21: DC / RMS Measurements
DC/RMS definitions
Averaging for DC/RMS measurements
Windowing for DC/RMS measurements

## Slide 22: DC Value – Definition
Continuous signal
Voltage
Time
t1
t2
(t2 - t1) represents the
Integration time or
Measurement time
Notes: Integrated DC value –
You can measure a more accurate value by averaging. In a continuous signal, the averaged value between two specific times (t1 and t2) is defined as the “signal integration between t1 and t2, divided by the measurement time, that is t2-t1”.
The area between the averaged value Vdc and the signal that is above Vdc, is equal to the area between Vdc and the signal that is under Vdc.

## Slide 23: DC Value – Sampled Signal
Sampled signal
Voltage
Time
p1
p2
….
pn
Notes: For a sampled system, the DC value is defined as the mean value of the samples acquired in the specified measurement time window.

## Slide 24: RMS Value – Definition
Continuous signal
Sampled signal
Notes: Root mean square is defined as the square Root of the Mean value of the Squared signal.
Note that where DC measurements can result in both positive and negative results, RMS measurements will always give positive results expect if your signal is exactly “null” resulting in an RMS value of zero.

## Slide 25: DC / RMS Measurements
Averaging
  Measurement time
  Linear averaging
  Exponential averaging
Applying a window
  Hann window
  Low-side-lobe window
Notes: Depending on your signal, you may or may not want to use a window for making DC/RMS measurements.
If you know the exact frequency of the sine tone contained in your signal, use a measurement time that contains an integral number of cycles of the sine tone. If you do not know the frequency of the sine tone, apply a window, such as the Hann window to significantly reduce the measurement time to achieve a specified accuracy.
If you have a noisy signal, do not use a window. In this case, you can increase the accuracy of your measurement by increasing the measurement time, or by applying a filter to your signal (to reduce the noise) before making the measurement.
You can also average the results of several measurements.

## Slide 26: Frequency Analysis
Based on the Fast Fourier Transform (FFT)
Dot product of signal with cosines & sines
FFT
Time
Domain
Signal
Frequency
Domain
Signal
Notes: Fourier’s theorem states that any waveform in the time domain can be represented by the weighted sum of sines and cosines. The same waveform can then be represented in the frequency domain as a pair of amplitude and phase values at each component frequency.
The Fast Fourier Transform (FFT) transforms digital samples from the time domain into the frequency domain. Each frequency component is the result of a dot product of the time signal with the complex exponential at the component frequency.
The DC component is the dot product of x(n) with [cos(0)-jsin(0)], or with 1.0.
The first “bin”, or frequency component, is the dot product of x(n) with cos(2n/N)-jsin(2n/N). Here, cos(2n/N) is a single cycle of the cosine wave, and sin(2n/N) is a single cycle of a sine wave.
In general, the real part of FFT bin k, Re[X(k)], is the dot product of x(n) with k cycles of the cosine wave, and the imaginary part of bin k, Im[X(k)], is the dot product of x(n) with k cycles of the sine wave.

## Slide 27: Frequency Analysis – Fmax, f
Highest frequency which you can analyze
Frequency resolution
fs : sampling frequency
T : total acquisition time
N : FFT block size
Notes: The use of the FFT for frequency analysis implies two important relationships.
The first one relates the highest frequency which can be analyzed to the sampling frequency (see Shannon sampling theorem before) .
The second one links the frequency resolution to the total acquisition time, which is related to the sampling frequency and the block size of the FFT.

## Slide 28: Frequency Analysis Tools
FFT gives magnitude and phase information (i.e. complex)
  Magnitude = sqrt(Real^2 + Imag^2)
  Phase = Tan-1(Imag / Real)
Power spectrum reflects the energy content
  Power spectrum = Mag^2
Notes: The FFT spectrum output is complex, that is, every frequency component has a magnitude and phase. The phase is relative to the start of the time record, or relative to a cosine wave starting at the beginning of the time record. Single channel phase measurements are stable only if the input signal is triggered. Dual-channel phase measurements compute phase differences between channels so that if the channels are simultaneously sampled, triggering is usually not necessary.
Normally the magnitude of the spectrum is displayed. The magnitude is the square root of the FFT times its complex conjugate, or the square root of the sum of the squares of the real and imaginary parts of the FFT.
The phase is the arctangent of the ratio of the imaginary and real parts of the FFT, and is usually between  and -  (between 180 and -180 degrees).

## Slide 29: Frequency Domain Averaging
RMS averaging
  Weighted mean of squared magnitude
  Reduces fluctuations, not noise
Vector averaging
  Real and imaginary part averaged separately
  Reduces un-correlated noise
  Single channel requires trigger
Peak hold
  Maximum magnitude on a bin-by-bin basis
Notes: Averaging successive measurements tends to improve measurement accuracy. Averaging is usually performed on measurement results or on individual spectra, but not directly on the time record.
RMS averaging computes the weighted mean of the sum of squared values. RMS averaging reduces fluctuations in the data but does not reduce the actual noise floor.
Vector averaging uses the complex FFT spectrum. The real part is averaged separately from the imaginary part. This can reduce the noise floor for random signals since they are not phase coherent from one time record to the next. For single channel measurements vector averaging requires a trigger so that the real and imaginary parts of the signal will add in phase instead of canceling randomly.
Peak hold averaging retains the peak of the spectral magnitudes on a frequency bin by bin basis. The peak values are stored in the original complex form so that all complex forms of the data can be displayed.

## Slide 30: Weighting Modes for Averaging
Linear weighting
  Combines N spectral records with equal weights
  For stationary signals
Exponential weighting
  Emphasizes new spectral data more than old
  For non-stationary signal
  Continuously averaged
Notes: Each new spectral record for RMS and Vector averaging can be weighted using either Linear or Exponential weighting.
Linear weighting combines N spectral records with equal weighting. When the number of averages has been completed, the analyzer stops averaging and presents the averaged results. Exponential weighting emphasizes new spectral data more than old.
Averaging takes place according to this formula –
New Average = New Spectra * 1/N + Old Average * (N-1)/N
where N is the number of averages for linear averaging, or is computed from the time constant for exponential averaging.

## Slide 31: Tone Detection
Detecting the frequency, amplitude, and start phase of an acquired sine waveform
Can detect a single tone or multiple tones
Fast and accurate results
Noise resistant

## Slide 32: Tone Detection (in LabVIEW 6i)
Input is a WDT but internally the detection is performed in the frequency domain
Finds the highest tone within a search range
Exports different results options
    Detected tone/spectrum
    Residual tone/spectrum

## Slide 33: Distortion Analysis (THD)
Total Harmonic Distortion (THD)
  Ratio between RMS(Harmonics) and RMS(fundamental)
Notes: If a signal of a specific frequency is applied at the input of a linear system, the output of the system will be at the same frequency, but of possibly different amplitude and/or phase. However, if a signal is applied at the input of a nonlinear system, the output of the system will be distorted and will contain frequencies that were not originally present at the input.
There are several measurements that you can make to measure the amount of distortion introduced by the nonlinear system. One of these measurements is the Total Harmonic Distortion (THD) which is the ratio between the RMS value of the harmonics to the RMS value of the fundamental (the frequency applied at the input of the nonlinear system).

## Slide 34: Distortion Analysis (SINAD)
Signal in Noise and Distortion (SINAD)
  Ratio between RMS(Signal) and RMS(Distortion + Noise)
Notes: Another measure of distortion is SINAD, which additionally takes into account the noise present in the output signal. It is the ratio of the RMS value of the signal (fundamental + harmonics + noise) at the output of the nonlinear system to the RMS value of the harmonics and noise (without the fundamental).

## Slide 35: Distortion in LabVIEW 6i
Input is a WDT. Distortion Components detected using the Extract Single Tone Information VI
Finds the highest Tone within a Search Range and its Harmonics (for THD)
Exports different results options
    Detected tone/spectrum, residual tones/spectrum, harmonics only and so on
Notes: LabVIEW 6i contains the following VIs for measuring the THD and SINAD.
Harmonic Distortion Analyzer.VI
SINAD Analyzer.VI

## Slide 36: Digital Filtering
Finite impulse response (FIR) vs. infinite impulse response (IIR)
Topology/design method
Type
Cutoff frequency(s)
Order
Notes: Digital filtering is generally applied when two signals are to be separated, and they are separable by frequency. The first decision to make is whether the Filter should be FIR or IIR.
FIR vs. IIR – If linear phase (constant group delay) is a requirement, then the filter must be FIR. If linear phase is not a requirement, then IIR is generally preferred due to lower computational overhead resulting from a lower order for a desired characteristic as compared to FIR filters. IIR designs are limited to lowpass, highpass, bandpass or bandstop. If a more arbitrary magnitude response is desired then a Parks-McClellan designed FIR filter is required.
Topology – the next design choice. With IIR filters this means a choice between Butterworth, Chebyshev, inverse Chebyshev, Elliptic, and Bessel. Butterworth filters yield a flat magnitude response and are preferred if ripple in the passband or stopband is not allowed. Because ripple is not allowed, a Butterworth filter rolls off at a slower rate than other designs. Chebyshev filters have ripple in the passband, but no ripple in the stopband. Inverse Chebyshev filters allow ripple in the stopband but not in the passband. Because both filters allow ripple, they have sharper roll-off. Elliptic filters allow ripple in both the passband and the stopband. Because of this ripple the Elliptic design has the sharpest roll-off. Deciding between the various IIR topologies is, generally, then, a decision about ripple. If no ripple is allowed, choose Butterworth. If passband ripple is allowed, choose Chebyshev. If stopband ripple is allowed, choose Inverse Chebyshev. If both passband and stopband ripple are allowed, choose Elliptic.
For FIR filters it means a choice between Equiripple, Windowed, or Parks-McClellan. The Equiripple yields an equal ripple magnitude in both the passband and stopband. The Windowed filter has the advantage of being very simple and fast, but yields a filter that has increasing ripple near the cutoff frequency(s). Generally equiripple is preferred.
Type – the type parameter allows a choice between Lowpass, Highpass, Bandpass, and Bandstop.
Cutoff frequency(s) – define the edges or boundaries of the passband and stopband regions. Cutoff frequency is not necessarily the –3dB point of the filter, but is defined in the context of the filter topology. For example, the cutoff frequency for a Butterworth filter is the –3dB point. For a Chebyshev filter, the cutoff frequency is the point at which the passband ripple exceeds the desired specification. For an inverse Chebyshev filter, the cutoff frequency is the point at which the filter response is always below the desired stopband attenuation.
Order – The order of the filter (along with passband/or stopband ripple) determines how much attenuation is possible at cutoff frequencies. Often it is not known what the order of the filter should be, and determining the filter order is an iterative process. A more useful way of specifying the filter is to specify the gain at the passband and stopband cutoff frequencies. LabVIEW allows this alternate specification, and uses it to determine the cutoff frequencies and necessary order.

## Slide 37: Digital Filtering in LabVIEW 6i
Waveform datatype (WDT)
IIR vs. FIR
Polymorphic
Reset
Parameter change
Notes: LabVIEW supports filtering of both arrays and Waveforms (WDT). We will discuss the WDT filtering VIs available. There are two VIs available on the Waveform Conditioning palette. They are Digital FIR Filter.vi and Digital IIR Filter.vi. The Digital IIR Filter.vi implements all the various IIR topologies mentioned in the previous slide, while the Digital FIR Filter implements the Equiripple and Windowed designs. These VIs are polymorphic, meaning that the VI changes depending on what is wired to the inputs of the VI. All the VIs are polymorphic on the signal in and filter parameters inputs. For example, the Digital IIR filter is really a mask or template that represents three separate VIs. The first VI takes a set of filter specifications and a signal waveform and filters the waveform. The second VI takes an array of waveforms and filter each according to the same specified filter, with separate filter states being maintained for each waveform. The last VI filters an array of waveforms according to an array of filter specifications (of the same size as the array of waveforms). The behavior of the Digital FIR Filter.vi is the same with regard to polymorphism.
These VIs have a reset input that allows you to force a redesign of the filter coefficients and a reset of the filter states to zero, but wiring this input is optional. LabVIEW 6.0i has a new primitive, namely the “First Call?” primitive, that allows subVIs to be aware of whether they are being called for the first time or not. So if the filter VI is being called for the first time then the filter is designed and the filter states initialized to zero. Subsequent calls to the VI (I.e. calling the VI in a loop) will be correctly identified as such, and will not implicitly cause a redesign or reset of the filter, unless the design parameters have changed or the incoming waveform is not continuous.
Parameter changes will therefore cause a redesign of the filter coefficients. Parameters are the required and optional clusters of filter specifications, and the dt value in the waveform. If the waveform dt changes from one iteration to the next, then the filter VI will redesign itself automatically and issue a warning that it has done so. If the waveform being acquired is not continuous from block to block, then the filter will automatically reset internal filter states to zero, but will not redesign filter coefficients.

## Slide 38: Triggering
Level based software triggering
  Account for channel hysteresis
  Detect trigger on rising or falling edge
Applications
  Use when hardware triggering is not available
  Analyze pre-trigger data
  Trigger on a signal with noise or multi-tone signal
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Basic Level Trigger Detection VI detects trigger point in an input waveform. You can specify a hysteresis value (h) for the channel, level (l) and slope (rising or falling edge). We define a low threshold level = l - h/2 and a high threshold level = l + h/2. On a rising edge, the VI detects a trigger when there is a low level to high level transition. On a falling edge, the VI detects a trigger when there is a high level to a low level transition. The trigger detected? output of the VI is TRUE if a trigger is detected and is FALSE if no trigger is detected. The trigger index output contains the array index at which the trigger is detected. You can perform triggering on single channel or multi-channel data.
Software trigger detection is extremely useful when a data acquisition board does not support triggering. If the hardware supports triggering, you can usually only acquire the post-trigger data. If you want to analyze the pre-trigger data, you will find software triggering to be useful. Software triggering is particularly useful for detecting trigger in a signal with noise or in a multi-tone signal. You can acquire such a signal, extract a single tone out using the VIs discussed earlier and then perform software triggering on that particular tone.
Another important application of software triggering is to “condition” the signal for Frequency Vector Averaging. Vector averaging requires triggered signals to ensure that the desired signals are phase coherent, thus reducing phase incoherent (noise) signals.

## Slide 39: Limit Testing
Limit Specification
  Also known as mask testing
  Create continuous or segmented limits
  Specify limits as number or formula
Perform Limit Testing
Applications
  Telecom pulse mask tests
  Analysis of filter design
Notes: SC3C-<number>
Measurement Acquisition and Analysis in LabVIEW
<header>
<date/time>
Limit Testing is also referred to as mask testing in a number of applications. You can use the Limit Specification VIs to create limits. You can specify the Y axis values for the limits in terms of numbers or formula. The X axis values for the limits can only be specified in terms of numbers. You can create either continuous limits or segmented limits. You can use the Limit Testing VI to perform limit testing. You can create the limit once and then use this VI continuously. The test passed? output of the Limit Testing VI is TRUE if the test signal falls within the region (mask) defined by the limits. It is FALSE otherwise. The output values output is an array that contains the test signal, the data points which fall outside the limits and the upper and lower limits. You can wire this output to a graph for easy visualization.
Limit or Mask testing is used in a wide range of test and measurement applications. For example, ADSL modem manufacturers use it to ensure that the power spectral density of ADSL signals meet the recommendations laid out in the ANSI T1.413 specification. Consider a second application in the area of digital filter design. Your aim is to design a low pass filter with a passband ripple of –10dB and stopband attenuation of 60 dB. You can use limit testing to ensure that the frequency response of the filter always meets these specifications.

## Slide 40: Measurement Recap
Viewing the Signal
DC/RMS
Frequency Analysis
Tone Detection
Distortion
Filtering
Triggering
Limit Testing
Measurement Recap
Signal Generation

## Slide 41: Signal Generation
Common test signals
Multitone
Noise
Notes: Common test signals form the basis for many tests and are used to measure the response of a system to a particular stimulus. Some of the common test signals are: Sine, Square, Triangle, Sawtooth, Pulse, Sinc, Impulse, and Ramp.
The common test signals, except for the sine wave, do not allow full control over their spectral content. For example, the harmonic components of a square wave are fixed in frequency, phase, and amplitude relative to the fundamental. On the other hand, multitone signals can be generated with a specific amplitude and phase for each individual frequency component. A multitone signal is the superposition of several sine waves or tones, each with a distinct amplitude, phase, and frequency.
A multitone signal is typically created so that an integer number of cycles of each individual tone are contained in the signal. If an FFT of the entire multitone signal is computed, then each of the tones falls exactly onto a single frequency bin. This means there is no spectral spread or leakage. For this reason a multitone signal is self windowing, and does not require the application of a window. Multitone signals are a part of many test specifications and allow the fast and efficient stimulus of a system across an arbitrary band of frequencies. Multitone test signals can be used to determine the frequency response of a device, and with appropriate selection of frequencies, can also be used to measure such quantities as intermodulation distortion.

## Slide 42: Signal Generation (Notes)
Notes: Noise signals may be used to perform frequency response measurements, or to simulate certain processes. Several types of noise are typically used, namely Uniform White Noise, Gaussian White Noise, and Periodic Random Noise.
The term white in the definition of noise refers to the frequency domain characteristic of noise. Ideal white noise has equal power per unit bandwidth, resulting in a flat power spectral density across the frequency range of interest. Thus, the power in the frequency range from 100 Hz to 110 Hz is the same as the power in the frequency range from 1000 Hz to 1010 Hz. In practical measurements, to achieve the flat power spectral density would require an infinite number of samples. Thus, when making measurements of white noise, the power spectra are usually averaged, with more number of averages resulting in a flatter power spectrum.
The terms uniform and Gaussian refer to the probability density function (PDF) of the amplitudes of the time domain samples of the noise. For uniform white noise, the PDF of the amplitudes of the time domain samples is uniform within the specified maximum and minimum levels. Another way to state this is to say that all amplitude values between some limits are equally likely or probable. Thermal noise produced in active components tends to be uniform white in distribution.
Periodic Random Noise (PRN) is a summation of sinusoidal signals with the same amplitudes but with random phases. It consists of all sine waves with frequencies that can be represented with an integral number of cycles in the requested number of samples. Since PRN contains only integral-cycle sinusoids, you do not need to window PRN before performing spectral analysis because PRN is self-windowing and therefore has no spectral leakage.

## Slide 43: Multitone Signals
Crest factor
Phase
Notes: Crest Factor – the relative phases of the constituent tones with respect to each other determines the crest factor of a multitone signal with specified amplitude. The crest factor is defined as the ratio of the peak magnitude to the RMS value of the signal. For example, a sine wave has a crest factor of 1.414:1. For the same maximum amplitude, a multitone signal with a large crest factor contains less energy than one with a smaller crest factor. Another way to express this is to say that a large crest factor means that the amplitude of a given component sine tone is lower than the same sine tone in a multitone signal with a smaller crest factor. A higher crest factor results in individual sine tones with lower signal-to-noise ratios. Proper selection of phases is therefore critical to generating a useful multitone signal. To avoid clipping, the maximum value of the multitone signal should not exceed the maximum capability of the hardware that generates the signal. This places a limit on the maximum amplitude of the signal. You can generate a multitone signal with a specific amplitude by different combinations of the phase relationships and amplitudes of the constituent sine tones. It is usually better to generate a signal choosing amplitudes and phases that result in a lower crest factor.

## Slide 44: Multitone Signals (Notes)
Notes: Phase – there are two general schemes for generating tone phases of multitone signals. The first is to make the phase difference between adjacent-frequency tones vary linearly from 0 to 360 degrees. This allows the creation of multitone signals with very low crest factors, but the multitone signals then have some potentially undesirable characteristics. This sort of multitone signal is very sensitive to phase distortion. If, in the course of generating the multitone signal, the hardware or signal path induces non-linear phase distortion, then the crest factor can vary considerably. In addition, a multitone signal with this sort of phase relationship may display some repetitive time domain characteristics that are possibly undesirable.
Another way to generate the tone phases is to vary them randomly. As the number of tones increases the multitone signal will have amplitudes that are nearly Gaussian in distribution. In addition to being more noise-like, this signal is also much less sensitive to phase distortion. Multitone signals with this sort of phase relationship generally achieve a crest factor between 10 and 11 dB.

## Slide 45: For More Information
NI Developer Zone
  http://zone.ni.com
LabVIEW measurements manual
Application notes
User solutions
White papers
Free evaluation software
  ni.com/labview
Notes: We hope you learned something about the tools now available to help you build your test, measurement and automation systems.
If you would like to find out more, there are many sources of information. Be sure to check out the contents of your seminar CD because it contains a great deal of useful technical information, including white papers, user solutions and application notes.
Our evaluation software gives you a chance to take an in-depth look at our software before you buy.
For the very latest information, be sure to check our Web site ni.com. This is constantly updated with product data and other valuable information.
If you still have questions, feel free to talk to us after the seminar or call your local National Instruments sales office.