## Slide 1: Measuring Temperature Accurately
Kenneth Kelly
Senior Engineer
Wed Aug 16
10:15-11:30 a.m., 1:45-3:00 p.m.
Cypress (8B)
Notes: At National Instruments, our goal is to make computer-based measurement and automation technologies to solve customer applications. While we strive to improve these products, we also strive to educate our customers about fundamental technologies and to assist them in their applications. Our goal in this presentation is to educate our customers on the fundamentals of measuring temperature and how to improve the accuracy of these measurements.

## Slide 2: Agenda
Temperature sensors
Instrumentation
Demo
Notes: WS1A- <number>
Measuring Temperature Accurately
<header>
<date/time>
In this presentation, we cover the measurement of temperature by discussing the operating principal behind each sensor type, the instrumentation required to condition the sensor’s output, and tips and techniques that will help you in selecting sensors, signal conditioning products, and applying them effectively to obtain accurate temperature measurements. This presentation will be concluded with a demo that illustrates software compensation for improving thermocouple measurements

## Slide 3: Temperature Sensors
RTDs
Thermistors
IC sensors
Thermocouples
Notes: We will discuss various popular types of temperature measurement sensors used with electronic instrumentation.
They are Resistance Temperature Detectors (RTDs), thermistors, IC sensors and thermocouples.
We will also discuss the instrumentation involved when using these sensors.

## Slide 4: RTD Basics
Resistance Temperature Detector
Resistance temperature
Platinum 100 at 0 °C
Up to 800 °C
Notes: RTDs are some of the most accurate temperature sensors. They are also known for their excellent stability characteristics. They are usually used to measure temperature from 0 °C to 450 °C, although some can be used up to 800 °C.
RTDs are made of fine coils or films of metals. When heated, the resistance of the metal increases – when cooled, the resistance decreases. This behavior allows their use as a temperature measurement device by measuring their resistance.
While RTDs are made from gold, silver, copper, tungsten and nickel, platinum (Pt) is the most commonly used material. Platinum is a noble metal and has excellent measurement integrity and stability, while having the highest resistivity (resistance/cross-sectional area) among the above metals.
Each of these metals has a particular resistance vs. temperature curve, and a different formula for converting the measured resistance to temperature.

## Slide 5: RTD Advantages
Very accurate
Very stable
Standardized among vendors
Large variety of packaging options
Notes: RTDs are used in applications requiring very high accuracy. RTDs are also known for their stability over time.
One of biggest advantages of using RTDs is that many vendors make RTDs which adhere to certain national and international standards. This allows users to pick and choose from many different vendors based on form factor and accuracy, without worrying about “temperature-to-resistance” characteristics.

## Slide 6: RTD Disadvantages
Costly
Require current source
Low resistance/small change
4-wire measurement
Self-heating
Slow
Notes: RTDs do have certain drawbacks and these have to be kept in mind while setting up the measurement system. Since platinum is expensive, RTDs made with Pt are expensive. Since the resistance of the RTDs is low (100 for Pt at 0 °C), and also because the resistance sensitivity is low (0.4 /°C ), parasitic lead wiring is a major issue. 4-wire measurement schemes are highly recommended as a means to circumvent this problem.
Most measurement instruments provide stable calibrated current excitation to the RTDs. This results in a power I2R dissipating inside the RTD, which heats the RTD, which increases its temperature. Note that when the measurement is made, there is no way to tell what component of temperature reading has come from the object being measured versus self-heating component. Every RTD has an associated dissipation constant (°C/W), which can be used to calculate this temperature rise due to self-heating. Always use the minimum excitation to reduce the effects of self-heating.
Most RTDs are slow to react to changing temperature because of thermal mass.

## Slide 7: Thermistor Basics
Thermally sensitive resistor
Resistance varies withtemperature
Semiconductor made from metal oxides
2,252 to 10,000 at 25 °C
Up to 300 °C
Notes: A thermistor is a piece of semiconductor made from metal oxides, pressed into a small bead, disk, wafer, or other shape, sintered at high temperatures and finally coated with epoxy or glass.
There are two kinds of thermistors – those with negative temperature coefficient (NTC), whose resistance decreases with temperatures, and positive temperature coefficient (PTC), whose resistance increases with temperature. NTC thermistors are more commonly used for temperature measurement applications.
Popular values are 2,252  , 5,000  and 10,000 . The 5,000  has about 200  /°C sensitivity at room temperature. Compared to 0.4  /°C for a Pt100, this is quite high.
Upper range of measurement can be up to 300 °C.

## Slide 8: Thermistor Advantages
Very accurate
Stable
High resistance
Good sensitivity
Low thermal mass
Notes: Very high accuracy thermistors with good stability are available (e.g. 0.1 °C).
Because thermistors have high resistance, lead wire resistance does not affect the accuracy of the measurements as much as in the case of RTDs. Two-wire measurement configuration are usually adequate for measuring thermistors.
The physical size of thermistors is quite small. They have low thermal mass and hence have very fast response to temperature changes. These features, combined with high resistance change with respect to temperature, allow sensitive measurements to be made with thermistors.

## Slide 9: Thermistor Disadvantages
Relatively recent standardization among vendors
Limited packaging options (epoxy bead)
Require current source
Self-heating
Notes: Compared to RTDs, interchangeable thermistors are available only in the recent years.
Because the resistance is quite high, the excitation current has to be small to prevent self-heating from affecting the accuracy of the measurement.
The limited range of the thermistor is due to the way thermistors are manufactured.
The semiconductor is soldered to its leads and if the temperature surrounding the thermistor goes beyond the melting point of this solder, then the thermistor will open circuit. Vendors have recently been switching to higher temperature solders to gain higher upper measurement ranges.

## Slide 10: IC Sensor Basics
Semiconductor devices
Voltage temperature
Analog (voltage or current) or digital outputs available
Up to 150 °C
Notes: IC sensors operate on the physical characteristics of p-n (transistor) junctions in semiconductors. They have integrated linearization circuits on the chip which gives a fairly linear output – proportional to temperature. Outputs are available in both analog and digital formats.
Because these devices are silicon-based ICs, their temperature range is limited to about 150 °C.

## Slide 11: IC Sensor (cont.)
Accurate near room temperature
Easy to interface
On chip linearization
Limited range
Needs power supply
Self-heating
Notes: IC temperature sensors are available with fairly high accuracy near room temperature (approx. 0.5 °C).
One of the main advantages of IC sensors is that they are easy to interface to.
The drawbacks are limited temperature range, the need for a power supply, and self-heating (unless they are mounted in a way that has low thermal resistance from the IC package).

## Slide 12: Thermocouple Basics
Junction of two dissimilar metals
Voltage rises with temperature
Non-linear
Notes: Thermocouples are, by far, the most commonly used temperature sensors.
A thermocouple is created when two dissimilar metals touch and the contact point produces a small open-circuit voltage as a function of temperature. This thermo-electric voltage is known as Seebeck voltage, named after Thomas Seebeck, who discovered the phenomenon in 1821. The voltage is non-linear with respect to temperature, however for small changes,
V = S * T,
where S is the Seebeck coefficient
S changes with temperature, causing the output voltages to be nonlinear over the operating range.
Several types of thermocouples are available. They are designated by capital letters that indicate their composition according to American National Standards Institute (ANSI) conventions. For example, J-type thermocouple is made from iron as one conductor and constantan (copper-nickel alloy) as another.

## Slide 13: Thermocouple (cont.)
Relative temperature measurement
Cold-junction compensation
  AB is measuring junction
  AC and BC are parasitic junctions
Notes: There is one fundamental difference between thermocouples and the earlier temperature sensors we discussed – with thermocouples, we can perform relative measurements.
Consider the thermocouple made of material A and B that is connected to the electronic instrumentation through copper traces (material C). There are three dissimilar metal junctions – AB, AC and BC. Junction AB generates a Seebeck voltage proportional to the temperature it is measuring. Junctions AC and BC have their own Seebeck voltages as well. The electronic instrumentation will measure a net voltage generated by ALL the junctions. To determine the contribution from junction AB, you need to know the temperature of AC and BC. Then, you can subtract the contributions of the parasitic junctions. This is called cold-junction compensation.

## Slide 14: Thermocouple (cont.)
Ice bath temperature reference
Two thermocouples (back-to-back)
Notes: Thermocouples require some form of temperature reference to compensate for unwanted parasitic thermocouples.
Traditionally, junctions AC and BC were held at 0 °C in an ice bath. The National Institute of Standards and Technology (NIST) created thermocouple reference tables using this setup.

## Slide 15: Thermocouple (cont.)
equivalent due to
Thermocouple Law of Intermediate Metals
Notes: Although an ice bath reference is accurate, a more practical approach is to measure the temperature of the reference junction with a direct-reading temperature sensor, such as a thermistor or an IC sensor, then subtract the parasitic thermocouple thermoelectric contributions. This process is called cold-junction compensation.
The thermocouple law of intermediate metals allows us to make the transition shown above.
Although you can use hardware compensation, it is now convenient to do software compensation.

## Slide 16: Thermocouple (cont.)
Popular thermocouple types: J, K, T
Other types: E, R, S, B, C, G
Selection based on:
  temperature range
  environment
  accuracy
  sensitivity
  cost
  form factor
Notes: There are a wide variety of thermocouple types available.
Each thermocouple type has it own set of coefficients for performing linearization. Be aware that the coefficients are typically only valid over a certain range. Look up tables are also available. The accuracy of thermocouples varies from type to type. The accuracy also varies from thermocouple to thermocouple of the same type. Slight deviations in the alloys used and tip construction can have an affect. Purchasing thermocouples made from the same batch of wire is one way to assure consistency in performance among all the thermocouples used in your system.
Thermocouples are readily available in a wide variety of form factors, such as weld tip, probe, film, and so on. Thermocouples also vary in cost depending on the form factor desired.

## Slide 17: Thermocouples Advantages
Self powered
Inexpensive
Rugged
Wide temperature range
Available in different form factors
Notes: Thermocouples are the most popular sensor among all temperature sensors for several reasons. Unlike other temperature sensors, they are self-powered – they do not need a power supply or current source. Since thermocouples are basically nothing more than two wires, they are inexpensive and rugged by nature. A variety of thermocouple types are available for wide temperature ranges.

## Slide 18: Thermocouple Disadvantages
Low voltage/small change
Need another temperature sensor
Not stable
Accuracy varies from one to another
Notes: Thermocouples do have certain drawbacks – they generate extremely low voltages, and the temperature sensitivity is also very small – needing very accurate and careful instrumentation. Since they are a relative measurement, another temperature sensor is required for cold-junction compensation. Also, thermocouples are not as stable as the other available temperature sensors.

## Slide 19: RTD Instrumentation
Notes: Typically, a precision current source is used for exciting an RTD. Since the resistance of the RTD is quite small, 4-wire connections are recommended to avoid errors due to lead wire resistance.
The current source value has to be small enough to avoid any self-heating errors.
Use high amplification or high resolution analog input instrumentation.

## Slide 20: Thermistor Instrumentation
Notes: Typically, a precision current source is used for exciting a thermistor. Alternately, a precision voltage source and a precision resistor in series is used in a voltage-divider configuration.
Since the resistance of the thermistor is quite high (e.g. 5000 ), 4-wire connections are recommended, but not necessary, to avoid errors due to lead wire resistance.
The current source value has to be small enough to avoid any self-heating errors.
Use high amplification or high resolution analog input instrumentation.

## Slide 21: IC Sensor Instrumentation
Power supply
Analog – filter and gain and ADC
Digital – DIO or micro-controller
Self-heating
Notes: All IC sensors need a power supply – it could be something ubiquitous, such as a 5V supply.
An IC sensor such as LM35 has a 10 mV/°C output. Use an appropriate gain on the input signal conditioning to get the best resolution.
Some other ICs also have digital outputs – those can be interfaced using digital I/O or directly with microcontrollers.
IC sensors are susceptible to self-heating. Appropriate care must be taken while mounting the sensor on the specimen to dissipate any heating which may cause the IC’s temperature to be higher than the specimen.

## Slide 22: Thermocouple Instrumentation
Amplification
High resolution
Filtering
Notes: Because thermocouple voltages are very low (10s of mV ~50 V/°C), use as much gain as possible for best resolution and noise performance. Make sure you select the gain so that the voltage does not exceed the maximum limit, when the thermocouple is at elevated temperatures.
In addition to or in lieu of using very high gain, you may use a very high resolution analog-to-digital converter (ADC) to get the best performance.
Since temperature is a slowly moving phenomenon, in most cases, use as low frequency filter as possible, to reject noise, specifically (50 Hz/ 60 Hz) power-line noise. Filtering also prevents aliasing – a common problem in sampled measurement systems. Use of sigma-delta ADC will also offer excellent power-line noise rejection with the built-in digital filters.
You can also take a large number of measurements over a power-line cycle, and then take the average of these readings – this will in effect, cancel out the effects of power-line noise.

## Slide 23: Thermocouple Instrumentation
Ground-referencing
Open-thermocouple detection
Cold-junction compensation
Isothermal connections
Notes: If the thermocouple is not ground-referenced, make sure your instrumentation provides some way to ground-reference the thermocouple. Otherwise, it may float away outside the instrumentation’s common-mode input limits.
To detect open or broken thermocouple, many instrumentation systems inject a small amount of current into the thermocouple. This current normally does not cause any effect on the measurement. But when the thermocouple is broken, this current charges the small parasitic capacitor, which in turn causes the input of the instrumentation to register a significantly higher voltage than the normal thermocouple voltage.
As discussed earlier, the thermocouple method is a relative measurement method, and cold-junction compensation is required.
A large percentage of the error in measurement is due to cold-junction compensation. It becomes even worse if the thermocouple connection to the instrumentation is not isothermal.

## Slide 24: General Instrumentation
Scanning
Switching (solid-state vs mechanical relay)
Differential measurements
Isolation
Notes: Even though temperature is a slowly moving phenomenon, if there are a large number of channels it is very useful to have high speed scanning. Correlation between different channels is also possible with high speed scanning.
For multiplexing a large number of channels, some instruments use solid-state multiplexers, while some others use mechanical relay-based multiplexers. While mechanical relays may seem cheaper to buy, they have a finite lifetime. Therefore, make sure your instrumentation allows easy replacement – without having to rewire your sensors – whenever you need to replace the multiplexer module. Otherwise, choose instrumentation with solid-state technology, which does not have this problem.
Always use differential measurement methods for low level sensors. Ground loops can make the errors very significant.
Isolation has several advantages – including breaking ground loops, rejecting large common mode voltages, protection from damaging high voltages, and most important – offering safety under possible accidental conditions.

## Slide 25: Which Sensor Should I Use?
RTD
    Very accurate
    Most measurement issues to contend with
Thermistor
    Very accurate over limited range
    Simple measurement setup
IC Sensor
    Easy interface
    Very limited range and accuracy
Thermocouple
    Lots of options
    High quality instrumentation readily available
Notes: WS1A- <number>
Measuring Temperature Accurately
<header>
<date/time>
This slide provides a brief overview of each sensor that has been discussed. As you can see, there are a number of issues involved with making a temperature sensor selection. Realize that accuracy is itself a relative term, in that what is considered accurate to one system may not be accurate enough for another.
Many systems only require about 1 °C of accuracy, in which case most any sensor will suffice and the decision needs to be made based on other merits. Any accuracy requirement of less than 1 °C will require some effort for every 0.1 °C desired. An accuracy of 0.7 °C is relatively easy to achieve, however 0.5 °C is a little more difficult, and less than 0.3 °C can require special efforts to assure measurement quality.

## Slide 26: Instrumentation
High precision
  NI-4350, NI-4351
  DMM based systems
High channel count
  SCXI
  Switch based systems
Distributed
  FieldPoint™

## Slide 27: High Precision
NI-4350, NI-4351
  24 bit ADC resolution
DMM switch system
  5.5 digit DMM with relay multiplexer

## Slide 28: High Channel Count
SCXI system
  Very high channel counts possible
  3,000 channels per system
  High speed scanning
  Conditioning for a variety of sensor types

## Slide 29: Distributed
FieldPoint
  Serial communication
  Ethernet based communication
  Bank isolated
  Low speed monitoring

## Slide 30: Conclusions
Sensor selection based on application
Signal conditioning & techniques for application
More Information
  ni.com/sensors
Questions?
Notes: WS1A- <number>
Measuring Temperature Accurately
<header>
<date/time>
We have given you information that should prove helpful in selecting the best sensor for you application, as well as information regarding the best choice of instrumentation and useful techniques to obtain accurate measurements.
Visit our Web site ni.com to obtain more information on our products as well application white papers on subjects such as the measurement of temperature using thermocouples, RTDs, and thermistors.