## Slide 1: Making Frequency Domain Measurements
Trent Butcher
Computer-Based Instruments Product Manager
Wed Aug 16
10:15-11:30 a.m., 1:45-3:00 p.m.
Cedar (8C)
Notes: In this seminar you will learn about how frequency domain measurements are made. We will discuss signal conditioning on time domain signals in preparation for frequency analysis, the Fast Fourier Transform (FFT), zoom FFT analysis, octave analysis, and swept sine analysis.

## Slide 2: Agenda
Basics of frequency measurements
Time domain considerations
Frequency analysis
National Instruments solutions
Notes: In this seminar we will discuss the various benefits of computer based instruments as they apply to making frequency based measurements. The basics of making frequency domain measurements will be discussed as well as a detailed discussion about time domain considerations and frequency analysis. Finally, we will discuss solutions provided by National Instruments.

## Slide 3: Time vs Frequency Domain
Notes: The basic concept of Frequency Domain Analysis is to analyze the components of a time domain signal (one that we could see on an oscilloscope).
Some measurements, such as harmonic distortion, are very difficult to quantify by inspecting the time domain signal on an oscilloscope. However, when you use a Dynamic Signal Analyzer to display a signal in the frequency domain, you can measure harmonic frequencies and amplitudes.
Fourier’s theorem states that any waveform in the time domain can be represented by the weighted sum of sines and cosines. The same waveform appears in the frequency domain as a pair of amplitude and phase values at each component frequency.
The Fourier transform converts a time domain representation of a signal into a frequency domain representation. The Fourier transform has its digital equivalent called the Discrete Fourier Transform (DFT) which can transform sampled time domain data into the frequency domain. The Fast Fourier Transform is an optimized implementation of a DFT.

## Slide 4
Sample Time
Domain Signal
FFT
Anti-Alias
Filter
Octave
Acquire
Waveform
Basics of Frequency Measurements
Signal
Conditioning
Frequency
Analysis
Notes: When making frequency domain measurements with a Dynamic Signal Analyzer, it is first necessary to acquire the waveform in the time domain. This is done via sampling and A/D converters.
For accurate frequency-domain measurements, your input signals will require signal conditioning. Some signal conditioning is applied to a signal prior to sampling. Anti-aliasing filters prevent high frequencies from appearing as lower frequency spurious signals in the acquired data. If the signal is out of the range of the analyzer or small within the range, overload detection or programmable gains may be applied. Next, an A/D converter samples the signal and they are transferred to host memory. When phase information between multiple signals is important, triggering and simultaneous sampling should be used.
After the waveform has been acquired, you can apply a Fourier Transform to a set of samples to convert it to a frequency-domain representation. Other common frequency domain alternatives include Chirp-Z and Fractional-Octave Analysis.
If the FFT and/or octave analysis is performed as fast or faster than the time domain acquisition, then the analyzer is considered to be “real-time.” With the performance improvement of PCs, some real-time analysis applications can be achieved using a time domain acquisition device and the processor on the PC.

## Slide 5: Time Domain Considerations Sampling Guidelines
Improperly sampled
Properly sampled
fN = fs/2 fs: sampling frequency
Nyquist Theorem
  sampling rate > 2 x maximum frequency of signal
Nyquist Frequency (fN)
  maximum frequency that can be analyzed
Frequencies above Nyquist Frequency cause aliasing
Notes: The Nyquist Theorem states that a signal can be reconstructed from its samples if the sampling rate is greater than twice the frequency of the signal. It follows then that the highest frequency (Nyquist frequency: fN) that can be analyzed is fN = fs/2, where fs is the sampling frequency. Any analog frequency greater than fN will, after sampling, appear as a frequency between 0 and fN. Such a frequency is known as an “alias” frequency. In the digital (sampled) domain, there is no way to distinguish these alias frequencies from the frequencies that actually lie between 0 and fN. Therefore, these alias frequencies need to be removed from the analog signal before sampling by the A/D converter.

## Slide 6: Time Domain Considerations Anti-Aliasing Filter
Removes frequencies higher than Nyquist frequency
Analog low-pass filter
Before sampling
Flat Frequency
Response
Sharp
Roll-off
Notes: An analog low-pass filter typically removes the components of a signal that exceed the Nyquist frequency. This anti-aliasing filter should exhibit a flat in-band frequency response with a good high frequency alias rejection and a fast roll-off in the transition band.

## Slide 7: Time Domain Considerations Alias Free Bandwidth
f1
f3
fs /2
fs
alias free bandwidth
f1
fs /2
fs
anti-aliasing
filter
f2
attenuated
f2
alias
f3
f4
RAW SIGNAL
ACQUIRED SIGNAL
Nyquist Frequency
Sample Frequency
Notes: The user needs to be aware of the relationship between the anti-aliasing filter cutoff frequency and sampling frequency. Some Dynamic Signal Analyzers (DSA) use Delta/Sigma converters to convert analog signals into digital data. These converters have built-in anti-aliasing filters whose cutoff frequency is typically 45 to 50% of the sample rate. This sort of converter provides very sharp rejection, and the filter cutoff frequency automatically tracks with the sample rate. Other DSAs use converters that do not have built-in antialias filters and must rely on external filtering with fixed filter cutoff frequencies.

## Slide 8: Time Domain Considerations Additional Items
 analog to digital converter
  High resolution
  Built-in anti-aliasing filters
  Suited for sound and vibration measurements
Simultaneous sampling and triggering
  Phase relationship between signals
Programmable gain
Overload detection
Notes: As mentioned previously, Delta/Sigma converters have built-in anti-aliasing filters that prevent frequencies higher than the Nyquist frequency from appearing as spurious noise. Delta/Sigma converters also have very high time domain resolutions. These characteristics make Delta/Sigma converters the best selection for making frequency measurements for sound and vibration applications.
Typically when making measuring multiple frequency measurements, the phase relationship between the signals is desired. In order to obtain this information, the inputs must be simultaneously sampled. Triggering is also needed to begin the acquisition of the signals at the desired times. The triggering modes can be selected in software on the PC.
To adjust to the limits of the input range, programmable gains are applied to signals smaller than the the input range. Using development software on the computer these gains can be adjusted programmatically. For signals larger than the input range, overload detection circuitry is used to prevent damage to the instrument.

## Slide 9: Time Domain Considerations Smoothing Windows
Nonintegral number of cycles
Reduces spectral leakage
Window selection depends on the application
PC Based instruments greatly facilitate transient analysis
No windowing
Windowing
Window
Notes: Because the amount of leakage is dependent on the amplitude of the discontinuity at the boundaries, you can use windowing to reduce the size of the discontinuity and hence reduce spectral leakage. Windowing consists of multiplying the time domain signal by another time domain waveform, known as a window, whose amplitude tapers gradually and smoothly toward zero at edges. The result is a windowed signal with no (or very small) discontinuities, and thus reduced spectral leakage. The window should be applied to the signal after it has been digitized by the A/D converter.
There are many different types of windows. The one you choose depends on your application. The typical trade-off between windows is between amplitude accuracy and frequency accuracy. For example, the flattop window provides good amplitude accuracy, but spreads the signal across the spectrum effectively reducing frequency resolution. Another type of window, the Hanning window, provides good frequency accuracy, but can produce up to a 16% amplitude error.
A different type of analysis that is only possible on a FFT analyzer is transient analysis. An exponential window can be used for this type of analysis. A decaying exponential is applied to the signal which reduces it to zero at the end of the time period, thus eliminating any spectral leakage. PC based instruments are key to making these measurements because they provide the necessary triggering to do this type of analysis.

## Slide 10: Frequency Domain Analysis
FFT analysis
Octave analysis
Swept sine analysis
Notes: There are three types of frequency domain analysis that we will discuss in the remaining section of this presentation.
FFT analysis
Octave analysis
Swept sine analysis

## Slide 11: FFT Analysis
Time domain in discrete values
  Use Discrete Fourier Transform (DFT)
Fast Fourier Transform (FFT)
  Optimized version of DFT
Highest frequency that can be analyzed
Frequency resolution
fs : sampling frequency
T : total acquisition time
N : FFT block size
Notes: The time domain signal is sampled and digitized into discrete values. Because of this the Fourier transform will not work on this data. Instead, the Discrete Fourier Transform (DFT) is used which produces, as its result, the frequency domain components in discrete values or “bins.”
The Fast Fourier Transform (FFT) is an optimized version of the DFT that takes less computation to perform. Below is the equation for the FFT.
Each frequency component produced through the FFT is the result of a dot product of the time signal with the complex exponential at the component frequency.
The DC component is the dot product of x(n) with [cos(0)-jsin(0)], or with 1.0.
The use of the FFT for frequency analysis implies two important relationships.
The first one relates the highest frequency that can be analyzed to the sampling frequency (see the previous discussion of the Nyquist theorem).
The second one links the frequency resolution to the total acquisition time, which is related to the sampling frequency and the block size of the FFT.

## Slide 12: FFT Analysis
FFT gives magnitude and phase information
  Magnitude = sqrt(Real^2 + Imag^2)
  Phase = Tan-1(Imag / Real)
Power Spectrum reflects the energy content
  Power Spectrum = Mag^2
Applications
  Vibration analysis
  Structural dynamics testing
  Preventative maintenance
  Shock testing
Notes: The FFT spectrum output is complex; that is, every frequency component has a magnitude and phase. The phase is relative to the start of the time record, or relative to a cosine wave starting at the beginning of the time record. Single-channel phase measurements are stable only if the input signal is triggered. Dual-channel phase measurements compute phase differences between channels so that if the channels are simultaneously sampled, triggering is usually not necessary.
The energy content of the signal over its frequencies is given by the Power Spectrum. The Power Spectrum is simply the magnitude of the total spectrum. The magnitude is the square root of the FFT times its complex conjugate, or the square root of the sum of the squares of the real and imaginary parts of the FFT. The phase is the arctangent of the ratio of the imaginary and real parts of the FFT, and is usually between  and -  (between 180 and -180 degrees).
The Power Spectrum is commonly used in many applications where phase information is not needed. This is similar to the information provided by swept spectrum analyzers which do not provide phase information.
Applications that use FFT analysis are quite extensive. Those applications requiring fast response times, such as transient analysis, vibration and shock testing use FFT analysis. Preventative maintenance and structural dynamics testing also use FFT analysis primarily in their applications.

## Slide 13: Zoom FFT Analysis
Concentrates (“zooms”) FFT on a narrow band of frequencies
Improves frequency resolution
Distinguishes between closely-spaced frequencies
Baseband analysis requires longer acquisition time for better resolution – requires more computation
Notes: Zoom FFT analysis is used to concentrate, or “zoom,” the FFT analysis on a narrow band of frequencies. This improves the frequency resolution and helps you to distinguish closely-spaced frequencies.
With standard FFT analysis, the frequency resolution is dependent on the total acquisition time. The frequency span always begins at zero and is spread to a specified frequency. This limits high resolution analysis at higher frequencies. One way to improve the resolution is to increase the acquisition time. However, this places a greater computational burden on the processor. The better solution is to start the frequency span at a value other than zero. This is what is called zoom FFT analysis. Zoom analysis is performed by using a digital mixer and simulated digital sine wave to effectively move the frequency range of the input signal down to a lower frequency. The “start” frequency is mixed down to DC and the “stop” frequency to what was the baseband end frequency. The FFT is then performed on this signal and consequently over the desired frequency range.

## Slide 14: Zoom FFT Analysis
Baseband FFT Analysis
Zoom FFT
Analysis
Notes: This example shows how zoom FFT analysis is helpful in detecting two closely spaced sinusoids. The top graph shows the signals as being viewed using standard FFT analysis. The frequency spectrum begins at zero and increases to some specified frequency.
With zoom FFT analysis, the range around the sinusoids can be “zoomed” in to view their respective frequencies. As seen on the chart, the two different frequencies are clearly revealed.

## Slide 15: Zoom FFT Analysis – How It’s Done
f1
f2
f1 + f2
f1 - f2
f1 - f2
f1
f2
f1 – f2
Notes: Zoom FFT analysis (also called band selectable analysis) is done by shifting the desired signal down in frequency to a range where a baseband FFT can be performed on the signal.
First the signal is passed through a mixer. A mixer is a device that combines signals of two frequencies at its two inputs and gives the sum and difference of the two signals at the output. At one input the signal to be analyzed, which is already digitized at this point, is connected. At the other input, a pure digital sinusoid is connected. In this case only the difference between the signals is desired so a digital lowpass filter is chosen that will select the proper signal. Now our desired frequency has shifted down in frequency to where the start frequency is now at zero and the FFT can be performed only over the desired frequency range, thus concentrating the resolution completely over the selected frequency span. This is the manner in which you “zoom” in on a desired frequency range.

## Slide 16: Octave Analysis
Analysis performed through a parallel bank of bandpass filters
  One octave corresponds to the doubling of the frequency
  Reference frequency is 1 kHz (audio domain)
220 Hz
440 Hz
880 Hz
A
A
A
Notes: In many cases, when dealing with sound measurement and analysis, the final “customer” is the human ear. And like most human senses, the ear exhibits a response based on a logarithmic scale for both the level and the frequency. So, to produce results that are somewhere related to this human perception, sound levels are expressed in decibels and frequency content measured with a logarithmic scale. A lot of research efforts are currently focused on this field of psycho-acoustics, but octave band analysis remains a technique of prime choice.
One octave band corresponds to the frequency range between two frequencies with a ratio of 2 to 1. A typical example of this is a piano keyboard, where two consecutive C tones are exactly separated by an octave. The reference center frequency for the audio range is at 1 kHz. Other midband frequencies are thus at 500 Hz, 250 Hz, … and 2 kHz, 4 kHz, … The octave band centered at any particular midband frequency fm extends from to (fm is the geometric mean).
Octave band analysis is performed through a parallel bank of bandpass filters. The real-time output of each filter is then averaged to compute the power in each band and displayed as a bar graph.

## Slide 17: Octave Analysis
Octave analysis gives log-spaced frequency information
  Similar to human perception of sound
  1/1, 1/3, 1/12, and 1/24 octave analysis
FFT gives linearly-spaced frequency information
Applications
  noise emissions testing
  acoustic intensity measurement
  sound power measurement
  audio equalization
Notes: FFT analysis measures the spectrum with linearly-spaced frequency bins. Octave band analysis measures the spectral energy with logarithmically spaced bandpass filters.
As already mentioned, this is similar to the way people perceive sound. Another advantage, especially when dealing with signals that exhibit a lot of different frequency components, is that the energy content is grouped by bands. Different bands could then be related to totally different phenomena. This also facilitates the comparison between signals and is a major reason why octave band analysis is used in numerous standards.
Fractional octave bands allow to select a frequency resolution well adapted to the signal of interest. Typical fractional bands are 1/3 octave with three bands per octave, 1/12 octave with 12 bands per octave, and 1/24 of octave with 24 bands per octave. Specifications regarding these octave band and fractional octave band filters are given by ANSI and IEC standards.
Applications are mainly in the fields of acoustics and audio signal processing. A specific example is in testing noise emissions in aircraft. Real-time 1/3 octave analyzers are required by federal aviation regulations to measure noise at different frequency bands. Audio equipment designers use octave analyzers to measure the sound power at various frequency bands to determine the quality of the equipment. Acoustic engineers determine the design of rooms or concert halls by measuring sound power levels over various frequency ranges.

## Slide 18: Swept Sine Analysis
Source steps through a range of frequencies
Analyzer measures frequency amplitude and phase at each step
Non-FFT based
Source
Device
Under Test
Frequency
Response
Notes: Swept sine analysis is another method of performing frequency analysis. The output of the instrument is configured to sweep a single sine wave over a range of frequencies. This is passed to the device under test (DUT). The response of the DUT is then measured at one of the analog inputs. The amplitude and phase of the response is measured at each individual frequency step. The steps can be linearly or logarithmically spaced. The amplitude spectrum of channel A is calculated, then the amplitude spectrum of B, then the frequency response B(f)/A(f). This is not done via the FFT. All measurements are taken in the time domain.

## Slide 19: Swept Sine Analysis
Auto-ranging: dynamic range optimized at each frequency
  Adjust source amplitude
  Adjust input range
  Both improve dynamic range at particular frequencies
    Can get 140 dB effective dynamic range
Channel B
Channel A
Notes: Comparing the swept sine measurement technique to the FFT measurement, the swept sine measures only one line at a time and is therefore much slower to the FFT which measures a group of frequencies simultaneously.
On the other hand, swept sine analyzers can optimize their measurements at each frequency by adjusting the source output power and input voltage range to increase the S/N ratio. Adapting to the system under measurement in this fashion can extend the effective dynamic range of the DSA to beyond 140 dB.

## Slide 20: Swept Sine Analysis
Auto-resolution
  Sweep optimized - more time at lower frequencies, less time at higher
  Increases frequency resolution on rapidly changing responses
Applications
  Speaker testing
  Cell phone testing
  Electronic equipment characterization
Notes: To optimize total measurement time when sweeping a large frequency range, the bandwidth at each frequency line can be configured to be inversely proportional to the frequency. This has the effect that more time is used for lower frequency measurements and less time for higher frequency measurements.
You can also optimize sweep time for narrow response functions by configuring the swept sine analyzer to skip frequencies where the response remains fairly constant, and to increase frequency resolution where the response is changing rapidly.
Applications that require a frequency response, but cannot use a noise source (one that outputs power across all frequency values), use swept sine analysis. These types of applications are speaker and cell phone testing. In these applications it is important to analyze the system at each frequency along the spectrum individually.

## Slide 21: Comparison of Frequency Analysis Methods
FFT analysis
  Very fast
  Linear frequency scale
  Based on discrete Fourier transform
Octave analysis
  Logarithmic frequency scale
  Set of filters dividing frequency into bands
  Similar to how human ear perceives sound
Swept sine analysis
  Good dynamic range
  Source and analyzer step across frequency range
  Slower response
Notes: FFT Analysis
  Applications requiring very fast response times are best suited for FFT analysis. With the increase in processor speeds, this method of frequency analysis has become a very fast method for obtaining frequency spectrums. This type of analysis can also be used to measure trends in frequency responses over time. This is useful in applications involving vibration and rotating machinery. The responses are typically viewed in a waterfall display.
  FFT analysis is also suited for transient analysis. In the example of obtaining the frequency response of a hammer striking a beam, this method is particularly useful. Typically the response to an impulse dies out quickly. FFT analysis can detect all the frequencies that occur during that short time because it is very fast. Since the impulse, as viewed in the frequency domain, contains energy over a large range of frequencies, FFT analysis is needed to detect these frequencies. Another need for FFT analysis with this test is triggering. FFT analysis needs to be triggered with the impulse input to the system.
Octave Analysis
  Since octave analysis closely resembles the way in which the human ear perceives sound, it is useful for applications involving audio testing, speech analysis, acoustics testing and environmental noise measurements. Like the human ear, an octave analyzer takes the frequency response on a logarithmic frequency scale over certain defined bands. Each octave may be divided into various bands such as 1/3, 1/12 or 1/24 octave. The time domain signal is passed through various digital filters as defined by these bands, and power level measurements are taken within each band.
Swept Sine Analysis
  Applications such as speaker and cell phone testing require testing at each individual frequency. To obtain a frequency response of a system, the source must step through each frequency instead of providing power across the entire frequency spectrum as with a noise source. With swept sine analysis, a single sine wave is applied to the input of the system and is stepped through a range of frequencies. The output of the system is measured for amplitude and phase at each frequency step. With this mode, a much higher dynamic range can also be achieved. This method of frequency analysis is slower than FFT analysis, but it is suited for these applications.

## Slide 22: National Instruments Solutions
NI 4551 & NI 4552 instruments
  On-board processing
  95 kHz alias free bandwidth
  90 dB dynamic range
  Simultaneous sampling
  FFT and swept sine modes
  Octave analysis software add-on
NI 44xx boards
  Host based in-line processing
  Same signal conditioning for frequency analysis as NI 45xx
  Flexibility of analysis tools
Powerful programming environments
  LabVIEW™, Measurement Studio™, Visual Basic, C/C++
Notes: The NI 4551 and NI 4552 instruments are high performance computer-based dynamic signal analyzers for the PCI bus. They are designed to make sound and vibration measurements with the same performance as stand-alone dynamic signal analyzers. These instruments perform real-time FFT, octave analysis, and swept sine analysis, implement anti-aliasing filters, use simultaneous sampling, have a 90 dB dynamic range, use 16 bit delta/sigma ADCs, and have 32 digital I/O lines. Because these instruments are connected to the PC via the high speed PCI bus, users can retrieve data faster than with stand-alone instruments and expect faster measurement times.
The PCI-445x boards have similar characteristics as the NI-455x instruments. They come with simultaneous sampling, anti-aliasing filters and the other analog input characteristics, but they do not include on board processing. All processing of the time domain data for frequency analysis and octave analysis, must be performed on the computer in software.
The NI-455x instruments can be programmed using the NI-DSA instrument driver in LabVIEW, LabWindows™/CVI, Visual Basic and many C/C++ programming environments. These instruments can also be programmed using the NI-DAQ driver, but without on board processing. The PCI-445x boards are programmed using the NI-DAQ™ driver. VirtualBench™ DSA is a ready to run application which takes advantage of the NI-DSA instrument driver or NI-DAQ depending on which instrument you are using.

## Slide 23: Sound and Vibration Toolset
Add-on software library for LabVIEW
VIs that give measurements
  Sound-level measurements
  Fractional octave analysis
  Transient analysis
  Includes source code
For use with the NI 44xx boards
Notes: The Sound and Vibration Toolset is an add-on for the LabVIEW programming environment. This toolset offers functions and visualization tools for fractional-octave analysis, sound level measurements, frequency analysis, transient analysis and calibration. It is for use with instruments or boards that lack on-board processing. For example, this toolset could be used in conjunction with the PCI-445x boards for frequency analysis.

## Slide 24: Conclusion
Several time domain considerations for frequency measurements
Frequency analysis
  FFT analysis
  Zoom FFT analysis
  Octave analysis
  Swept sine analysis
National Instruments has the computer based instrument solution!
Notes: In this presentation you have seen the basic considerations to make when performing frequency domain measurements. Several time domain considerations such as sampling rate, anti-aliasing filters, windowing, resolution and gains are necessary to perform good frequency analysis. You have also seen various types of frequency analysis such as zoom FFT analysis and octave analysis as well as what some of their applications are. Finally, we have discussed the various benefits of using computer based instruments to perform these measurements.