## Slide 1: Measuring Strain and Displacement Accurately
Mark Whittington
Senior Engineer
Wed Aug 16
12:00-1:15 p.m., 3:30-4:45 p.m.
Cypress (8B)
Notes: At National Instruments, our goal is making computer-based measurement and automation technologies solve your applications. While we strive to improve these products, we also strive to educate all of our customers about fundamental technologies and to assist them in their applications. Our goal in this presentation is to show you the fundamentals of measuring strain and displacement.

## Slide 2: Agenda
Strain and LVDT fundamentals
Transducer options
Instrumentation requirements
Tips and techniques
Notes: In this presentation, we cover the measurement of strain and displacement by first discussing the fundamental operating principals, the instrumentation required to condition strain gage and LVDT signals, tips and techniques that will help you in selecting sensors, signal conditioning products, and applying them effectively to obtain accurate measurements.

## Slide 3: What is Strain?
L+DL
W-DW
force
force
principal strain : e =
DL
L
transverse strain : eR =
DW
W
Poisson’s ratio : n = (~ .285 for steel)
eT
e
Notes: Strain is the amount of deformation of a body due to an applied force. More specifically, strain () is defined as the fractional change in length.
It is important to remember that strain occurs both parallel and perpendicular to the direction of the applied force. In the figure, we see that applying a tensile force to the bar causes the bar’s length to increase by L and the bar’s girth to decrease by W. The principle strain ,  = L/L, is the strain in the direction of the applied force, and the transverse strain, T = W/W, is the strain perpendicular to the applied force. The ratio of the two strains,  = T / , is called Poisson's ratio, and depends on the material. For example, Poisson’s ratio is 0.285 for steel and 0.32 for aluminum.
Note that strain is a dimensionless quantity. It is usually expressed in microstrain (). One microstrain represents a deformation of one one millionth. The value, 10,000 microstrain, represents a deformation of 1 percent.

## Slide 4: What is a Strain Gage?
Backing
Principal Axis
Solder Tabs
Notes: The metallic strain gage consists of very fine wire or, more commonly, metallic foil arranged in a grid pattern. The grid pattern maximizes the amount of foil subject to strain along the principal axis of the gage. The cross sectional area of the grid is minimized to reduce the effect of shear strain and Poisson Strain. The grid is bonded to a thin backing, called the carrier, which is attached directly to the test specimen. Therefore, the strain experienced by the test specimen is transferred directly to the strain gage, which responds with a linear change in electrical resistance.

## Slide 5: Types of Strain Gages
Foil – most widely used
Wire – very high temperature use
Semiconductor
  High gage factor ~100
  High temperature drift
  Fragile
Notes: By far, foil gages are the most widely used. Foil gages are manufactured using the same photo-etching techniques used to produce printed circuit boards. The metal foil is usually either constantan or a modified nickel-chromium, alloys with properties favorable for measuring strain such as high resistivity, high strain sensitivity, good fatigue life, and high elongation capability. Foil gages offer two significant advantages over wire types: 1) For the same cross sectional area, a foil gage has a much greater surface area through which strain is transmitted from the test specimen to the foil, hence the backing and adhesive undergo much less shear stress. 2) For the same cross sectional area, heat conductivity between the conductor and test specimen is much higher so the gage operates at much higher power levels.
Prior to the 1950’s, gages were predominantly of the wire type. Today, wire strain gages are mostly used in high temperature applications.
Semiconductor gages have strain sensitivities 10 to 50 times greater than metal foil gages, but because of fragility, low ductility, and high temperature sensitivity, they have not become widely used except for high sensitivity pressure transducers and accelerometers

## Slide 6: Strain Gage Sensitivity
GF = =
DR/R
DL/L
GF = 2 for most foil gages
3500 m e (.35 %) DR/R = .7%
Gage factor:
DR/R
e
Notes: The parameter describing the sensitivity of a strain gage to strain is gage factor and is defined as the ratio of fractional change in resistance to applied strain. Note that the parameter is dimensionless.
For most foil gages, GF is between 2 and 2.15. Some types used for dynamic work have gage factors as high as 3.7. Semiconductor gages, by contrast may have gage factors as high as 150.
Most strain measurements are below 3,500 microstrain. This represents a 0.7 percent change in resistance. Because the change is relatively small, strain measurements are usually performed using wheatstone bridges rather than direct resistance measurements. A small change in bridge element resistance produces a voltage deflection away from zero at the bridge output. Hence, the large offset of a direct measurement is eliminated.

## Slide 7: Quarter Bridge
- 4Vr
GF(1 + 2Vr)
e =
where
EO
EI
Vr = ( ) - ( )
EO
EI
strained
unstrained
0.5 microvolts per microstrain per V of excitation
Nonlinear
Temperature drift
RG
R1
R2
R3
EO
EI
+
-
-
+
Notes: Where the strain output of a single gage is measured (as is most common in experimental stress analysis), the quarter bridge circuit is used. The gage is placed in one arm of the circuit. The resistor in the lower arm is selected to have the same resistance as the gage (RG = R3). The two resistors in the opposite leg must equal one another (R2 = R1) but need not have the same resistance as the gage.
We may easily solve for strain given the measured input voltage, Ei (both strained and unstrained) the excitation voltage, Eo, and gage factor, GF. The formulation shown takes into account the non-zero input voltage in the unstrained case, resulting from imperfect matching of bridge components. Notice that the formula is not linear since there is the denominator term (1 + 2Vr). Hence, it is desirable to use this formula in software rather than merely assuming strain is equal to Ei multiplied by a proportionality constant. (The effect of the nonlinearly is small, however, and can be ignored with little error for levels of strain below 5,000 ).
The quarter bridge offers no inherent temperature drift cancellation. Several techniques to minimize temperature drift in quarter bridge set-ups are discussed later.

## Slide 8: Half Bridge
- 2Vr
GF
e =
Twice the sensitivity of quarter bridge
Temperature drifts cancel out
Linear
RG1
RG2
R
R
RG2
RG1
EO
EI
+
-
+
-
Notes: In the half bridge, two gages are placed in adjacent arms. If the physical location and force applied to the test specimen is such that one gage receives compressive strain while the other receives tensile strain at the same magnitude, the output of the half bridge is twice that of the quarter bridge. Similarly, the output of the bridge is zero if both gages change resistance by the same amount. As a result, the half bridge remains balanced when both gages undergo the same resistance change from thermal drift. Also notice that the half bridge is inherently linear.
The slide shows a set-up for measuring the bending moment through the beam. Note that axial strain (strain along the axis of the beam) would cause both strain gauges to change resistance by the same amount and in the same direction. Hence the bridge output remains at zero and unresponsive to the axial stress field which is not of interest to the experimenter in this case.

## Slide 9: Full Bridge
- 4Vr
GF
e =
Twice the sensitivity of half bridge
Temperature drifts cancel out
Linear
RG1
RG2
RG3
RG4
RG1
RG2
RG3
RG4
EO
EI
+
-
+
-
Notes: In the full bridge, active gages are placed in all bridge arms. If opposing gages are like-strained and adjacent gages oppositely strained, the output of the full bridge is twice that of the half bridge and four times that of the quarter bridge. Like the half bridge, the full bridge remains balanced when all gages undergo the same resistance change, hence the full bridge is also self-temperature compensating. Also, like the half bridge, the full bridge is inherently linear.
Most strain gauge-based sensors use full bridges.
The slide shows a set-up for measuring the torque. Note that the set-up is inherently unresponsive to axial stress and bending moment. Only torque, the variable of interest, is measured.

## Slide 10: Causes of Measurement Error
Temperature related
  Uncorrected thermal output
  Excessive gage self-heating
  Improper wiring
Mechanical related
  Poor bonding
  Misalignment
Notes: Most errors in strain gage measurements are related to temperature change or differences in temperature between different gauges in the same bridge. The quarter bridge is particularly sensitive to temperature variation since the bridge itself offers no inherent temperature compensation. Also, resistance changes in the wiring can have the same effect as temperature changes in the gauges and introduce errors if not properly canceled out.
Improper bonding of the gage to the test specimen can also introduce errors. Causes might be creep or fracturing in the glue line. Finally, errors can occur when there is a misalignment of the gages from the intended direction.

## Slide 11: Temperature Compensation
Dummy gage tracks active
EO
EI
active
dummy
+
-
+
-
NO
STRESS!
Notes: The simplest way you can correct for temperature drift is by using a half bridge or full bridge configuration where all strain gages are at the same temperature and mounted on the same material. In the circuit shown here, the dummy gage tracks the active gauge with temperature and thus causes no shift if the output voltage is related to temperature.
It is important to remember that if there is any force in the direction of the dummy gage, the measurement of strain along the direction of the active gage will be in error. Therefore, you should take care to mount the dummy gauge in a stress free region.

## Slide 12: Strain Gauge Rosettes
Single plane rosette
Completely unknown
stress fields
“Tee” rosettes
One stress field, known direction, use 1 half bridge
Bi-axial stress fields, known directions; use 2 quarter bridges
Notes: We could place a two element or “tee” rosette gauge in the half bridge configuration of the last page to obtain a temperature compensated measurement. We must remember that both gauges respond to a single stress field along the axis of one gauge because both an axial and transverse strain field are generated.
If the test object experiences sheer stress in addition to axial stress, at least two independent measurements are necessary at a single point to determine the stress fields at that point. If the directions of the stress fields are unknown, three independent measurements are necessary. In these cases, you should useeach gauge of the rosette in a quarter bridge configuration. The dummy technique of the last page might not be practical or convenient due to the difficulty of mounting the dummies in a stress free location.
We must find another way of temperature compensating the quarter bridge.

## Slide 13: Quarter Bridge Temp. Compensation
Use correct “S-T-C” number
Correction polynomial spread:
+ 0.15 me/oF at 2s
Temperature (oF)
Thermal output (me)
Thermal output, constantan
gage type “06” ( for steel )
500
400
300
200
100
-100
-200
-300
0
0
100
200
300
400
Notes: Since the advent of self-temperature-correcting strain gages, the use of the quarter bridge has become common place in experimental stress analysis. Gages made of constantan or nickel-chromium are heat-treated to have a temperature coefficient of resistance which cancels the apparent strain caused by thermal expansion in the test specimen. Strain gage manufacturers offer standard product already compensated for steel or aluminum.
As is seen in the above figure, the resulting compensation is not constant with temperature. If no further temperature compensation were done, measurements could be in error by -100  over the temperature range -40 oF to 260 oF. Further correction is possible by measuring temperature and using this curve to correct the data. Manufacturers print the polynomial coefficients to fourth order on each package of gages. With the coefficients you can conduct temperature correction in software. The spread of the correction curve (from gage to gage) over the range +32 oF to +350 oF is stated as + 0.15  at a 95 percent confidence level.

## Slide 14: Avoiding Self Heating Problems
Use just enough excitation voltage
Guideline:
steel – 2-5 watt / in2
aluminum – 5-10 watt / in2
Determine experimentally
Larger gage = less heat
Use 350 ohm gages, not 120 ohm
Notes: The change in output voltage for a given level of strain increases in direct proportion to the excitation voltage. Increasing excitation voltage thus improves signal to noise ratio, but a practical limit is reached when the ill-effects of gage self-heating become predominant.
For steel, an excitation level between 2 to 5 watts/in2 is recommended. An optimum excitation voltage is best determined by an experimental procedure. With no load applied, you should examine the zero point of the channel while excitation level is progressively raised. Once instability in the zero reading is observed, you should lower the excitation until stability returns. It is best to perform this experiment at the highest temperature over which measurements are taken.
Obviously, using larger gages and higher resistance gages (350 ohms instead of 120 ohms) decreases the power per unit area dissipated, and higher excitation voltage is possible.

## Slide 15: Use Correct Wiring
TC of copper – 2,000 PPM per oF
Quarter bridge – use three wires
+
-
+
-
RG
RL1
RL2
RL3
EO
EI
<-- wire --->
Notes: Lead wire resistance can have a great effect on the accuracy of strain gage measurements. Even if you have initially balanced the system and you have taken into account lead wire desensitization, changes in the lead wire resistance with temperature can still introduce large errors in static measurements. That this can occur even though the lead wire resistance is much smaller than the bridge resistance is not surprising considering that the temperature coefficient of copper is about 2000 PPM/oF. This is more than two orders of magnitude greater than the temperature coefficient of the gauges.
Although you could connect to the gauge with two wires in a quarter bridge circuit, lead wire temperature drift would be enormous. For this reason, static measurements are always performed using the three wire circuit, shown here. The bridge remains balanced as long as lead wire resistances RL1 and RL3 track with temperature.

## Slide 16: Gage Bonding
Surface preparation
Cyanoacrylate –
fast and easy, for short term work
Epoxy –
heat curing
Weldable strain gages
Notes: Obtaining a good bond between the strain gage and test specimen is crucial for good measurements.
A smooth surface free of oils, contamination, or oxides insures a good bond. Manufacturers usually call out a procedure of degreasing, abrading, conditioning, and neutralizing the surface before applying the adhesive.
Cyanoacrylate-based adhesives (super glue is a cyanoacrylate-based adhesive) form strong thin bonds very quickly at room temperature. The glue lines, however, become brittle after a few months, and are very sensitive to humidity. Thus you should use cyanoacrylate-based adhesives only for short term work. Most adhesives used in long term and transducer work are special two part epoxies which form very thin glue line, typically 0.1 to 0.5 mils thick. Thin glue lines improve shear strength, reduce creep, and improve thermal conductivity. A heat curing process lasting several hours is usually required.
For difficult installations, weldable strain gages are sometimes used. The gauge is mounted on a thin strip of stainless steel that is spot welded to the test specimen.

## Slide 17: Bad Bonds
Correct reading
Excessive creep
– uncured glue
Fractured, air pockets
Notes: The behavior of poor bonds is unmistakable.
Most materials exhibit some creep, but excessive creep is indicative of a poor bond. The adhesive probably never fully cured, or the glue line is excessively thick.
Another clear indication of a poor bond is the “jumpy” behavior of the last plot above. Chances are there are fractures or air pockets in the bond.
Note that most poor bonds are determined by visual inspection.

## Slide 18: Instrumentation
high
gain
excitation
+
-
bridge
completion
null
comp.
shunt
cal.
gauge
Notes: All strain gauge signal conditioners have an excitation source and a high gain amplifier. The source is usually constant voltage, but some conditioners also have constant current. The high gain amplifier is necessary to boost the weak bridge output to a level that can be easily digitized.
Other features commonly found are bridge completion, null compensation, and shunt calibration.
Bridge completion is simply the provision to replace bridge elements with fixed-valued resistors when fewer than four active strain gauges are used.
Null compensation is an adjustment which removes the initial offset at the bridge output due to imperfect balance in the bridge arms. The offset adjustment is also performed in software, but removing the offset at the bridge output can greatly increase dynamic range.
Shunt calibration provides a simple means for calibrating system gain.

## Slide 19: Shunt Calibration
+ Vin -
1743 microstrain
100K
350
350
+
Vex
-
350
350
For gain calibration
Gain and vex need
not be known
Use separate wires
if cable long
Quarter bridge –
shunt the dummy
Notes: To perform shunt calibration, a resistor of known value is placed across a bridge arm causing a know reduction in resistance. The known resistance change corresponds directly to a specific level of simulated strain. With the shunt calibration circuit enabled while the mechanical input is zero, the system gain is adjusted to obtain a reading equal to the simulated strain.

## Slide 20: Remote Sensing
Through supply Ratiometric compensation
V+
V-
V-
V+
V+
V-
-







Notes: You can avoid the problems caused by I*R drop in the lead wires entirely by using sensing of the excitation voltage. You can do this in one of two ways.
The most common way is to establish a Thevenin connection to the bridge through the sense leads. No I*R drop exists across the wires leading back to negative inputs of the amplifiers because the input impedance is very high. Through negative feedback, the amplifiers drive the excitation terminals of the bridge to voltages exactly equal to those on their positive inputs. Thus, the desired excitation voltage across the bridge is achieved with lead wire resistance having no effect.
The other way is sometimes referred to as “ratiometric compensation.” The excitation voltage at the bridge terminals is monitored through a second input channel. The voltage from this second measurement is then used in the equations which convert voltage to strain. This technique has the added advantage that drift in the excitation source is cancelled out if a change occurs while measurements are taken. Also, continuously monitoring the excitation voltage is useful in detecting complete loss of excitation. This might happen as a result of an open or short in the wiring.

## Slide 21: SCXI™-1520 Eight Channel Strain Gage Module
0 to 10 V excitation, independent per channel
Gain of 1 to 1000, 48 settings
High resolution hardware null compensation
Two independent shunt calibration circuits
Four pole low pass filter – 10 Hz, 100 Hz, 1 KHz, 10 KHz, or full bandwidth
Remote sense
Track and hold
Notes: The SCXI-1520 is the ideal module for conditioning strain gauge signals and signals from strain gage-based transducers. You can adjust the excitation voltage from 0 to 10 V in 0.625 V steps, so the user can obtain the optimum excitation setting for a particular application. You can set the DC gain of the input channel at any one of 48 settings between X1 and X1000. Gain increases by about 15 percent for each progressively higher setting. This is much finer than most signal conditioning systems, where the 1:2:5 gain scheme is used. With high resolution hardware null compensation, you remove large input signal offsets before they reach you DAQ board. Two independent shunt calibration circuits are provided for each channel. The shunt calibration resistors are in the terminal block and user changeable. You can set the four pole low pass filter at 10Hz, 100Hz, 1KHz, 10KHz, or full bandwidth so that a wide range of sampling rate is permitted with adequate alias protection. The module also features remote sense operation and simultaneous track and hold across all channels.

## Slide 22: Measuring Strain Demonstration
Notes: We demonstrate null compensation and shunt calibration of a quarter bridge circuit using an SCXI-1520 module.

## Slide 23: LVDT
Linear Voltage Differential Transformer
Coil assembly
Core
Notes: The LVDT (linear voltage differential transformer) is a sensor which measures linear displacement. All LVDTs consist of a stationary coil assembly and moveable core. The Coil assembly houses the primary and two secondary windings and has an internal bore open at both ends. Because the core is a steel rod of high magnetic permeability and is usually somewhat smaller in diameter than the internal bore of the coil assembly, you can mount the rod in such a way that no contact is made with the coil assembly. Thus the rod may move back and forth with no friction or wear.

## Slide 24: LVDT Operation
Core at center, Eout is zero
PRI
SEC2
Core
Eout
+
-
+
-
SEC1
Notes: The figure shows a cross sectional view of an LVDT. Two secondary windings are located on either side of the primary winding. The core causes the magnetic field generated by the primary to be coupled to the secondaries. When the core is at center as shown, the voltage induced in each secondary is equal, and the LVDT output (for the series-opposed connection shown in this case) is zero because both voltages cancel.

## Slide 25: LVDT Operation
Core left-of-center
Eout non-zero and in phase with input
PRI
SEC1
SEC2
Core
Eout
Ein
Eout
+
-
+
-
Notes: Move the core to the left and the first secondary is more strongly coupled to the primary than the second secondary. The greater voltage of the first secondary causes an output voltage which is in phase with the primary voltage.

## Slide 26: LVDT Operation
Core right-of-center
Eout non-zero and out-of-phase with input
PRI
SEC1
SEC2
Core
Eout
Ein
Eout
+
-
+
-
Notes: Move the core to the left and the first secondary is more strongly coupled to the primary than the second secondary. The greater voltage of the first secondary causes an output voltage which is in phase with the primary voltage.

## Slide 27: LVDT Characteristics
Advantages
  Extremely robust
  Wide temperature range
  “Infinite” resolution
  Noncontact, frictionless operation
  Perfect repeatability (no hysterysis)
  Accurate small travel, <.1”
  Absolute measurement at power on
Disadvantages
  Special signal conditioning requirements
  Limited linearity (~0.1 to 0.5%)
  Limited bandwidth
Notes: Attributes
  LVDTs sustain extreme levels of shock and vibration and operate in filthy environments with no degradation in operation because there are no delicate parts. And because the LVDT contains no active electrical components, it operates well past 150 °C, the maximum operating temperature of most encoders. The resolution of an LVDT is theoretically infinite since its operation is based on a continuous analog principal. Also, you can mount LVDTs in such a way that no contact is made to the system being measured. Thus, frictionless operation results with perfect repeatability. LVDTs that accurately measure displacement less than 0.1 in. (2.5 mm) are widely available at a reasonable cost.
Liabilities
  LVDTs require signal conditioning electronics that add to the overall cost of the system. Overall cost generally is higher than the minimum requirements for a potentiometer or encoder. LVDT linearity is on the order of 0.1 to 0.5 percent and lower than high resolution encoders.
  Finally, as compared to potentiometers and encoders, LVDTs have limited bandwidths. The bandwidth of a typical LVDT signal conditioner is one tenth the carrier frequency.

## Slide 28: Compare LVDT with Other Sensors
Pot Encoder LVDT / RVDT
Travel limited limited unlimited
Resolution infinite 8 to 16 bits infinite
Accuracy (F.S.) .1% 2-n .1% to 1%
Electronics simple digital moderate
Response excellent excellent moderate
Endurance poor-good good excellent
Hysterysis poor good excellent
Friction high low zero
Cost (low end) $10-50 $100-1000 $50-500
Notes: The chart above summarizes the relative performance of potentiometers, encoders, and LVDTs.

## Slide 29: Scaling LVDT Signals
LVDT sensitivity – signal volts per
excitation volts per unit measure of
displacement
  V / V / in. same units mV / V / 0.001 in.
  V / V / mm same units mV / V / m
  mV / V / degrees
displacement =
( Vrms from secondaries)
(Vrms to primary) * (sensitivity)
Notes: The LVDT sensitivity factor on the manufacturer’s data sheet expresses the ratio of signal voltage output per excitation voltage input per unit of mechanical displacement. In the United States, LVDT manufacturers specify the sensitivity of their LVDTs in units of “volts per volts per inch” or “millivolts per volts per 0.001 inch.” In either case, the units are the same. For example, an LVDT sensitivity of 10 V/V/ in. is the same thing as 10 mV/V/.001 in. Similarly, in the metric system, “volts per volts per meter” is actually the same physical unit as “millivolts per volts per millimeter.”
The formula shown above gives the measured displacement as a function of the measured voltage. “Vrms from secondaries” is the signal from the LVDT; “Vrms to primary” is the excitation voltage applied to the LVDT, and “sensitivity” is the manufacturer’s sensitivity factor from the LVDT data sheet.

## Slide 30: Two Point System Set-up and Calibration
1) Mechanically zero sensor, hardware at low gain
2) Mechanism at full scale ---> adjust hardware gain for full scale output.
3) Mechanism at zero ---> software offset
4) Mechanism at full scale ---> span constant
Notes: National Instruments recommends using the procedure described above when you install a new LVDT, gauge head, or signal conditioning channel. Consider the example of installing an RVDT to measure the steering angle of an automobile. We assume that a linkage has been installed between the RVDT and some moving component of the steering system.
Step 1: Center the car’s steering wheel and set the hardware gain at a low setting so that in-range readings are guaranteed. Rotate the RVDT in its mount until a reading of zero is obtained. Then lock down the RVDT in its mount.
Step 2: Turn the car’s steering wheel all the way to the left and increase the instrumentation gain to obtain a near full scale voltage at the digitizer.
Step 3: Turn the car’s steering wheel back to center and zero the residual offset in software.
Step 4: Turn the car’s steering wheel all the way left again and set the software scaling factor to obtain the car’s actual maximum steering angle.
The system is now fully calibrated with the sensor properly aligned and the hardware gain set for the best possible resolution.

## Slide 31: LVDT Response Curve
Typical linearity – travel linearity
travel
output
50 % 0.15 %
100 % 0.25 %
125 % 0.35 %
150 % 0.50 %
Notes: Let us now consider the shape of the LVDT response curve.
The figure shows a typical response. As the core reaches the limit of the displacement range specified by the manufacturer, the slope begins to decrease. The LVDT becomes increasingly nonlinear as the core travels beyond the specified limit.
If the core is travels far enough beyond the manufacturer specified limits, the slope changes sign, thus forming an “S” curve. An S curve response can be highly undesirable in control applications and should be duly noted by the control engineer.

## Slide 32: Linearize the Curve with LabVIEW™
Macro-sensors E750-100 ( 100 % travel is .1” ) polynomial y = a0 + a1*x + a3*x3
travel before after
50 % 0.02 % 0.01 %
100 % 0.4 % -0.12 %
150 % 2.0 % -0.16 %
200 % 3.0 % 0.34 %
Notes: You can greatly reduce the nonlinearity in the response curve for an LVDT by using low order polynomials. Not only can you improve the linearity in the manufacturer’s advertised range, but you can also improve the useful operating range of the device. We were able to increase the useful range of a + 0.1 in. travel LVDT to + 0.2 in. by using only a third order correcting polynomial.

## Slide 33: RVDTs and Resolvers
RVDT
  Rotational version of an LVDT
  +30 o to +70 o range
  Phase reversal every 180 o
Resolver
  +360 o range
  Two coils 90 o apart
Notes: RVDTs are the rotational version of LVDTs and generally operate over a range of + 30° to 70o. They are available in servo mount and may rotate through 360° without encountering a stop. Like there LVDT counterpart, they have no brushes (all windings are stationary).
A position sensor very similar to the RVDT is the resolver, which has a sine stator winding and a cosine stator winding. By demodulating each secondary coil separately, an angle measurement through 360° is possible.

## Slide 34: Resolver
Notes: The two stator windings and rotor winding are shown in the drawing. The arrangement of the windings is such that the sign and magnitude of the voltage induced in a stator winding from the rotor winding is the sine or cosine of shaft angle. Having two stator windings separated at 90º permits measurement of rotational position at constant resolution over 360° and with no 180° ambiguity.
Typical operating frequencies are between 2 and 20 KHz, and typical excitation levels are between 1 and 20 Vrms. In practice, frequencies and excitation levels may vary widely with little change in sensor performance.
Resolvers are available in standard servo housings. Brushless resolvers are also available which have complete clearance between the rotor and stator. You mount the rotor to the shaft and the stator to the fixed, non rotating part of the system. Thus, the sensor imparts no friction on the shaft.

## Slide 35: Synchros
Three secondaries at
120 ° spacings
Typically 60 or 400 Hz,
26.5 or 115 Vrms
Not 3 phase devices
Notes: Synchro transmitters are rotational sensors with three windings spaced at 120 º angles. Synchros typically operate at 60 or 400 Hz, and at excitation voltages of 26.5 or 115 Vrms. Synchros are often thought to be three phase devices. They are not. Three AC signals are involved, but all are at the same phase. Position information is carried only in the amplitudes of the three signals, not the phases.

## Slide 36: LVDT Instrumentation
Dedicated LVDT signal conditioner
  SCXI-1540 – eight channel LVDT module
Part of sensor – DC LVDT
  More expensive, not as robust
  Must supply power
Notes: You have several alternatives when choosing instrumentation for LVDT signal conditioning.
You can choose to use a product specifically designed for LVDT signal conditioning such as National Instruments SCXI-1540 eight channel LVDT signal conditioner. The SCXI-1540 has all features necessary to accurately condition common LVDTs, RVDTs, and resolvers in the flexible SCXI signal conditioning architecture.
You can purchase a so-called DC LVDT which has the signal conditioning built-in to the sensor. The output level is several volts and may be interfaced directly to a general purpose data acquisition board. Note that the operating temperature range is much less than a passive LVDT since the electronics cease to function at about 150 °C.

## Slide 37: Resolver Instrumentation
Resolver to encoder converters
  No absolute position at power on
Multi-channel LVDT signal conditioner
  SCXI-1540 eight channel LVDT module
  Synchronize channels
Notes: Due to the great popularity of encoders, there are products available which interface directly to resolvers but output the same digital signals as encoders. This approach may be convenient, but the resolver’s advantage of making absolute measurements immediately after power on is lost.
Alternatively, you may use an LVDT signal conditioner if there are at least two channels that you can synchronize to one another. The inputs of each of the two channels are connected to each of the stator windings. You may use either channel’s excitation output to excite the resolver’s rotor winding. In software, you obtain absolute angular position by simply using a four quadrant arctangent function (a function that has two arguments corresponding to the relative values of the sine and cosine for the angle which is returned). Note there is no need to scale voltage anywhere in the system in order to obtain an angle measurement.

## Slide 38: Four Wire LVDT Instrumentation
DC
output
X
LPF
Diff. amp
Excitation
Demodulate
Filter
+
-
3 Vrms
10 KHz
3 Vrms
full scale
Demodulation
signal obtained
directly from
excitation source
Four wire connection sensitive to phase shift in wiring and sensor but
insensitive to common mode noise
Notes: The figure shows a block diagram of an LVDT signal conditioner with a four wire hook-up. The hardware generates a sine wave which is applied to the LDVT primary. The return signal is obtained from the LVDT secondaries wired in series and applied to a differential amplifier where common mode noise picked up in the cable is rejected. The signal is then synchronously demodulated in a multiplier block. Notice that the demodulating signal comes directly from the excitation source. The demodulator output is passed through a low pass filter to remove high frequency ripple. The resulting output is a DC voltage proportional to core displacement.

## Slide 39: Four Wire Nulling Problem
4-wire demodulation
signal, same as excitation
A
null voltage – (A-B) signal partially in-phase
with the demodulation signal. Null is shifted.
B
A-B
Phasor diagram
A – 1st secondary voltage
B – 2nd secondary voltage
A-B – Four wire return signal
Notes: A four wire demodulation scheme is subject to shift in the indicated null output. The phasor diagram shows why. When the LVDT core is at center, the magnitude of the the two secondary voltages are equal. In an imperfect LVDT, however, the two secondary voltages are not necessarily in phase. The small phase difference between A and B causes a non-zero (A-B) signal, even though the core is at center. This is referred to as the null voltage. It is sometimes specified in a manufacturer’s data sheet. Where there is also a large phase shift in the LVDT, the null voltage can shift to the extent that a substantial component is in phase with the demodulation signal. As a result, a small DC signal is present at the demodulator output, and a zero reading is not obtained when the LVDT core is actually at center.

## Slide 40: Five Wire LVDT Instrumentation
DC
output
X
LPF
Diff. amp
Excitation
Demod.
Filter
+
-
+
+
A+B
Demodulation
is common
mode signal
from LVDT
Five wire connection rejects LVDT null voltage – A-B signal 90 ° out of phase with A+B demodulation signal
A
B
A-B
Notes: The five wire connection avoids the null shift problem altogether.
Operation is identical to that used for the four wire connection except that the demodulation signal is obtained from the common mode signal returned from the LVDT secondaries. Notice that the fifth connection simply grounds the center connection between the secondaries so that the common mode return signal is properly referenced to ground.

## Slide 41: Five Wire Connection – Null is Good
5-wire demodulation
B
A-B null voltage 90 ° with respect to
demodulation signal. Output is zero.
A
A-B
A+B
Phasor diagram
A – 1st secondary voltage
B – 2nd secondary voltage
A-B – return signal
A+B – signal used as phase
reference for demodulation
Notes: In the phasor diagram, we see that the (A+B) signal is shifted 90° with respect to the (A-B) null voltage. The A+B signal is the signal used for demodulation. The demodulator output is zero in this case since both inputs are separated in phase by 90°. Hence, the 5 wire connection avoids the null shift problem altogether. An output of zero is obtained when the core is exactly at center.

## Slide 42: Four Wire vs Five Wire
Four Wire
  Use with 4-wire sensors
  Use when cable run is short
  Use when environmental noise is high
Five Wire
  Use when LVDT null voltage is high
  Use when cable run is long
Notes: In the majority of cases, either the four wire or five wire connection works equally well. You generally obtain better null performance with the 5-wire connection and should opt for this connection whenever the LVDT null voltage is high or the cable run is long. The manufacture’s data sheet sometimes gives the null voltage. Note that the null voltage can be much higher than listed if you use the LVDT at a frequency other than that for which it is specified.

## Slide 43: SCXI-1540: 8-Channel LVDT Module
2.5, 3.3, 5.0, or 10 KHz at 1 or 3 Vrms
0.05 to 6 Vrms input levels
Four or five wire connection
Synchronize multiple channels
Field calibration
Notes: The SCXI-1540 has all features necessary to accurately condition common LVDTs, RVDTs, and Resolvers in the flexible SCXI signal conditioning architecture. The excitation frequency may be set at 2.5, 3.3, 5.0, or 10 KHz and the amplitude set at 1 or 3 Vrms. You can set the gain over a wide range to accommodate input signal levels ranging from 0.05 to 6 Vrms full scale. Both four wire and five wire connections are supported. You may synchronize the excitation of multiple channels so that beat frequencies are avoided. (A beat frequency occurs when the excitation signal from one channel couples into the input signal of another. Coupling happens in multiconductor cable). Lastly, you can calibrate the SCXI-1540 in the field though driver software.

## Slide 44: SCXI-1540 External Sync
CH0+
CH0-
SYNC0
EX0+
EX0-
CH1+
CH1-
SYNC1
EX1+
EX1-
Resolver
Slave –
configure for external sync.
Master –
do not configure for external sync.
Notes: This slide shows how to synchronize a slave channel to a master channel. Synchronization is required when interfacing to a resolver.

## Slide 45: LVDT Demonstration
Notes: We now demonstrate the two point system set-up and calibration procedure using an LVDT with the SCXI-1540.