## Slide 1: Measuring Vibration Accurately
Paul Lennous
Hardware Engineer
Thurs Aug 17
12:00-1:15 p.m., 3:30-4:45 p.m.
Cypress (8B)
Notes: At National Instruments, our goal is to make computer-based measurement and automation technologies to solve customer applications. While we strive to improve these products, we also strive to educate our customers about fundamental technologies and to assist them in their applications. Our goal in this presentation is to educate our customers on the fundamentals of measuring vibration as accurately as possible.

## Slide 2: Agenda
Signal conditioning requirements
Setting up a system
Demonstration
Notes: In this presentation, we discuss in detail the various elements in a vibration data acquisition system, as well as their individual tradeoffs. We then cover the procedure to determine exactly how to set up such a system. Finally, we demonstrate both a single channel data acquisition system and a two dimensional system. Finally, we will demonstrate both a single dimension and a two dimension data acquisition system.
For brevity, this presentation will discuss the requirements for using current excited active accelerometers.

## Slide 3: Signal Conditioning Requirements
Accelerometer
Current source
Grounding
AC coupling
Instrumentation amplifier
Low pass filter
Simultaneous sample and hold
Notes: The signal conditioning circuitry for measuring vibration is fairly straightforward. It consists of the accelerometer itself, a current source to excite the accelerometer, proper grounding to eliminate noise pick-up, AC coupling to remove DC offsets in the system, an instrumentation amplifier to boost the accelerometer’s signal level, a low pass filter to reduce noise and prevent aliasing in the data acquisition system, and finally, simultaneous sample and hold circuitry to keep the signals properly timed with respect to each other.

## Slide 4: Accelerometer
Converts an acceleration into a voltage signal
Notes: The most obvious starting point in the system is the accelerometer itself. It is simply a transducer that converts an applied acceleration to its body into a proportionate displacement of voltage on its output.

## Slide 5: Types of Accelerometers
1D or linear
  Measure acceleration along only one axis
3D (tri-axial) accelerometers
  Measure acceleration along all three axes
  Have three separate outputs
1D
3D
Notes: Accelerometers come in two axial types. The most common accelerometer measures acceleration only along a single axis. This type is often used to measure mechanical noise levels. For example, you may use it on a compressor, and the results may help determine if the compressor is operating properly, or is in need of maintenance.
The second type is the tri-axial accelerometer. This accelerometer is capable of creating a three dimensional vector of acceleration in the form of orthogonal components. This type is used when it is necessary to determine the type of vibration (lateral, transverse, rotational, etc.) that a component is undergoing, or if the acceleration of the component in any direction must be known. Installation of this type of accelerometer must be done more carefully in order to get an accurate indication of the direction of acceleration. This type of accelerometer also has three separate outputs, each output needing its own signal conditioning.

## Slide 6: Active Accelerometer Anatomy
Front end circuitry supplied inside of sensor
Energized with a constant current source
Example – ICP by PCB piezotronics
Notes: The active accelerometer contains a piezoelectric crystal bonded to the casing of the accelerometer. Bonded to the other side of the crystal is a small seismic mass. When the accelerometer is moved, this mass either compresses or stretches the crystal, inducing a charge on its surfaces. This charge is picked up by a small charge sensitive amplifier built inside the sensor. This amplifier is powered by an external constant current source, and modulates its output voltage with respect to the varying charge on the piezoelectric crystal. Thus, the accelerometer uses only two wires for both sensor excitation (current) and signal output (voltage).
This also means that the active accelerometer has a relatively low output impedance, usually on the order of a few kilo-ohms.

## Slide 7: Active Accelerometers
Advantages
Less noise sensitivity
Built-in signal conditioning circuitry
Simple to use
Disadvantages
Fixed signal sensitivity
Limited temperature range (~250 °F)
Notes: The active accelerometer has advantages over the conventional passive element accelerometer. Because the active accelerometer has a relatively low output impedance, it is less sensitive to noise being picked up in its cabling. The built in signal conditioning circuitry further reduces noise pick-up by being placed as close as possible to the crystal. The net result is that the active accelerometer is easier to use, especially in very noisy environments.
Despite these advantages, there are some disadvantages which may prevent one from choosing an active accelerometer. Because the front end circuitry is provided inside the accelerometer, it has a fixed signal sensitivity independent of the excitation current. This circuitry also limits the usable temperature range of the accelerometer since most of the signal conditioning circuitry will fail above approximately 250 °F.

## Slide 8: Accelerometer Parameters
Measuring range
  Maximum measurable accelerations
  Given in + g’s
Sensitivity
  Ratio of output voltage per g of acceleration
  Given in mV/g
Resonant frequency
  Frequency that accelerometer starts ringing
  Given in kHz
Notes: Active accelerometers have three major parameters to consider when selecting which model to use.
The first parameter is the sensor’s measuring range. This number represents the maximum acceleration that the sensor can reasonably measure, represented in +g’s. (One g is one Earth gravity.) Thus, a sensor with a measuring range of +5 g’s can measure 5 g’s of acceleration in either direction along its measuring axis.
The next parameter is sensitivity. This is the sensor’s “gain” given as a ration of volts per g. Thus a sensor with a sensitivity of 200 mV/g displaces 300 mV when undergoing 1.5 g of acceleration.
The final parameter is the accelerometer’s resonant frequency. This represents the frequency that the sensor rings at if bumped abruptly, much like a goblet rings when tapped. Because the sensor wants to naturally resonate at this frequency, it is very important to make sure that you do not use the accelerometer to measure any signals near the resonant frequency.

## Slide 9: Current Source
Most accelerometers want between 2 and 20 milliamps of excitation current
Source should have enough compliancy to drive accelerometer through its full range
Three common types – battery and resistor, battery and current diode, and active source
Notes: The power source to excite the accelerometer is a current source. Most accelerometers can operate on a current between 2 and 20 mA. While it may seem more economical to use a 2 mA source, it should be noted that the maximum frequency and amplitude that an accelerometer can drive is directly related to the current that is exciting it. Thus, long cable lengths need higher currents in order to drive high frequencies at large voltage ranges. Systems with short cables or low frequency measurements can tolerate being driven with lower currents.
In addition to the amount of current the source can supply, it must have a high enough voltage compliancy to power the sensor over its entire operating range. Voltage compliancy is a measure of the voltage range that a current source properly functions. For example, a 24 V compliant current source delivers a constant current over a range of 0 to 24 V across its terminals.
Most accelerometer systems use one of three types of current sources, resistor and a battery, diode and a battery, and active. These will be discussed in detail.

## Slide 10: Battery and Resistor
Very simple and inexpensive
Usually three 9 V batteries and a resistor
Very nonlinear response
Drifts with time
Notes: The simplest current source used for exciting accelerometers is a battery and resistor system – three 9 V batteries in series with a resistor.
While it is very simple and inexpensive, it is also very inadequate. The resistor maintains a constant current flow as long as there is no signal in the system. The moment a signal changes the voltage level at the output terminals, the current also changes in accordance with Ohm’s law. To put it bluntly, the battery and resistor current source is 0 V compliant by its very design! This dynamic changing of current betrays itself as non-linearity in the system’s response. As a result, this current source produces a lot of harmonic distortion.
The necessarily low value of the resistor to supply the necessary current means that the dynamic impedance of this current source is also quite low, usually around 5 kilo-Ohms. (An ideal current source has an infinite dynamic impedance.)
Finally, the nominal current and peak output from this current source slowly changes over time as the batteries discharge.

## Slide 11: Battery and Current Diode
Also simple and inexpensive
Usually three 9 V batteries and a current diode
Better response than resistor circuit
Performance drops with frequency
Notes: The battery and current diode is an improvement on the battery and resistor current source. It replaces the resistor with a constant current diode. This current source is much more compliant than the resistor source, and produces much less harmonics due to non-linear distortion. In addition, it has a much higher dynamic impedance, usually on the order of a half a meg-Ohm. However, because of parasitic capacitances in the current diode, this current source has a drop-off in performance at higher frequencies.
Like the battery and resistor current source, the battery and current diode source slowly drops in performance over time as the batteries discharge, lowering the current source’s voltage compliancy.

## Slide 12: Active Current Source
Closest to an ideal source
More expensive than the other sources
Gives the best performance
Notes: The active current source is the best source to use for energizing accelerometers. Its compliancy can be very high, and does not degrade with either time or frequency. There are no batteries to replace, thus eliminating the need to service it. Depending on the complexity of the design, it can have a dynamic impedance on the order of a half a meg-Ohm and up.
The only drawback is that it is more expensive than the other current sources because of its greater complexity.

## Slide 13: Sensor Grounding
Essential to minimize noise
If done improperly, can make matters worse
Proper grounding is simple to determine
  If sensor is floating, ground the system’s input
  If sensor is grounded, do not ground the system’s input
Notes: Proper grounding is essential to minimizing noise in the vibration measuring system. Improper grounding, however, can cause more noise problems than it can eliminate. There are two main causes of noise as a result of grounding, ground loops or floating nodes.
A ground loop occurs when both the sensor and the signal conditioning input are grounded. The loop forms between the sensor cable and the building ground. This huge loop is perfect for picking up any induced currents produced by machinery, lights, and more. Because the loop resistance is finite, the result is a large 50 or 60 Hz common mode noise signal on the signal conditioning’s input. Thus, a large common mode noise is picked up in the system. If the situation is bad enough, the common mode voltage can get high enough to saturate the signal conditioning’s front end, causing a multitude of noise problems.
Likewise, floating nodes which occur when neither the sensor nor the signal conditioning system are grounded, induced charges can build up on the sensor to the point of overloading the front end of the signal conditioning system.
The solution is simple. Either ground the sensor or the signal conditioning input, but not both.

## Slide 14: AC Coupling
Removes large DC offset and drift
Increases signal resolution
Make sure that low frequency rolloff is low enough
Notes: The next stage, AC coupling, removes the DC voltage offset that is present on the accelerometer. This dramatically increases the resolution of the system, because the system no longer has to accommodate the offset. For example, an accelerometer with an output voltage swing of +2 V and an offset voltage of 10 V would require a digitizer with a 12 V input range, despite only using 4 V for the actual signal. The same sensor AC coupled would require a 4 V input range on the digitizer, increasing the resolution by a factor of three.
In addition, the AC coupling removes the long term DC drift that accelerometers have due to age and temperature effects.
The only care one needs to have with AC coupling is to make sure that the low frequency roll-off is lower than the range of frequencies desired.

## Slide 15: Instrumentation Amplifier
Amplifies signal from sensor
Rejects common mode noise
Gain depends on accelerometer sensitivity and signal strength
Notes: The instrumentation amplifier serves three purposes. It amplifies the signals coming from the accelerometer, if necessary. It also removes any common mode noise remaining in the system. Finally, it provides a low enough output impedance to drive the rest of the signal conditioning system.
Some instrumentation amplifiers have programmable gains, so they can accommodate several different input voltage ranges, thus increasing their versatility.

## Slide 16: Low Pass Filter
Reduces noise in all systems
Used for anti-aliasing in data acquisition systems
Three common types include:
  Elliptic
  Butterworth
  Bessel
Come in any number of orders or poles
Notes: After amplification, the signal is filtered to remove any high frequency noise components. This low pass filtering also prevents any aliasing of high frequency components when the signal is digitized.
Low pass filters have two main characteristics, type and order. The filter type determines how the filter behaves in the pass band, stop band, and its phase response. Three filter types, the elliptic, Butterworth, and Bessel filter will be discussed in detail.
The filter order determines the rate of frequency roll-off in the stop band.

## Slide 17: Elliptic Filter
Maximum rolloff in the stop-band
Used in very noisy environments
Large phase error
Notes: For a maximum roll-off just after the pass band, one should choose the elliptic filter. This filter is designed to maximize roll-off at the expense of other parameters. As the graph of frequency response shows, the elliptic filter rolls off much faster than the Butterworth or Bessel filters. This would be the filter to choose in very noisy environments, or when a fast roll-off is needed.
The elliptic filter’s main drawbacks are some ripple in attenuation in the pass band, and poor phase response, both which will be discussed later.

## Slide 18: Butterworth Filter
Maximum flatness in pass-band
Used when a flat gain response is necessary
Moderate phase error
Notes: If maximum flatness in the passband is desired, then the Butterworth filter is the ideal choice. The Butterworth filter is designed to have a maximally flat gain in the passband at the expense of phase error and roll-off. It is nearly perfectly flat for most of the passband, compared to the rippled elliptic filter or the immediately falling off Bessel filter. This filter would be a good candidate for audio applications.
The Butterworth filter’s drawbacks are having a somewhat more gradual roll-off than the elliptic filter, and more phase error than the Bessel filter. Otherwise, it is a very popular general-purpose filter.

## Slide 19: Bessel Filter
Minimum phase error in pass-band
Used to preserve phase or timing
Very gradual rolloff
Notes: If signal phase or timing needs to be preserved, then the Bessel filter is the clear choice. The Bessel filter was designed to have a minimum of phase error at the expense of pass band flatness and stop band roll-off. It is the filter to use when timing is critical, because all signals pass through it at a constant delay rate, despite their frequency components.
The Bessel filter’s drawbacks are an immediate roll-off occurring in the pass band, despite having a gradual roll-off in the stop band.

## Slide 20: Filter Step Responses
Elliptic – large overshoot and ringing
Butterworth – some overshoot and ringing
Bessel – no overshoot and ringing
Notes: Sometimes it is necessary to refer to the step response of each filter to determine which filter to use. Some systems are sensitive to overshoot and ring. The graph above can be helpful in selecting the right filter. As the graph shows, the elliptic filter has the largest amount of overshoot and ring for a step response. The Butterworth is next, with a little less overshoot and ring, and a little faster rise time. If the system is very sensitive to overshoot and ring, then a Bessel filter would be the best choice, having no overshoot or ringing, and having the fastest rise time.

## Slide 21: Filter Order Effects
Higher order filter = steeper rolloff
Higher order filter = more expensive
Balance cost versus performance
Notes: As previously stated, a filter’s order affects its rate of roll-off on the stop band. The higher order the filter, the steeper the roll-off. Roll-off refers to the rate in which signals are attenuated with increasing frequency. For every order in the filter, it rolls off 6 dB/octave or 20 dB/decade. Thus, a four pole filter has a roll-off of 24 dB/octave or 120 dB/decade in the stop band.
Higher order filters are more expensive to build, but they may be necessary in environments that have a large amount of high frequency noise, or when a noise component is very close in frequency to the desired bandwidth.

## Slide 22: Simultaneous Sample and Hold (SSH)
Many channels = large delay between channel scans
Delay looks like a shift in time in data
SSH “freezes” time while all channels are scanned
Result – no time skew between channels
Notes: The final step in signal conditioning before digitizing is to pass through a simultaneous sample and hold circuit. This circuit reduces timing errors in multichannel systems.
Consider a multichannel system without a sample and hold, and assume that two channels are monitoring the same signal. If the channels are digitized one at a time, some time passes between the digitizing of the first channel and the second. Thus there is a slight skewing in time on the output of these two channels. If both channels are plotted, the second channel looks like its signal lagged the first channel, despite both channels looking at the identical signal.
By putting the SSH in the path, the SSH circuit “freezes” time by holding all channels at their last sampled level until the digitizer finishes digitizing all of them. The result is that every channel in the scan represents the exact same point in time. Thus, the two channels have overlapping outputs, just as they should.

## Slide 23: Setting Up the System
Choose an accelerometer
Provide an adequate current source
Ground the system appropriately
Adjust the instrumentation amplifier gain
Select a filter type, bandwidth, and order
Use SSH for multiple channel systems
Account for sensor sensitivity in the data acquisition system
Notes: Now that each element in the system has been described, one may set up a vibration measurement system in a straightforward manner.

## Slide 24: Choosing an Accelerometer
Sensor’s measuring range must be larger than application range
Sensor’s resonant frequency must be much higher than measured system’s
Sensor’s sensitivity should be as large as possible
Notes: From the range of accelerometers available, narrow down the list to those which have an measuring range that encompasses the desired application’s range. From the remaining selection, eliminate any that have a resonant frequency in or near the desired bandwidth of frequencies. Finally, choose an accelerometer that has the largest sensitivity that the signal conditioning system can accommodate. The maximum sensor voltage output can be determined by multiplying the measuring range by the sensitivity.

## Slide 25: Grounding the System
If floating sensor, ground the system
If grounded sensor, leave system floating
Beware using grounded accelerometers on single-ended systems
Notes: The next task is to properly ground the system. If the sensor is grounded, use a differential input conditioning system. If the sensor is floating with respect to ground, ground the signal conditioning system’s inverting input. If you do not know if the sensor is floating or grounded, you can simply determine the answer by measuring the resistance between the sensor’s cable shell and ground.

## Slide 26: Instrumentation Amplifier Gain
Input range = sensitivity * measuring range
Adjust instrumentation amplifier to accommodate input range
If input range is too large, select a less sensitive accelerometer
Notes: Determine the necessary input range for the instrumentation system by multiplying the sensor’s sensitivity by its measuring range. Set the instrumentation amplifier’s range to the maximum gain that accommodates the input range. If you forgot to do this earlier, you may find that your sensor has too great an output range, and you may have to select a less sensitive accelerometer.

## Slide 27: Select Filter Type
Choose type of filter based on desired parameters
Select a bandwidth just above the desired frequency range
For slower sampling rates and very noisy environments, choose a higher order filter
Notes: Choose the necessary filter type depending on the signal conditioning needs. A Bessel is recommended for multichannel systems. Set the filter’s bandwidth to the lowest possible setting above the desired signal’s bandwidth. If operating in a very noisy environment, it may be necessary to choose a high order filter.

## Slide 28: Adjust for Sensor Sensitivity
Raw data output is in volts, not g’s
Divide voltage by sensitivity to get acceleration in g’s
Notes: Finally, be sure to account for the accelerometer’s actual sensitivity in the data acquisition system. Do not assume that an accelerometer has its nominally rated sensitivity given in the specifications sheet. Accelerometers are serial numbered, and each one has its own sensitivity measured.
To get the value of acceleration in g’s simply divide the signal conditioning system’s voltage output by the accelerometer’s sensitivity.

## Slide 29: Conclusions
Choose your sensors to best match your application
Select each stage of the system based on desired signal and surrounding environment
More information
  ni.com
Notes: We have given you information that should prove helpful in selecting the best sensor for your application, as well as information regarding the best choice of instrumentation and useful techniques to obtain accurate vibration measurements.
Visit our Web site ni.com for more information on our products as well as application white papers on subjects such as the measurement of strain and position.