PCI-Based Control System for Laser Etching an IntraOcular Lens

by

Roy E. Williams, Ph.D.

President

FEO Medical, Inc.

United States

Category: 
Biotechnology/Life Sciences

NI Products Used:

PCI-1409 Frame Grabber

PCI-6034E DAQ Card

PCI-6518 DIO Card

LabVIEW

LabVIEW Database Connectivity Toolset

Internet Toolkit

Vision Development Module

The Challenge:

Developing a high resolution, deep UltraViolet (UV) laser etching system to produce highly complex, custom acrylic or silicone Intraocular Lenses for implantation in the human eye after cataract removal, or for refractive correction, thus allowing the physician to better fit the patient versus present methods.

The Solution:

Creating a high-speed, mixed signal PCI data acquisition and control system, which interfaces to an ophthalmology-based topographic wave-front measurement systems, a UV-enabled digital micromirror device and a deep-UV laser system, all through a NI LabVIEW software program.

Abstract:

The human visual system is complex - consisting of the cornea, lens, and retina. Over time, the lens forms cataracts preventing it from focusing light on the retina, resulting in vision loss. Surgeons correct cataracts by removing the lens and inserting an artificial lens (IOL).  Current methods for manufacturing IOLs are limited to producing simple refractions - sphere and cylinder.  Physicians can now measure higher-order aberrations of the eye, but these cannot be corrected by current manufacturing.  It would be advantageous to produce IOLs to correct these aberrations, providing the patient with superior vision.  No systems existed to accomplish this.

Introduction:

A well-known cataract surgeon commissioned FEO Medical (FEOM) to develop a prototype of such a system providing: mass manufacturing capability; interface to commercial eye measurement systems; data import from remote locations; and scalable, maintainable, user-friendly software. The customer required delivery of the prototype in nine months. FEOM could only spare one physicist and one electro-optical engineer. It took a commercial refractive surgery company nearly two years, with multiple engineers, to interface one wavefront measurement system to their refractive laser surgery system. Based on previous work with NI hardware, and LabVIEW software, for an eye tracker on a laser refractive surgery system, FEOM felt confident it could provide the deliverable.

System Description:

Each part of the system had a unique requirement, which was met by NI hardware and LabVIEW. In the Eye Measurement interface section, the software had to interface with many systems using different communication protocols and data sets. In the Profile Generation section, the software had to calculate several equations and merge the results to construct a 3-D model of the etch pattern. The IOL is etched using (1) the Laser Control section, where the software had to precisely control the UV laser fire and laser energy, and monitor the beam uniformity; and (2) the Laser Beam Shaping section, where the software had to control 786,432 micromirrors on the UV-enabled DLP device (UV-DLP).

INTERNET

OPTICS

IOL BLANK

DMD

Optical System and Target

NI PCI-6518

Isolated DIO

NI PCI-1409

Frame Grabber

DEEP UV IMAGER

& ENERGY METER

EXCIMER UV LASER

Display

PC System High Speed USB

Laser Control

PC System

RS-232

Eye Measurement Interface

Profile Generation

DMD

Control

Internet Interface

NI PCI-6034E

DAQ

Beam Homogeneity

PC System

Clinic 1

Clinic 2

Clinic 3

Clinic 4

This block diagram describes the system, showing LabVIEW software, NI HW and the external systems’ interface.

ISP

(Database)

HW

LabVIEW Code

IOL Etch Control

Remote Clinics

Eye Measurement Interface:

The first task is to measure the patient’s visual system. This is performed with systems in clinics located remotely from the IOL etching system. Systems for measuring corneal topography, or wavefront components, output their data in a proprietary format, which is available from the manufacturer. The remote clinics upload the patient’s measurement data to a common ftp-based database. The LabVIEW-based IOL etching system accesses the ftp database (written using the LabVIEW Database Connectivity Toolset) via the Internet to get the patient’s data (using National Instruments Internet Toolkit).

The data is then parsed based on the manufacturer’s format. This format is contained in a LabVIEW configuration file, which allows for the easy addition of other measurement systems’ files as they become available. The parsed information is saved to the hard disk and then passed to the Profile Generation code.

Profile Generation:

Next, the system must generate a 3-D volumetric profile based on the measurement data. Profile generation begins with a basic IOL formula to calculate the IOL power. The IOL equations are implemented using LabVIEW’s math functions. Once the IOL power is known, the code converts the power to a radius-of-curvature (ROC) based on an index of refraction. From the ROC, a 3-D shape defining the basic IOL shape is generated. This profile is then altered by adding the higher order aberrations captured by the wavefront system (e.g., spherical aberration and coma). This 3-D profile is displayed at every step, using LabVIEW 3-D surface graph functions, to allow the surgeon to review the profile (figure below). The resultant 3-D profile is then sent to the Laser Beam Shaping code.

IOL Etch - Laser Beam Shaping:

Based on the etch depth per laser pulse for each IOL material, the code “slices” the 3-D profile into a number of individual 2-D layers (e.g., for a spherical shape, each layer will be represented by a “circle,” each with a different diameter). Each individual 2-D slice is presented as an “image” to the UV-DLP device via a high-speed USB port.  Each mirror will then be set to reflect or not reflect the laser energy, based on the “image.” The 2-D mirror configuration is read back to confirm correct positioning. Each mirror in the device’s 1024x768 array is 13-µm2, and thus, when the laser is fired, nearly 768,000 individual 13-µm2 laser beams are directed to the IOL surface.

IOL Etch - Laser Control:

With the mirror configuration set, the excimer UV-laser is ready to fire. The LabVIEW code first sends the desired energy setting to the laser via the RS-232 link. When the laser is ready, it enables a digital “ready” signal. The code reads the “ready” signal and sends a digital “fire” signal to the laser. Once the laser fires, it enables a “fired” signal. Upon receiving the “fired” signal, the code reads the laser energy output, via the RS-232 link, and a second energy reading from the deep-UV Imager via the PCI-6036E DAQ board. The code continually monitors these two energy readings and adjusts the laser accordingly to maintain the required energy (PID loop).

When the laser pulse is fired, and reflected from the UV-DLP device mirrors to the IOL blank, the surface is etched as described by each 2-D slice “image.” For the example above the result will be an IOL with a spherical shape.

Note that this laser uses digital I/O and RS-232 for monitoring and control. Future lasers may use Ethernet or USB – LabVIEW allows easy modification to use these new laser interfaces.

IOL Etch – Laser Beam Uniformity:

Laser beam uniformity is checked before each procedure begins.  It may also be checked during the procedure to ensure optimum results.  The system checks beam uniformity using the deep-UV Imager.  Here, an NI PCI-1409 captures a single laser shot from a custom designed UV camera.  Using LabVIEW Vision Development Module, the laser beam intensity is calculated at every point (individual camera sensor).  This intensity information is then used in the laser beam shaping section to decide which micromirrors will be turned OFF, and how long they will be turned OFF, in order to accommodate for any beam non-uniformities.  This method allows for the smoothest etch of the IOL surface.

This figure describes the Profile Generation code showing a vertical cross section of the 3-D profile.

Conclusion:

FEOM successfully constructed the system and etched an IOL blank in August 2005, thus meeting the customer’s deadline.  FEOM continues to refine the system while preparing it for manufacture. FEOM is also implementing an automated laser alignment and calibration tool for the system, simplified by LabVIEW’s scalability.