System Based on NI FieldPoint Used to Monitor the Feeding Activity of Laboratory Animals

Category:

Biotechnology/Life Sciences

Authors:

Matthew R. Williams, Mack Information Systems, Inc. and PV Knowledge, LLC

Edward A. Ulman, PhD., Research Diets, Inc.

NI Products Used:

FieldPoint

cFP-2010,

cFP-SG-140,

LabVIEWt

LabVIEWt Real-Time Module

The Challenge:

Monitor and record the feeding activity of laboratory animals, including total consumption, meal duration, meal construction, and the effects of different treatment protocols on feeding behavior.

The Solution:

Build a multiple-channel data acquisition and analysis system based on National Instruments FieldPoint hardware with a PC-based user interface. The NI FieldPoint controller operates as a stand-alone device, recording feeding activity according to a user-configurable algorithm. The PC provides further data analysis, report generation, and long-term storage.

Abstract

The feeding activity of laboratory animals, typically mice and rats, can be significantly influenced by treatment protocols. Being able to quantify these influences is critical to understanding the effects of various drugs under investigation. This system provides researchers with data describing exactly when and how much each animal eats, almost down to the individual ‘nibble.’

Introduction

Adult laboratory mice typically eat three to six grams per day; however, very little is known about meal size, frequency, duration, or timing. While it is easy to observe gross changes in feeding activity on a daily basis, many treatment protocols have a significant impact on feeding behavior without affecting the total consumption over extended periods. Understanding these changes in feeding characteristics is essential to understanding the effects of drugs, nutrients, or other treatment protocols.

The system allows a researcher to observe the feeding behaviors of the animals in much greater detail than was previously possible. Traditional methods involved periodic hand weighing of food, which disturbs the animals and is error prone, or direct human observation of feeding activity. Neither method is easy to automate or scale up to many animals.

The FieldPoint system permits the simultaneous observation of 32 animals with a single piece of hardware. Furthermore, because FieldPoint is a stand-alone system, multiple FieldPoint units can be monitored by a single PC, providing an easily scaleable automation solution to a previously labor-intensive activity.

# System Architecture

The system has three logical components:

- The FieldPoint node

- The monitor application

- The analysis application

The FieldPoint node gathers data from the primary sensors and stores it in its memory. The monitor application retrieves the data from the FieldPoint node, performs some preliminary data processing, and then stores the results onto a local or network disk. The analysis application, which can run on the monitor application PC or a different PC, retrieves the data from the disk, performs analysis as interactively directed by the user, and then presents the data visually or as various output files.

# FieldPoint and National Instruments LabVIEW Real-Time Module are Key

By implementing the primary measurement algorithm in the FieldPoint controller, we were able to gather all of the raw data within the FieldPoint node. The PC may be left connected, but it is not required. This creates a robust solution, as network or PC outages do not impact the actual data collection.

Strain gages are used to measure the amount of food consumed as well as to indicate the animal’s activity level. Significant development effort went into the algorithm itself, with the result being that the system can accurately measure feeding activity as small as ~0.02 grams, even in the presence of temperature changes. The accuracy, resolution, and capability of the FieldPoint platform were essential to the development of the system.

# Flexibility of NI LabVIEW Speeds Development Process

Our primary measurement, the strain gage, produces a lot of noise relative to the size of our desired signal. By creating our application in a modular fashion we were able to quickly generate and evaluate different measurement algorithms to extract the desired signals. Additionally, because the algorithms could be developed directly on the PC before being targeted to the FieldPoint, development could proceed very rapidly through a number of algorithm candidates. The algorithm and various aspects of the system are covered by issued and pending patents.

# Modular Architecture Results in Flexibility and Reliability

The application was designed in a modular fashion. LabVIEW made it very easy to segment the application into manageable chunks and to perform unit testing. Data flow, in-placeness, and data compartmentalization let us rapidly and reliably prototype new features. It was not uncommon during our development process for alpha builds to be distributed to selected end users around the world the next day.

This modularity also let us extract pieces of the application for use in other tools such as QA fixtures and system evaluation tools. Using a shared library ensured that the results reported by the various tools will accurately reflect what the primary application would show. This also made it easy to develop spin-off products such as a small channel count unit for use as a sales demo. Pop out the FieldPoint I/O calls, replace with USB-6009 equivalents, build for the PC, and we had a portable system for demonstration purposes.

# Graphical Representation of Data

The system can generate records of thousands of bouts per day. Presenting the data in a manageable manner is essential to the researcher’s understanding. The graphing tools in LabVIEW are key to making a researcher-friendly user interface.

We created a data viewer tool,that allows the researcher to very quickly compare, organize, analyze, and extract the data. By initially presenting the data graphically, the user can rapidly identify which animals or time periods are of interest. Grouping functions allow the user to see averages or individual animals at the click of a (computer) mouse. The most common calculations are built into the viewer, and the raw data is also available for external processing.

Each graph line represents the feeding activity of a group of animals. The effects of the treatment protocol on group A, the red trace, are readily apparent compared to the two other groups. All of the detailed data behind the group totals can also be displayed.

Conclusion

Baseline feeding patterns have been established for various species of laboratory animals.  While the overall consumption rate for mice was well known, we can now show that within that total, the animals will eat 10 to 20 meals of approximately 0.3 to 0 .6 grams each.  The individual meals are further divided into 30 to 60 feeding bouts of approximately 0.01grams. Additionally, we can show quite clearly the relationship between external influences and feeding activity; for instance, rats typically eat three distinct meals during the night, while mice munch more.

This system gives the researcher data that was simply unavailable previously. The level of detail about the animal’s feeding, combined with the ease of formatting and presenting the data, allows the researcher to quickly and accurately assess the effects of treatment protocols.

Figure 1: The BioDAQ system architecture.

Figure 2: The data viewer analysis tool shows the effects of treatment.