Remote Experimentation with MEMS Devices

Category:

Mechatronics

Authors:
Bill Diong,, Jamie Smith, and Edward Kolesar, Texas Christian University.

NI Products Used:

Laboratory Virtual Instrument Engineering Workbench (LabVIEW 7.1.1)

NI Educational Laboratory Virtual Instrumentation Suite

PCI-6251 DAQ board

Measurement and Automation Explorer

PCI-1405 Image acquisition (IMAQ) card

IMAQ software

The Challenge:

Develop the means for a class of undergraduate students to conduct experiments involving unique and delicate instruments to characterize various given Micro-Electromechanical System (MEMS) devices.

The Solution:

Use National Instruments LabVIEW and other NI software and hardware to enable student access over the Internet to command appropriate inputs to the MEMS device under test, and to capture data from and images of that device.

Abstract:

Our main goal was to enable those students enrolling in the materials science course at Texas Christian University (TCU) to conduct an experiment via the Internet on MEMS devices that are located in a TCU research laboratory. A remote experimentation system based on NI hardware and software was developed and integrated in a timely and cost-effective manner. Subsequently, the system’s capability to allow the intended experiment to be performed by students over the Internet has been verified. Furthermore, the students have been surveyed concerning their experience, and the responses were overwhelmingly positive.

Introduction:

Mid-way through TCU’s ENGR 30014 – materials science course, the students learn about the electrical and thermal properties of materials. Recently, the instructor had been thinking it would be helpful for the students to carry out an experiment that served to demonstrate the synergistic linkage between these topics, while simultaneously illustrating the concept of device characterization. One tantalizing possibility was to have these undergraduate students run experiments on MEMS devices to determine the in-plane tip displacement versus absorbed power characteristics of a MEMS electrothermal actuator. .However, the risk of those students damaging delicate and expensive instruments, typically used exclusively for MEMS research, was an unacceptable one.

Knowing that NI had products that allowed remote access to and data acquisition from test equipment over the Internet, we decided to embark on a project to enable our students to conduct an experiment via the Internet on electrothermal actuators that are located in the TCU MEMS research laboratory. This would provide the students with a unique learning experience while imposing constraints on what they could do with the supporting equipment so as to minimize the risk of damaging it..

Experiment VI

This experiment VI was then developed successfully by the second author, who was a senior undergraduate student at the time with no prior NI LabVIEW experience, after only a short learning curve. The main functions of the VI are to

(i)	Apply a user-commanded voltage

(ii)	Capture images of the unpowered and powered actuator (so that its displacement can be determined by the user)

(iii)	Acquire and calculate the voltage, current, and power values associated with the MEMS device

(iv)	Plot the displacement versus power characteristic curve for the device under test.

NI ELVIS

The NI ELVIS was used as the hardware interface between the PCI-6251 DAQ board and the MEMS device. The user-commanded voltage on the VI front panel is applied by the DAQ board via a power amplifier to the series connection of a resistor (so that the supplied current can be determined from its voltage drop as acquired by the DAQ board) and the MEMS device; therefore the device's voltage is actually somewhat less than the commanded voltage. The exact value is calculated by subtracting the voltage drop across the current sense resistor from the commanded voltage, and that device voltage value is then displayed along with the device current and power values on the VI front panel.

Image Capture

It was decided that the images of the unpowered and powered MEMS device, as captured by a video camera through a microscope, should be embedded into the VI front panel.  Hence, an IMAQ PCI-1405 board and the Measurement and Automation Explorer (MAX) were used to perform the image acquisition by LabVIEW.

Displacement versus Power Plot

Our solution to implementing the MEMS tip displacement versus power plot was to use shift registers rather than local variables. By using shift registers, one has the ability to append the new data to the old data for any desired number of points.

Remote Access

LabVIEW 7.1.1 features the capability to have a VI viewed and controlled remotely over the Internet by using a Web browser. Therefore, we took the steps necessary to publish the MEMS experiment VI described above on the Web. Once the .htm file was available on the VI server PC, we were able to view and control the MEMS experiment VI front panel with a Web browser on a different PC. On some Windows PCs, we encountered the problem of being unable to view (and therefore to control) the VI front panel in the Web browser window; but this was easily remedied by installing the LabVIEW Run-Time Engine 7.1.1, which can be downloaded for free from the ni.com KnowledgeBase.

Access Limiting

Because the Web publishing tool described above does not allow the access limiting desired for this experiment (to prevent plagiarism), some modifications needed to be made.  With the built-in capabilities, either all users can view the VI front panel or none can.  But for our purposes, when one student is using the experiment VI, no other student should be able to view the first student’s work during his/her experiment. Our solution to the access limiting problem was to slightly alter the HTML code generated by the LabVIEW Web Publishing Tool and embed it in an ASPX file, which for security reasons resides on a gateway Web server for the TCU domain.

Consequently, the MEMS remote experiment URL provided to the students links them to a gatekeeper Web page (as constructed by the ASPX file) where they must press a “start Ttst” button.  Pressing that button then causes the ASPX file to call the experiment VI, which is on the MEMS lab PC and internal to the TCU network, by using that computer’s IP address.  If someone is already using the VI, the gatekeeper page then displays a message that asks the user seeking access to try again later. Otherwise, the user is connected to the experiment VI, and a time limit countdown begins for completion of the experiment. We do need to instruct the user that he or she should press the available “end test” button when the experiment has been completed rather than simply closing the browser window, otherwise the experiment VI will still appear to be in use to others seeking access until the time limit has been reached.  The time limit, button labels, width, and height of the VI front panel display and user prompts can all be changed in the ASPX file by modifying the appropriate lines of code.

Conclusions

NI hardware and software was readily applied to implement a remote experimentation system  that provides undergraduate students at TCU the unique opportunity to perform experiments on state-of-the-art MEMS devices while minimizing the risk of damage to delicate and expensive equipment. Furthermore, the students have been surveyed concerning their experience, and the responses were overwhelmingly positive. We believe that our broader, longer-term goal of enabling other TCU students to conduct experiments on systems located in any of the university’s science and engineering laboratories via the Internet is a step closer to reality.

Figure 1:  A flow chart of the MEMS experiment VI’s operations.