Jump-Start Your Sound and Vibration Application Development

Sound and Vibration Measurement Suite

Hands-On

NIWeek 2008

- Hands On with LabVIEW and the Sound and Vibration Express VIs

Exercise 1: Sound Level and Octave Analysis with the LabVIEW Sound and Vibration Toolkit

In this exercise you will use LabVIEW and the Sound & Vibration Toolkit to measure the sound level from a microphone and then perform third-octave analysis.

- Connect the microphone to channel 0 of the NI-9233.

- Launch LabVIEW and open a blank VI.

- Go to the block diagram by pressing Ctrl+E.  Right-click and place a DAQ Assistant from the Express > Input palette.

- Select “Acquire Signals”, “Analog Input”, and then “Sound Pressure” for the measurement type.  Next, a screen will show you the NI-9233.  Expand it so that you can see all four of the channels and select channel 0.  Click Finish.

- Configure the acquisition with the following settings:

- Enter the microphone sensitivity which is listed on the calibration certificate for your microphone.  If you don’t have a certificate then just leave it at “10”. Click OK.

- Press No when prompted about Automatic Loop Creation.

- Right-click on the data terminal of the DAQ Assistant and select Create > Graph Indicator.  This will create a graph on the Front Panel which displays the time domain data from the microphone.

- Go to the Sound and Vibration Express function palette by navigating to Addons > Sound & Vibration > S&V Express Measurements.

- Go to the Time-Domain Measurements palette and place the Sound Level Express VI on the diagram after the DAQ Assistant.  In the configuration page, select “A Weighting” on the Weighting tab.  On the Averaging tab, select only “Exponential” and uncheck all the other averaging options.  Click OK to apply your changes.  Note:  If you get a dialog asking you to save changes to the VI, click “Save”.

- Wire the data output from the DAQ Assistant to the input signal of the Sound Level Express VI.

- Go to the Front Panel of the VI and place a Meter indicator.  You can find this in the Numeric indicator palette.

- Change the maximum value of the meter from 10 to 120.  You can do this by double-clicking on the 10 and typing a new value.

- Return to the block diagram and connect the exponential output of the Sound Level Express VI to the meter indicator.

- To make this VI run continuously, we need to add a While Loop around all of the code we have written so far.  The While Loop can be found on the Programming > Structures palette.  Here’s what your VI should look like by now:

- Now you are ready to run this VI.  Go to the Front Panel and hit the Run button.  As the VI is running, make some noise that your microphone will pick up and watch your time domain signal and meter change values.  After you are done with this, save your VI as Sound Level Meter.vi on the Desktop.

- Let’s add a 1/3 Octave Analysis.  Go back to the block diagram and place an Octave Analysis Express VI inside the While Loop.  This can be found in the same S&V Express palette under Frequency-Domain Measurements.

- Configure the Octave Express VI to perform third-octave analysis over the frequency range of 20Hz to 20kHz with A-Weighting.  Click OK.

- Connect the data output of the DAQ Assistant to the input signal of the Octave Express VI.  Right-click on the octave output and select Create > Indicator.  The block diagram should look similar to this:

- Go to the Front Panel and run the VI.

Challenges:

- Add a limit test to the sound level measurements

- Compare a full octave analysis to 1/3 octave analysis

- Calibrate your microphone

Exercise 2: Vibration Analysis with the LabVIEW Sound and Vibration Toolkit

The DSA Demo Box has two accelerometers which measure the vibrations caused by an onboard unbalanced fan.  In this exercise you will use the Sound and Vibration Toolkit to measure the vibration level of these accelerometers and perform a power spectrum during live acquisition.

- Connect channel 0 to the X Acceleration and channel 1 to the Y Acceleration of the Demo Box.

- Open a new VI in LabVIEW.

- Go to the Block Diagram and place down the DAQ Assistant.

- Configure the DAQ Assistant for an “Acceleration” measurement task.  Hold down the Ctrl key to select both channel 0 and channel 1 from the NI-9233.  Click Finish.

- Configure the acquisition with the following settings:

Note: Make sure to set the sensitivity for both accelerometers (174mV/g).

- Press “No” when prompted about Automatic Loop Creation.

- Go to the Sound and Vibration Express palette and place the Vibration Level Express VI after the DAQ Assistant.

- Configure the Vibration Level Express VI to work for N channels and to calculate only the Running RMS level over a 1 second period.  Click OK to apply your settings.

- Wire the data output from the DAQ Assistant to the input signals of the Vibration Level Express VI.

- Right-click on the Running RMS output and select Create > Indicator.

- Place a While Loop around the code.

- Go to the Front Panel and run the VI.  The chart displays the running RMS vibration levels of each accelerometer.  On the DSA Demo Box, switch to the unbalanced fan and use the knob to change the speed of the fan.

- Save the VI as Vibration Level.vi.

- Go to the block diagram and place the Power Spectrum Express VI inside the While Loop.

- Configure the Power Spectrum measurement to work for N channels.  Choose how you would like the units to be displayed (RMS, Peak, etc.) and if you would like any averaging.  If you have no preferences, just use the default settings.  Click OK to apply the settings.

- Connect the data output of the DAQ Assistant to the “input signals” input of the Power Spectrum Express VI.  Right-click on the spectrum output and select Create > Indicator.

- Go to the Front Panel and run the VI.  Vary the speed of the fan and watch how the power spectrum changes.

Challenges:

- Find the peak amplitude and corresponding frequency in the spectrum

- Perform a Zoom FFT around this peak

- Calculate the power between 50Hz and 75Hz.

- Hands On with the Sound & Vibration Assistant

Exercise 3:  Sound Level and Octave Analysis with the Sound and Vibration Assistant

In this exercise you will use the Sound and Vibration Assistant to measure the sound level from a microphone, perform a third-octave analysis on the data, and save the data to file.

- Connect the microphone to channel 0 of the NI-9233.

- Select Start >> Programs > National Instruments > Sound and Vibration > Sound and Vibration Assistant to Launch the Sound and Vibration Assistant.

Notice that the Sound and Vibration Assistant is split into three main views: the Project View on the left, the Data View in the middle, and the Help on the right.

- The Sound and Vibration Assistant uses Steps to acquire, generate, analyze, load, and save signals.  Place a step that acquires sound pressure data from the NI 9233 by right clicking on the Project View and selecting Acquire Signals > DAQmx Acquire > Analog Input > Sound Pressure

- Select channel “ai0” and press OK.  When you place the DAQmx Acquire Step, the Data View tab will automatically change to the Step Setup tab.  This view allows you to configure all aspects of your acquisition step including number of channels, excitation levels, and timing information.

- Configure the DAQmx Acquire step with the following settings:

- Enter the microphone sensitivity which is listed on the calibration certificate for your microphone.  If you don’t have a certificate then just leave it at “10”.

- Before going any further, you can preview the time domain data from the Step Setup tab by clicking the Run button near the top of the screen.

- Click on the Data View tab and notice that it is blank.  You can configure the Data View to show only the signals that you wish to see.  Use your mouse to click on the Step output labeled “Sound Pressure” and drag it onto the blank data view.  Your time domain data should immediately appear in the Data View.

- Use an Analysis Step to take a Sound Level Measurement from this time domain data.  Right click on the Project View and select Analysis > Time Domain Measurements > Sound Level.  Again you will notice that the view tab automatically changes to Step Setup.

Note:  This is a good place to mention that every step in the Sound and Vibration Assistant has inputs and/or outputs.  The DAQmx Acquire step has only one output, Sound Pressure data.  The Sound Level step however has several sound level outputs and requires one input.  Sound Pressure has been chosen by default but you can change the source (if applicable) by clicking on the input’s drop down menu.

- This particular step has three configuration tabs near the bottom of the screen: Input, Weighting, and Averaging.   Configure this step to include A Weighting and to only compute Exponential averaging.

- Switch back to the Data View.  Click on the Sound Level output exponential and drag it onto the Data View.  Notice that a new chart and table was created to display this data.  Since the “exponential” output is a different data type than your currently displayed “Sound Pressure”, the Sound and Vibration Assistant automatically creates a new display in the Data View.

Right click the sound level chart and select View As > Meter.  Right click the meter and turn off Autoscale.  You can then double click the maximum value on the meter and change it to 100.

- Now let’s add a 1/3 Octave Analysis.  Place an Octave Analysis step from the Frequency Domain Measurements.

- Configure the Octave Analysis step to perform third-octave analysis over the frequency range of 20Hz to 20kHz with A-weighting.  Notice how the step’s preview adjusts to reflect the changes that you make.

- Switch back to the Data View and drag the “octave” output onto the Data View.

- Stop the Project.

- Finally, let’s save the octave analysis data to file.  Place a Save to ASCII/LVM step as the last step on the project view.  The Step Setup tab for this step should have two configuration tabs: Signals and File Settings.  Configure this step to save the octave data to Exercise1_Data.txt.   Also configure the step to Overwrite once, then append to file if the file already exists.

- Run your project and accumulate a few seconds worth of data.

- The data is being saved as a tab delimited text file and can be viewed in Microsoft Excel.  Stop your project, launch Microsoft Excel and open the file Exercise1_Data.txt.

Challenges:

- Add a limit test to the sound level measurement.

- Compare each of the different types of sound level measurements on the same chart.

- Record the sound pressure data and analyze it in Playback Mode.

  - Press the Record Button and include the Sound Pressure signal.

  - Run the project for a few seconds and then stop it.  A new tdms Log will appear in the bottom left hand corner of the Project View.

  - Now in the top left hand corner of the Project View change your mode from Monitor/Record to Playback.   You can place new analysis steps in the playback mode and analyze recorded data offline.

Exercise 4:  Vibration Analysis with the Sound and Vibration Assistant

The DSA Demo Box has two accelerometers which measure the vibrations caused by an onboard unbalanced fan.  In this exercise you will use the Sound and Vibration Assistant to measure the vibration level of these accelerometers and perform a power spectrum during live acquisition.

- Connect channel 0 to the X Acceleration and channel 1 to the Y Acceleration of the Demo Box.

- Launch the Sound and Vibration Assistant.

- Place a DAQmx Acquire step on the Project View, but this time let’s use the Channel View.  The Channel View is a convenient way to organize and configure all your National Instruments hardware for a project.  Select channel 0 and channel 1 from the Channel View.

- Notice how a DAQmx Acquire step was created and it automatically included channels 0 and 1 from your NI 9233.  From the channel view, change the measurement type of each channel from “Voltage” to “Acceleration.”  This will expose extra options for configuring current excitation and sensor sensitivity (for data scaling).

- You can also use the Channel View to quickly configure both channel and device settings.  Change the timing and channel settings to reflect the following:

Hint:  Timing information is considered a device property, so you will need to select the 9233 itself to configure those characteristics in Channel View.

- Double click the DAQmx Acquire step and notice your changes show up here as well.  Click the Run button and preview the time domain data.  Switch to the Data View and plot both channels of the Acceleration output onto the blank display.

- Place a Vibration Level analysis step from the Time Domain Measurements palette.  Configure the step to calculate only the Running RMS level over a 1 second period.

- Drag the “running RMS” output onto the Data View to display the vibration level for both channels.

- Place a Power Spectrum analysis step from the Frequency-Domain Measurements palette.  Choose how you would like the units to be displayed (RMS, Peak, etc.) and if you would like any averaging.

- Drag the “spectrum” output onto the Data View to display the Power Spectrum data for both channels.

- Vary the speed of the fan and watch how the power spectrum changes.

- Stop the Project.

- Finally, save the vibration data.  You could use the Save steps that were used in Exercise 3 but this time let’s use a drag-and-drop feature of the Sound and Vibration Assistant to move the current iteration of data directly to Excel.  Open Microsoft Excel and place it next to your Project View.  Click on the DAQmx Acquire output labeled “Acceleration” and drag it over to the blank Excel spreadsheet.  You can use this same technique to save data for any step’s output.

Challenges:

- Find the peak amplitude and corresponding frequency in the spectrum

- Perform a Zoom FFT around this peak

- Calculate the power between 50Hz and 75Hz.

- Log the vibration data with the record button, and analyze it in playback mode (see page 9 for more info).

- Hands On with LabVIEW and the Order Analysis Express VIs

Exercise 5:  Order Analysis with LabVIEW and the Sound and Vibration Measurement Suite

Use the NI-9233 and the Order Analysis Express VIs to make some basic rotating machinery measurements such as a speed profile and order power spectrum as well as use a color map to display results.

- Connect the Y Acceleration output of the DSA Demo Box to channel 0 of the NI-9233 and connect Tach Out to channel 1.

- Open a new VI in LabVIEW.

- Go to the block diagram and place down the DAQ Assistant.

- Configure the DAQ Assistant to acquire an “acceleration” measurement on ch 0.  Click Finish.

- Configure the measurement with the following settings:

- Now we’ll add the tachometer channel to the same task.  Click on the button for creating a new virtual channel.

- Select the measurement type to be “voltage” and from channel 1 of the NI-9233.  You should now see two different measurement types in your task.

- Click OK to apply your settings to the DAQ Assistant and press No when prompted about Automatic Loop Creation.

- Right-click on the data output terminal and select Create > Graph Indicator.

- Place the Split Signal function next to the DAQ Assistant.  You can find this function at Express palette > Signal Manipulation >Split Signal.  Expand the split signal so that it has two outputs.

- All Order Analysis Express VIs can be found in the palette located at Addons > Sound and Vibration > S&V Express Measurements > Order Analysis Express Measurements.

- Place an Analog Tacho Processing Express VI on your block diagram.  The tachometer from the fans in the DSA Demo Box has 2 pulses per revolutions.  Change the Pulse/Revolution to 2 and select OK.

- This Express VI calculates the speed profile of the machinery from the tachometer signal.  Wire the tachometer signal (bottom half of the Split Signal VI) into the analog tacho signal input.

- Next place the Statistics Express VI, located on the Express palette > Signal Analysis > Statistics.  Configure the VI to calculate only arithmetic mean.

- Wire the speed profile to the Signals input of the new express VI.

- Go to the front panel and place down a gauge numeric indicator.  You can find this on the Express palette > Numeric Indicators.   Increase the maximum value of the gauge to be 6000.  You can do this by double clicking on the value 10 and then typing 6000.  We will use the gauge indicator to display the current speed of the fan.

- Return to the block diagram.  Wire the output of the Statistics Express VI to the gauge indicator.

- To make this program run continuously, add a While Loop around all of your code.  Now your VI should look similar to this:

- Go to the front panel and run the VI.  Use the knob on the Demo Box to change the fan speed.  When you are finished testing your VI save it as Order Analysis.vi.

- Next we will add an Order Power Spectrum to our VI.  Retrieve the Order Spectrum Express VI and place it on your block diagram.  Press OK and keep all of the default settings.

- Wire the input signal terminal to the top half of the Split Signal VI and also connect the speed profile to the output of the Analog Tacho Processing VI.

- Right-click on the spectrum output and select create > indicator.

- Your VI should now look something like this:

- Go to the Front Panel and run the VI.  Vary the speed of the fan and switch between the balanced and unbalanced fans.

- Hint: Turn off AutoScaling on the Order Power Spectrum graph so you can easily see how things are changing.  Set the scale on the X-axis to be 0 to 10 and the Y-axis to be from -100 to 0.

- The last thing we will add is a color map.  Retrieve the Spectral Map Express VI and place it on your Block Diagram.

- Configure the Plot type to be Frequency – Time and select OK.

- Wire the input signal and speed profile inputs in the same way as you did with the Order Spectrum VI.

- Go to the Front Panel and place a color map from the Addons > Order Analysis palette.

- To finish the VI go back to the Block Diagram and wire the color map to the output of the Spectral Map VI.  Here is what your finished VI should look like:

- Run the VI and vary the speed of both the balanced and unbalanced fans.

Hands On with the Sound & Vibration Assistant and Order Analysis

Exercise 6:  Order Analysis with the Sound and Vibration Assistant

In this exercise you will use the Sound and Vibration Assistant with Order Analysis to make rotating machinery measurements such as a speed profile, order power spectrum, and order tracking.  You will also learn how to display the results using a color map or a waterfall plot.

- Connect the channel 0 of the NI-9233 to the X Acceleration output of the Demo Box.  Likewise connect channel 1 to the Y Acceleration output and channel 2 to the TACH OUT output.

- Select Start > Programs > National Instruments > Sound and Vibration >> Sound and Vibration Assistant to Launch the Sound and Vibration Assistant.

- Place a DAQmx Acquire step on the Project View, but again let’s use the Channel View.  The Channel View is a convenient way to organize and configure all your National Instruments hardware for a project.  Select channel 0, channel1, and channel 2 from the Channel View.

- Notice how a DAQmx Acquire step was created and it automatically includes channels 0, 1 and 2 from your NI 9233.  We want to measure the vibration from channels 0 and 1 and a tachometer pulse from channel 2.  From the channel view, change the measurement type of channel 0 and 1 from “Voltage” to “Acceleration.”  This will expose extra options for configuring current excitation and sensor sensitivity (for data scaling).  Leave channel 2 as a voltage measurement.

- You can also use the Channel View to quickly configure both channel and device settings.  Change the timing and channel settings to reflect the following:

Hint:  Timing information is considered a device property, so you must select the device itself (see above) to configure those characteristics in Channel View.

- Double click the DAQmx Acquire step and notice your changes show up here as well.  Click the Run button and preview the time domain data.  Switch to the Data View and plot all three channels of the Demo Box output by dragging the “Acceleration” and “Voltage” step outputs onto the blank Data Display.

- To perform Order Analysis, first add the tachometer processing step to the project.  Right-click on the Project View and select Analysis > Order Domain Measurements > Tachometer Processing.

- On the Input tab of the Step Setup, choose the channel acquiring the tach pulse.  In this case the tach pulse is being acquired by channel 2.

- On the Configuration tab change Pulse/Revolution to be 2 (see Demo Box).

- Return to the Data View where you are currently viewing the time domain data (see step 6).  Right-click on the current data display and select Data View > Add Display Below.  Drag the “current speed” step output from the Project View onto the new data display.  The current rotational speed is now displayed in RPM.  Right click the new display and select View As > Gauge.

- Next we will add an Order Power Spectrum to our project.  Right click on the Project View and select Analysis > Order-Domain Measurements > Order Spectrum.

- Again the Step Setup tab automatically comes up for you to configure your settings.  Both accelerometer channels are selected by default as is the speed profile generated by the Tachometer Processing step.  Configure the Spectrum and Processing Settings to your liking and display the order spectrum on a new data display on the Data View.

- If you have not done so yet, Run the project and adjust the size of your data displays to your preference.  Use the knob on the demo box to change the fan speed and switch the demo box outputs between the balanced and unbalanced fans.  Disable the Order Spectrum y-axis autoscaling to better see the changes in order magnitude.

- Now let’s add a colormap of the frequency vs. time data.  Right click on the Project View and select Analysis > Order-Domain Measurements > Spectral Map.  When the Step Setup tab appears, configure the Input Signal to be either accelerometer of the demo box.  The speed profile from the Tachometer Processing step is automatically configured.  Make sure “Export colormap” is checked and change the Time Interval(s) to 10ms under the Configuration tab.

- Display the colormap under the order spectrum display.  To better view the data, double-click the maximum y-axis value and change it to 1000Hz.  Your Project View and Data Display should look similar to the screenshot below.

Challenges:

- Use the Order Tracking step to track the magnitudes of both accelerometers’ second order over time.

- Extract the first order from both accelerometers and graph them on an X-Y plot.

- Log the accelerometer and tachometer data with the record button, and analyze your data in playback mode (see page 9 for more info).

Thank you for test driving the Sound and Vibration Measurement Suite.  Please let us know what questions, suggestions, and comments you may have.  Visit www.ni.com/soundandvibration for more information.